General Response,

Dear Editor,

We would like to thank yourself and all of the reviewers for taking the time to provide much appreciated feedback of our work. We believe that with your help we have been able to produce a superior manuscript that will provide the Journal of Visualized Experiments and wider bioprinting and food printing communities with a valuable piece of work. We have taken the time to address each of the concerns.

We noted a few recurring themes in the feedback,

- The title mentioned that we designed a “Bioink” printer but did not use a bioink to characterise the performance.

- More information about how to go about printing an object is needed.

- The accuracy of the printer needed to be validated through our results.

- The layout of the protocol was hard to follow and did not proceed in a stepwise manner.

Firstly, to make sure that our printer has wide appeal, we have introduced another material into our results section, Pluronic F-127 (Sigma Aldrich) as a model for bioinks to demonstrate the application of our printer to bioprinting.

Secondly, we have provided more information on how to go about printing an object by introducing a new section into the protocol, ‘8. Preparation for 3D Printing’. This will help bridge the gap between constructing an ADDME 3D printer and generating gcode to print an object.

Thirdly, to better characterise materials in a way that displays the accuracy and fidelity of our 3D printer we have added a new section in the results. In this section we use a 3D point cloud comparison of a computer generated object and the same object that has been 3D printed and 3D scanned.

Lastly, we have gone through our protocol and made sure that it proceeds in a step wise manner that is better aligned to the journals standards.

Again thank you for your constructive feedback, we hope our paper is now suitable.

Kind regards,

Matthew Lanaro

| Editors Comments | Comment | Response | Line Numbers |
| --- | --- | --- | --- |
| E1.1 | Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues | We have made a number of changes to all the sections within the manuscript. | All |
| E1.2 | For each step/substep, please ensure you answer the “how” question, i.e., how is the step performed? Alternatively, add references to published material specifying how to perform the protocol action. If revisions cause a step to have more than 2-3 actions and 4 sentences per step, please split into separate steps or substeps. | Thank you for your feedback, we have gone through and made significant changes to the wording of the protocol. | Protocol |
| E1.3 | Do you mean the Table of Materials here? | Yes, the name has been changed to table of materials. | Line 128-129 |
| E1.4 | Should part 9 also be listed here? | Yes, part 9 has been added. | Line 163 |
| E1.5 | Can you expand a little more here about how to actually print something? | We have added an additional section into the protocol “8. Preparation for 3D Printing” in addition to figure 11. | Lines 377 - 400 |
| E1.6 | Please correct the inverted parentheses around volume numbers. | The inverted parentheses have been corrected. | References |
| E1.7 | Please ensure the Table of Materials has information on all materials and equipment used, especially those mentioned in the Protocol. | We have included the materials and equipment used throughout the protocol. | Table of Materials |

| Reviewer 1 | Comment | Response | Line Numbers |
| --- | --- | --- | --- |
| R1.1 | The work in question does not reach an appropriate level to be published. | We have made additional changes to further support our work. Firstly, we have included an ink (Pluronic F-127) as a model for bioinks to demonstrate the use of our 3D printer in the biofabrication field. Secondly, we have added a 3D scanning comparison of a 3D printed object with the CAD model to verify the accuracy and fidelity of the printer. We trust that these changes raise the quality of our work to an acceptable level. | General |

| Reviewer 2 | Comment | Response | Line Numbers |
| --- | --- | --- | --- |
| R2.1 | The title indicated the bioink (polylactic acid) was also suitable for this equipment. However, author only used cream and chocolate as the materials, no evidence can prove polylactic acid can be the materials if printing by this equipment. | We have included Pluronic F-127 (Sigma Aldrich) as an ink (a model for bioinks) to demonstrate our printers usefulness to bioprinting. Please note that polylactic acid is only used as a material to construct the printer, not test it. We acknowledge that this is confusing and we have changed our language in the protocol to make this clearer. | Line 467 |
| R2.2 | Author mentioned the cost of this equipment was “cheap”, as far as I know, the cheapest extrusion 3D printer was less than 500 US dollars. Cost accounting is necessary for this manuscript. | Yes, we found and note in the introduction that the RepRap v2 is $400 USD. In the discussion we mentioned that the cost of our 3D printer was reduced to $245 USD. We have provided a detailed cost breakdown of each component in the table of materials. | Lines 84, 460 and the table of materials |
| R2.3 | Except the accuracy, some other parameters such as printing speed, size of syringe and compactness is also important for 3D printing, author need more statement for this part. | We have added expanded the discussion to include information on the importance of various printing parameters. Additionally, we have added a section in the results comparing a 3D scan of an object to the CAD model. This will characterise the geometric accuracy of the printer. | Line 504 - 506 |
| R2.4 | More optional accessories are good for this equipment. For instance, nozzles of different calibre and material, or syringes of different size. It can extend the scope of application of this equipment. Author can added more modularity components for this equipment to make it easy to use. | We aim to support the modularisation and customization of our equipment by releasing it as open source. This gives researchers full access to all the CAD files, enabling further accessories to be made. |  |
| R2.5 | The introduction part needs to be reversed. Author should focus on the development of 3D printer, listing the deficiencies of 3D printer existed, at the same time emphasize the superiorities of designed 3D printer. | We thank the reviewer for their suggestions regarding the layout of the introduction. We have made several changes to the arrangement of the introduction to better assist the reader to introduce our subject matter. | Introduction |

| Reviewer 3 | Comment | Response | Line Numbers |
| --- | --- | --- | --- |
| R3.1 | Please take out "bioinks" from the title. Bioinks is essentially defined as softer biomaterials printed along with live cells. Any hand cream cannot be considered as a bioink. Unless the authors would like to demonstrate the bioprinting with a more popular bioink material like alginate, GelMA etc. (https://www.sciencedirect.com/science/article/pii/S0734975016301719) | We have included Pluronic F127 as an ink (a model for bioinks) to display the versatility of our 3D printer as a bioprinter. | Results |
| R3.2 | The printed structures have not been measured (width/height of the lines in Fig. 11). There is no mention of the number of printed samples. It is crucial to demonstrate the print fidelity of this low-cost assembled printer. The print tolerances need to be mentioned. | We have added a results section comparing a CAD model with a 3D scan of the same printed model. We are able to report the geometric difference between the 3D model and 3D scanned model down to 50 µm. | Line 441 - 453 |
| R3.3 | The operation of the printer is not written clearly for a user to follow. Ideally, the operation should be in a step-wise protocol format. | We thank the reviewer for this comment, we note that the editor had a similar comment. We have gone through our protocol and made changes to insure that the protocol is clearly written and better follows the step-wise format. | Protocol |
| R3.4 | For any extrusion-based printing, shear-thinning is a major property of the material/biomaterial. The manuscript current doesn't mention anything about the material properties. | Yes, shear thinning is an extremely important materials property, especially for chocolate. We have aimed to make the results section about validating the protocol (in this case the printer) and not the materials. |  |
| R3.5 | What does the parameter "flow" imply in the table. | We have changed the term “flow” for extrusion. Extrusion (calculated from equation 2) is a dimensionless value that tells the electronics control board how much material to extrude. We can apply a scalar value to this extrusion, to extrude more or less than the value calculated from equation 2. |  |
| R3.6 | Is there any way to calculate the extrusion pressure? | Yes, for newtonian fluids the relationship between shear rate is easy to calculate, shear rate = 8* velocity/diameter. Other research papers measure the extrusion pressure directly with a sensor. We aimed to include results that validated the design of our printer and focused on geometrical accuracy. |  |
| R3.7 | What are the sensitivities for the machined heating jacket and silicon heating jacket? | The thermal limits of the heating supplies should not matter as the maximum temperature is limited by the 3D printed housing materials. The geometric sensitivities are forgiving as the housing, not the heater is the load bearing element. Put another way, the heaters are low tolerance parts. |  |
| R3.8 | Can the authors provide more detailed instructions for electronics and wiring? It will be nice to include an illustrated figure on this. | We have included an image of the wiring in figure 10 and added some explanation. | Line 344 – 347, figure 10 |
| R3.9 | For the software setup, it will be nice to provide screenshots of the installation process so the readers know where exactly to change the settings and configurations. | We decided to include additional navigation information so the readers can better follow along. | Line 325 - 328 |
| R3.10 | Please include important notes and safety instructions for the assembly protocol. | We have expanded the safety notes in a number of locations throughout the document. |  |
| R3.11 | Can the authors provide a summary or list of basic tools needed to assemble the 3D printer? For example screwdriver, Hex keys, etc. | We have updated the table of materials to include the tools needed to assemble the 3D printer. |  |
| R3.12 | It will be helpful to add numbers, callouts for each of the parts used in the machine. | We have added a number near each component in the figures which links to each respective component in the table of materials. |  |
| R3.13 | Line 352- The labels should be in lower case. | We have changed the labels. | Line 352, now Line 393 |
| R3.14 | It would be helpful to put all the online links as an Annex at the end of the paper | We have included a link to all online links in a note at the start of the paper. | Line 111 - 121 |
| R3.15 | Line 408 - Please take out "biofabrication". The printer has not been demonstrated for any biomaterials printing. | We have included an ink (as a model for bioinks) (Pluronic F-127) to demonstrate the versatility of our printer for biofabrication. |  |
| R3.16 | Some of the figures have a dark background and hence when printed on paper do not come out clearly. Since it is a protocol paper and it is reasonably expected for potentially interested users to print it and follow for assembly. | We recognise this may be a problem, unfortunately we do not have any other colour of acrylic available (such as black) to remake those figures. Due to time constraints we are not able to wait until new acrylic arrives. |  |