- Prepare specimen, Labware, and reagents

2.1   Transfer 5 L pooled healthy human plasma into to a Polypropylene, 96-Round Deep Well Plate.

Note: See Materials table for reagents and supplies used in this protocol. TPCK Treated Trypsin was purchased and trypsin to substrate ratio and incubation time was optimized specifically. If a different grade of trypsin is used, such as sequence grade recombinant trypsin, the enzyme to substrate ratio and incubation time should be tested and optimized.

- Operating Procedure

- Double-click the software icon.

- Under the “Method” tab, select “Home All Axes” to orient and prepare the automated liquid handler. Ensure all workstation syringes contain no visible air bubbles.

- Under “File”, select “Open”, and then “Method”.

- Select the method (Figure 6).

- Start the method by clicking the green triangle-shaped “Run” icon.

- To have an autosampler plate prepared at the end of the method, enter the value ‘true’ in the “Enter a value to use for ‘Autosampler’” prompt and then click “OK”. If an autosampler plate is not being prepared, enter the value ‘false’, and then click “OK”.

- Enter the value ‘1’ in the “Enter a value to use for ‘firstcolumn’” prompt and then click “OK”.

- Enter the value ‘12’ in the “Enter a value to use for ‘lastcolumn’” prompt, and then click “OK”.

Note: This step and step 4.7 tells the workstation to perform a digestion on 12 columns of a 96-well plate, resulting in all wells of the plate being used for digestion

- If a sample plate is being used with sample volumes of at least 20 µL, enter the value ‘true’ in the “Enter a value to use for ‘SamplePlate’” prompt, and then click “OK”. If a sample plate is not being used, enter the value ‘false’, and then click “OK”.

Note: If a sample plate is not being used, add 5 µL of sample to each corresponding well of the Reaction Plate.

  - Follow the directions in the Guided Labware Setup window and click “Continue”.

Note: The next window describes the layout for the “Reagent Plate” to be prepared (Figure 7, supplemental table 1). The following volumes will be aliquoted into each of the 8 wells of a single column in the 1 mL Deep Round 96-well plate:

Column 1: 540 µL of Reaction Mix 1.

Column 2: 50 µL of Cysteine Block.

Column 3: 730 µL of Reaction Mix 2.

Column 4: 130 µL of Trypsin.

Column 5: 130 µL of Quench Solution.

Column 5: 1090 µL of Mobile Phase Solution (If the user entered ‘true’ in step 6).

  - After clicking “Next”, the following window describes the minimum volume (20 µL) needed to be aliquoted into the Sample plate. If the user entered the value ‘false’ for the prompt involving the Sample Plate (step 9), the user will be prompted to add 5 µL sample to the Reaction Plate.

  - Click “Next”.

Note: The next window shows that an empty 90 µL tip box must be placed on the automated liquid handler deck.

  - Click “Next” again, laying out the automated liquid handler’s deck as specified and illustrated (Figure 8) including the Reagent Plate, the Reaction Plate, the Sample Plate, the Autosampler Plate, 6 full 90 µL tip boxes, an empty 90 µL tip box, and a full 230 µL tip box.

  - Click “Finish” to begin the method.

Note: After clicking “Finish”, the user cannot reach into the automated liquid handler. This will ‘break the light curtain’ of the machine and stop the liquid handler as a safety precaution.

  - At the ‘Continue after centrifugation’ prompt, retrieve the Reaction Plate and centrifuge it for 30 minutes at 3,000 RPMs.

  - After the centrifugation is complete, return the Reaction Plate to the automated liquid handler to the position it was in before the centrifugation, and click “Continue”.

Note: The automated liquid handler will stop again when the method is complete. If the user entered ‘true’ for step 6, the Autosampler Plate can then be removed from the workstation and placed into the autosampler of a liquid chromatography instrument for analysis.

- LC-MSMS

4.1 	Resolve peptides on a high flow HPLC comprising of a C18 2.1mm x 100mm, 3.5µm column with a flow rate of 0.25 mL/min and analyzed inline on a triple quadrupole mass spectrometer.

Note: After peptide digestion, the targeted LC-MSMS method to quantify peptides from robotic prepared samples is described in detail1 following a brief description of LC-MSMS

4.2	 Set the column oven temperature to 40 °C. Use buffer A (2% ACN, 98% water, 0.1% formic acid) and buffer B (95% ACN, 5% water, 0.1% formic acid) as the two mobile phases.

4.3	Equilibrate the column with 5% B for 5 minutes after loading. Elute the peptides over 30 minutes with a linear 5% to 35% gradient of buffer B.

4.3.1    Recycle the column before loading the next sample by washing with 98% B for 10 minutes and then 5% B for 5 minutes.

4.3.2    For the on-line diversion, divert the post-column eluent to waste before it entered the ion source using a two-phase switching valve.

4.4 	Process the MRM data.