TITLE:

A Novel Non-Invasive Method for the Detection of Elevated Intra-Compartmental Pressures of the Leg.

AUTHORS AND AFFILIATIONS:

Matthew J. Herring, MD; Erin Donohoe, MD; Meir T. Marmor, MD

Department of Orthopedic Surgery, University of California, San Francisco, CA, USA

matthew.herring@ucsf.edu

erin.donohoe@ucsf.edu

Corresponding Author:

Meir Marmor, MD

Meir.marmor@ucsf.edu

KEYWORDS:

Compartment Syndrome, Non-invasive pressure monitoring, intra-compartmental pressure, tibia fracture, decompressive fasciotomy, ultrasound

SUMMARY:

An ultrasound probe coupled with a pressure sensor is used to assess the intra-compartmental pressure of the leg by directly measuring the compartment fascia flattening pressure (CFFP). This non-invasive protocol will provide reliable assessment of the pressure inside the anterior muscular compartment of the lower leg.

ABSTRACT:

Acute Compartment Syndrome is a devastating consequence of musculoskeletal trauma. Currently the diagnosis is based on clinical signs and symptoms, and while adjuncts such as invasive intra-compartmental pressure measurements are often used to corroborate the physical exam findings, there remains no reliable objective test to aid in the decision to perform a decompressive fasciotomy. In a cadaver model of compartment syndrome, an ultrasound (US) based method has been shown to be a reliable measurement of increased intra-compartmental pressure. An absolute pressure of > 100 mbar or a difference of 50 mbar in the CFFP between the legs can be considered pathologic. Using an ultrasound transducer, coupled with a pressure sensor, the pressure needed to flatten the superficial fascia of the anterior compartment of lower legs (Compartment Fascia Flattening Pressure [CFFP]) can be measured. The CFFP of the injured leg is compared to the CFFP of the uninjured leg. This US measured index can then serve as an adjunct to the physical exam in evaluating injured lower extremities and assessing the need for decompressive fasciotomy. The advantages of this protocol include: being a non-invasive method and an easily reproducible technique.

INTRODUCTION:

The purpose of this novel protocol is to evaluate the intra-compartmental pressure of the leg in a non-invasive way by using an ultrasound transducer coupled with a pressure sensor. Compartment syndrome is a well-known complication in musculoskeletal trauma and is the result of elevated intra-compartmental pressures which compromise the perfusion of tissues leading to irreversible ischemic injury if not intervened upon. Its diagnosis is primarily based on clinical exam. Unfortunately, clinical exam alone has been shown to have poor sensitivity and specificity in the diagnosis of compartment syndrome1. Invasive intra-compartmental pressure measuring by needle manometry is commonly used in the assessment of compartment syndrome2-4. The disadvantage of needle manometry is that it is an invasive procedure, being quite uncomfortable for the patient and not amenable to serial measurements. Furthermore, invasive intra-compartmental pressure measuring measurement is not routinely performed by some surgeons due to considerable disagreement over the appropriate threshold pressure for diagnosis of ACS and the high (35%) false positive rate and variability of a single invasive measurement5-7.

It has been previously demonstrated in artificial, animal, and human cadaver models that a rise in the intra-compartmental pressure results in a decreased elasticity of the fascia overlying the anterior compartment suggesting a correlation between the intra-compartmental pressure and compliance of the facial compartment as measured by ultrasound8-10. Further work has described using the pressure required to flatten the fascia of the anterior compartment as measured by ultrasound (Compartment Fascia Flattening Pressure [CFFP]) as a surrogate measure of the intra-compartmental pressure11. The anterior compartment fascia can be readily visualized with standard ultrasonography and, with the addition of a pressure sensor to the probe, CFFPs can be easily and reliably measured. This protocol offers a non-invasive diagnostic alternative for evaluating intra-compartmental pressures and can be performed quickly and serially with little to no patient discomfort.

PROTOCOL:

This experimental protocol has been reviewed and approved by our hospital’s institutional review board (IRB) and follows all guidelines set forth by our institution’s human research ethics committee.

1. Identify the fascia of the anterior compartment of the leg

1.1. Identify a starting site for the ultrasound probe. Place the probe directly over the anterior compartment, just lateral to the tibial crest and one handbreadth distal to the tibial tubercle.

1.2. Identify the fascia as the bright, white strip just below the subcutaneous tissue on the ultrasound. Move the probe until the fascia is seen at the point where it attaches to the tibial crest.

1.3. Note the geometry of the fascia. With minimal pressure applied, this should be a convex shape.

2. Measure the Compartment Fascia Flattening Pressure of the injured leg

2.1. Once the fascia of the anterior compartment is identified, apply slow and steady pressure to the ultrasound probe until the fascia changes shape from convex to completely flat. Make several attempts including passing the flattening point to produce a concave shape in an effort to gently hone in on the exact CFFP.

2.2. Once the anterior compartment fascia is perfectly flat, record the pressure sensor measurement.

2.2. Repeat the CFFP measurement 2 more times. Record and average all 3 pressures to define the final CFFP.

3. Calculate the delta CFFP compared to the uninjured (contralateral) leg

3.1. Measure and record the CFFP in the contralateral leg in the same manner described above.

3.2. Calculate the delta CFFP by taking the difference between the CFFP of the injured leg the well leg.

REPRESENTATIVE RESULTS:

This technique has been used to measure 10 consecutive patients with injured legs without evidence of compartment syndrome and 3 patients with a clinical diagnosis of compartment syndrome necessitating decompressive fasciotomies. The average delta CFFP for patients without compartment syndrome was 10.7 ± 10.6 mbar compared to 157 ± 51.7 mbar for patients with compartment syndrome and single-tailed t-test identified the differences in pressures between the groups as being statistically significant (P<0.02).

FIGURE AND TABLE LEGENDS:

Figure 1: Ultrasound image of anterior compartments. The bright white fascia overlying the anterior compartment is demonstrated with the white arrows. Note the convex appearance of the fascia (A) with gentle pressure with the ultrasound probe versus the flattened appearance of the fascia (B) seen with applying increased pressure.

DISCUSSION:

This article demonstrates a novel, clinically applicable technique for the non-invasive measurement of intra-compartmental pressures in the leg using an ultrasound coupled with a pressure sensor. This technique only requires a hand-held ultrasound and a commercially available pressure sensing transducer. It is quick and easy to perform and is well tolerated by patients. The entire diagnostic procedure can be performed in under 5 minutes with no additional resources. The calculated delta CFFP may then be used in conjunction with the physical exam to assess the need for decompressive fasciotomy in an injury leg.

The critical step of this protocol requiring the greatest attention to detail is capturing the pressure value at the point where the fascial layer is completely flat. Subtle convexity or concavity may markedly inflate or deflate the CFFP value. Obtaining 3 serial measurements will aid in improving the accuracy of the test. Additionally, practicing the technique will improve reproducibly and reliability of this exam.

The next step, which is currently underway, is prospectively evaluating the utility of this noninvasive ultrasound method for assessing intra-compartmental pressures. We are collecting a prospective series of patient without compartment syndrome to obtain a baseline reference for CFFPs. We are also collecting prospective data on patient clinically diagnosed with compartment syndrome to identify a delta CFFP threshold for predicting the need for decompressive fasciotomy.

ACKNOWLEDGMENTS:

None.

DISCLOSURES:

The authors have nothing to disclose

REFERENCES:

- Ulmer, T. The clinical diagnosis of compartment syndrome of the lower leg: are clinical findings predictive of the disorder?. Journal of Orthopaedic Trauma. 16 (8), 572-577 (2002).

- Whitesides, T.E. et al. Tissue pressure measurements as a determinant for the need of fasciotomy. Clinical Orthopaedics and Related Research. 113, 43-51 (1975).

- McQueen, M.M., Court-Brown, C.M. Compartment monitoring in tibial fractures: the pressure threshold for decompression. The Journal of Bone and Joint Surgery. British Volume. 78 (1), 99-104 (996).

- McQueen, M.M., Christie, J., Court-Brown, C.M. Compartment pressures after intramedullary nailing of the tibia. The Journal of Bone and Joint Surgery. British Volume. 72 (3), 395-397 (1990).

- Masquelet, A-C. Acute compartment syndrome of the leg: pressure measurement and fasciotomy. Orthopaedics & Traumatology: Surgery & Research. 96 (8), 913-917 (2010).

- McQueen, M.M., Duckworth, A.D., Aitken, S.A. The estimated sensitivity and specificity of compartment pressure monitoring for acute compartment syndrome. Journal of Bone and Joint Surgery. 95 (8), 673-677 (2013).

- Whitney, A. et al. Do one-time intracompartmental pressure measurements have a high false-positive rate in diagnosing compartment syndrome?. Journal of Trauma and Acute Care Surgery. 76 (2), 479-483 (2014).

- Sellei, R. M. et al. Assessment of elevated compartment pressures by pressure-related ultrasound: a cadaveric model. European Journal of Trauma and Emergency Surgery. 41 (6), 639-645 (2015).

- Sellei, R.M. et al. Compartment elasticity measured by pressure-related ultrasound to determine patients “at risk” for compartment syndrome: an experimental in vitro study. Patient Safety in Surgery. 9 (1), 4 (2015).

- Bloch, A. et al. Compression sonography for non-invasive measurement of lower leg compartment pressure in an animal model." Injury. 49 (3), 532-537 (2018).

- Marmor, M. et al. Use of standard musculoskeletal ultrasound to determine the need for fasciotomy in an elevated muscle compartment pressure cadaver leg model. Injury (2019).