# Submission ID #:  59923

# Scriptwriter Name: Michael Linnes

# Project Page Link: http://www.jove.com/files_upload.php?src=18276563

# Title: Design, Instrumentation, and Usage Protocols for Distributed in Situ Thermal Hot Spot Monitoring in Electric Coils using FBG Sensor Multiplexing

# Authors and Affiliations:  Anees Mohammed, Siniša Djurović

School of Electrical and Electronic Engineering, University of Manchester, Manchester, UK

# Corresponding Author:

Siniša Djurović (Sinisa.Durovic@manchester.ac.uk)

# Email addresses for Co-authors:

Anees Mohammed (Anees.Mohammed@manchester.ac.uk)

Author Questionnaire:

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique? (N)

2. Does your protocol include software usage? (Y)

If yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps.

3. Which steps from the protocol section below are the most important for viewers to see? Please list 4-6 individual steps using the step numbers listed in this document. This information is important to prepare your Videographer for your shoot. (You do not need to include steps that will be screen captured. Please do not list entire sections.)

Authors, please answer this question with the steps listed here in the Protocol section below for use by the videographer. 2.3 - 2.5,  4.1 - 4.3, 5.2 - 6.2

4. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1-2 individual steps using the step numbers listed in this document. (Please do not list entire sections.)

Authors, please answer this question with the steps listed here in the Protocol section below for use by the videographer.

5. Will the filming need to take place in multiple locations? (N)

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

Authors, these headshots will be used for the JoVE Dedicated Author Webpage. Here is one example if you wish to take a look.

- REQUIRED Interview Statements:

  - Siniša Djurović: Monitoring thermal hot-spots within electric coils is critical in the power conversion area as it allows better understanding of device health, lifetime and proximity to design limits. [1]

    - INTERVIEW: Author says above statement interview style looking slightly off camera

  - Anees Mohammed: The reported technique enables in-situ distributed monitoring of thermal hotspots within the electric coil structures based on the application of multiplexed EMI immune and power passive fibre optic sensing.  [1]

    - INTERVIEW: Author says above statement interview style looking slightly off camera

OPTIONAL Interview Statements:

  - Siniša Djurović: The advanced FBG thermal sensing location and performance described in this video is unique and cannot be matched by conventional thermal sensors, such as thermocouples, or resistance based estimation techniques. [1]

    - INTERVIEW: Author says above statement interview style looking slightly off camera

  - Anees Mohammed: FBG sensors are inherently responsive to thermal and mechanical excitation and fragile, hence their application for exclusively thermal sensing within electric coils structure requires a special procedure explained in this protocol. [1]

    - INTERVIEW: Author says above statement interview style looking slightly off camera

# Section - Protocol

- Fiber Optic Thermal Sensor Design

  - First identify the sensor design and specifications based on your target coil structure and the interrogation system features. [1] The test coil shown here is a standard IEEE class H motorette, typical of electric machine coils.[2]

    - WIDE: Talent stands at bench with a few options laid out in front of them

    - Talent displays the described test coil

  - Anees Mohammed: “When designing the sensing scheme, ensure that the optical sensing fiber remains operative in the thermal and mechanical environments typical of the wound coil embedded sensing application.” [1]

    - INTERVIEW: Talent says above statement interview style

  - Using the standard bend-insensitive polyimide-coated single mode fiber ensures that the sensor is able to operate in temperatures in excess of 200 degrees Celsius and that it has the mechanical properties that allow it to be bent to conform to a desired coil geometry. [1]

    - Talent shows a section of the fiber to illustrate its general appearance and flexible nature.

  - In this application four thermal sensing points are chosen to be installed in four test coil cross sectional center locations. [1]

    - Talent points out thermal sensing points on the test coil. LAB MEDIA: 2.4.1.jpg (Author Comment: Shot has been filmed, but we think it will not clearly demonstrate the sensing locations. Therefore, a figure has been uploaded which provides clear illustration – this can either be combined with the shot or replace it as the shot on its own is not sufficiently clear.)

  - Anees Mohammed: “Individual sensing locations are identified based on relevant thermal monitoring standards for electric machines.  The distances between the individual sensing heads are based on the test coil geometry and the chosen sensing locations. [1]

    - INTERVIEW: Talent says above statement interview style while holding and pointing to the fiber

  - Next, specify individual FBG heads to be 5 millimeters in length and grated with different wavelengths spaced in a bandwidth from 1529 to 1560 nanometers to match the used commercial interrogator rating and to prevent shifted wavelengths interference. [1]

    - LABMEDIA: Figure 4 Figure 2.6.1 - Video Editor: Highlight the y-axis with mention of the bandwidth. (Author Comment: Figure 4 is for calibration and is not related to statement 2.6. A figure has been uploaded to illustrate the general pattern of reflected wavelengths by the sensor array illustrating four distinct reflected wavelength points)

  - Here, the overall fiber length is specified to 1.5 meters.  The initial 1.2 meters is packaged in Teflon and allows connection to the external interrogator device. The additional length of 0.3 meters [1] contains the four unpackaged sensing heads. [2]

    - LABMEDIA: Figure 3a- Video Editor: Highlight the 1200 mm region with the beginning of the 2nd sentence.  Then highlight the connector when “connection to the external interrogator device” is mentioned.  Finally, highlight the 300 mm section with the 3rd sentence.

    - LABMEDIA: Figure 3b - Video Editor: Highlight the 4 FBG sections

  - Shown in this video is the specified array sensor which was commercially manufactured.[1]

    - Talent shows the manufactured sensor measuring the sensing and connection areas and pointing the locations of individual FBG sensing heads.

- Interrogation System and Sensor Configuration

  - First, remove the protective cap from the FC/APC connector ferrule. [1] Then, clean the connector end-face by wiping it gently with an optical connector cleaner. [2]

    - Talent removes the cap

    - Talent cleans the connector

  - Next, ensure that the keyway is correctly aligned and plug-in the cleaned FBG probe connector to the interrogator channel connector. [1]

    - Talent carefully sets the connector orientation and plugs it in as described

  - Turn on the interrogator and run the configuration software. [1] On the instrument set-up tab, observe the reflected wavelength spectrums from the FBG array probe. Four peaks should be observed in the related channel spectrum. [2]

    - MED Over the Shoulder: Talent at turns on the interrogator and opens the software

    - SCREEN: Screen capture video as talent goes to the instrument setup tab, displays the reflected wavelengths/peaks.  Authors: Please upload this screen capture video to your article’s server page as 3_3_2_ Mohammed.mp4/mov

  - In the software, set the sampling frequency to 10 Hertz and set spectrum boundaries between FBG to prevent measurements interference.   Then, in the measurements setting, name the FBG heads as FBG1, FBG2, FBG3 and FBG4. Choose the wavelength as a type of quantity to be presented graphically at this stage.[1]

    - SCREEN: Screen capture video as talent performs the above steps in the order listed. Authors: Please upload this screen capture video to your article’s server page as 3_4_1_ Mohammed.mp4/mov

- Packaging Preparation

  - Appropriately package the sensing area where the FBG heads are imprinted in the array fiber using a PEEK capillary. This will protect the glass fiber and ensure that the sensing head is isolated from mechanical excitation and will yield an exclusively thermal excitation responsive sensor. [1]

    - Talent shows the PEEK packaging capillary.

  - Cut an adequate length of commercial PEEK tubing to the length of the target coil structure with a few extra centimeters to allow for fiber insertion and to cover the Teflon to PEEK capillary joint.  [1-TXT]

    - Talent measures and cuts the PEEK tubing TEXT: PEEK capillary tube, OD: 0.8 mm, Wall thickness: 0.1 mm

  - Next, take careful measurements of the FBG array and the PEEK capillary to accurately identify sensing locations on the outer surface of the PEEK capillary. This allows for positioning of FBG sensing heads in target locations within the motorette test coil. [1]

    - Talent slowly measures and marks the array

  - Then, prepare an appropriately sized shrink tube for later use [1]. [2]

    - Talent cuts shrink tubing

    - Talent places shrink tube over the specified joint

- Free Thermal Calibration

  - Insert the fiber sensing area into the PEEK capillary and maintain PEEK and Teflon connection using Kapton tape. [1]

    - Talent inserts the fiber and applies Kapton tape as described

  - Calibrate the packaged FBG array sensor by inserting it into the thermal chamber to extract its discrete temperature versus wavelength points. The FBG array sensing area is formed based on the coil geometry. [1]

    - Talent inserts array as described

  - Next, connect the grated optical fiber to the interrogator and launch the pre-configured interrogator software routine. [1]

    - Talent connects the fiber to the interrogator.

  - Operate the oven in a sequence of thermal steady state points. [1-TXT] Create a table from the measured reflected wavelengths of each individual FBG in the array for every constant temperature emulated in the chamber. [2]

    - Talent places setup into oven and turns it on TEXT: Room temp to 175°C, steps of 10 degrees

    - SCREEN: Screen capture video as talent creates tables as described.  Authors: Please upload this screen capture video to your article’s server page as 5_4_2_ Mohammed.mp4/mov

  - Then, use the recorded shifted wavelength versus temperature measurements to determine the optimal temperature-wavelength shift fit curves and their coefficients for each FBG. [1]

    - LABMEDIA: 5_4_2_ Mohammed.mp4/mov - Video Editor: Show a screenshot from the video in step 5.4.2 including the relevant coefficient numeric values.

  - Input the calculated coefficients in the relevant settings of the interrogator software to enable on-line temperature measurements from the FBG array. [1]

    - SCREEN: Screen capture video as talent enters the values relating to figure 4 into the software.  Authors: Please upload this screen capture video to your article’s server page as 5_6_1_ Mohammed.mp4/mov

- Test Coil Build and FBG Instrumentation

  - First, build and instrument the motorette random wound coil.  To accomplish this, set the selected, class H enameled copper wire reel in the winder device and wind half of the coil turns at a low speed.[1]

    - Talent performs the above step

  - Then, fit the prepared PEEK capillary in the center of the coil using Kapton tape. [1] Once properly positioned, wind the rest of the coil turns.[2]  Place the finished coil into the motorette frame.[3]

    - Talent tapes the capillary into the coil

    - Talent winds the rest of the coil

    - Talent places the coil into the frame

  - Next, bind the motorette coil end-windings. [1] [2]

    - Talent binds the coil end-windings

    - Talent connects the terminals

  - With the FBG array connected to the interrogator, carefully insert the sensing area fiber into the PEEK capillary until the end openings of Teflon and PEEK capillaries are in contact.[1]

    - Talent performs the above step in the order listed

  - Move the shrink tube to cover the capillaries ends and appropriately heat it until the desired fit is achieved.[1]

    - Talent positions the shrink tube and heats it

- Static and Non-Uniform Thermal Condition Tests

  - To begin the static test, connect the motorette to a DC power supply and control the DC power supply to inject the motorette with a DC current. [1]

    - Talent connects the power supply to the motorette and turns it on

  - Record measurements until the motorette coil thermal equilibrium is reached.[1]

    - SCREEN: Screen capture video as the measurements are recorded.  Authors: Please upload this screen capture video to your article’s server page as 7_2_1_ Mohammed.mp4/mov

  - Next, perform a non-uniform thermal condition test.  For this test, first wind the external coil containing 20 turns around a selected test coil section.[1]

    - Talent winds the external coil around the test coil section

  - With the external coil connected to a separate DC power supply, energize the motorette with the same DC current used in the static test.  [1] Once the thermal equilibrium is reached, begin recording thermal measurements.[2]

    - Talent sets the power as described

    - MED Over the Shoulder: Talent at computer, takes measurements.

  - Finally, energize the external coil with a DC current to provide non-uniform thermal conditions by delivering localized thermal excitation on the test coil. [1] Stop recording measurements once thermal equilibrium is reached. [2]

    - Talent energizes the coil

    - Talent disconnects device after turning off power

# Section – Results

- Results: In-situ Thermal Hot Spot Monitoring in Electrical Coils

  - During this representative static thermal test, the four internal temperature readings were taken by respective array FBG heads in their corresponding coil locations. [1] The readings are closely similar with a slight variation between the reported individual measurements of less than about 1.5 degrees Celsius [2-TXT].

    - LABMEDIA: Figure 5 without inset

    - LABMEDIA: Figure 5 including the zoomed in inset. TEXT: Ave. internal coil temp at steady-state ≈ 75.5°C

  - Once the external 20 turn coil was excited to emulate a non-uniform thermal condition within the test coil structure, a clear change was observed in the thermal measurements with redistribution of the coil internal temperature.[1]

    - LABMEDIA: Figure 6

  - The sensing point in closest proximity to the external coil, FBG4, measured the highest thermal level [1] and that furthest away sensing point, FBG2, measured the lowest. [2]

    - LABMEDIA: Figure 6- Video Editor: Highlight the FBG4 line

    - LABMEDIA: Figure 6- Video Editor: Highlight the FBG2 line

  - The observed readings clearly relate to variations in the individual sensing head distribution in the examined test coil geometry. This demonstrates the functional capability of the coil embedded array sensor for monitoring and identifying thermal hotspot distributions in random wound coils.[1]

    - LABMEDIA: Figure 6

# Section - Conclusion

- Conclusion Interview Statements:

# What is most important thing to remember when attempting this procedure? Please indicate the steps (e.g., 2.4., 2.5.) in the Protocol section this advice correlates to.

  - Siniša Djurović: In this video, we have demonstrated how a single optical fiber, using FBG technology, can provide measurements of distributed thermal hot-spots within a coil structure.  This would be extremely challenging using conventional sensors.[1]

    - INTERVIEW: Author says above statement interview style looking slightly off camera - Video Editor: Show Figure 6 as an inset during the second half of the first sentence

# Following this procedure, what other methods can be performed? What questions would these additional methods answer?

  - Anees Mohammed: To ensure accurate measurements, take special care with the packaging, installation, and calibration procedures.  These are needed to mitigate thermo-mechanical FBG cross-sensitivity, protect the sensing fiber, and allow for reliable thermal readings to be taken.[1]

    - INTERVIEW: Author says above statement interview style looking slightly off camera

# After its development, did this technique pave the way for researchers to explore new questions within a specific scientific field? If so, how?

  - Siniša Djurović: The reported technique provides new opportunities for the development of dedicated in-situ thermal monitoring applications in energy conversion devices.  These devices are the core elements of modern industrial applications, where conventional sensors are challenged. [1]

    - INTERVIEW: Author says above statement interview style looking slightly off camera