7-7-2019

Dear Editor:

On behalf of my co-authors, we appreciate editor and reviewers very much for the constructive comments and suggestions on our manuscript entitled “Three-dimensional printing guide template assisted percutaneous vertebroplasty”. We have read comments from you and reviewers carefully and made corresponding changes. We have tried our best to revise our manuscript according to the comments as followed.

We would like to express our great appreciation to you and reviewers for comments on our paper. Look forward to hearing from you soon.

Sincerely,

Qi Fei, MD

Department of Orthopedics, Beijing Friendship Hospital,

Capital Medical University,

Beijing, China, 100050

86-13641225060

spinefei@126.com

Response to editor: Thank you very much for giving us constructive comments.

- Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues. The JoVE editor will not copy-edit your manuscript and any errors in the submitted revision may be present in the published version.
There are many missing articles (a, an, the, etc.) that make some of the text awkward. Furthermore, some phrase usage is incorrect especially with the numbered lists.

Answer: We have made according changes in revised manuscript.

- Please remove the embedded figure(s) from the manuscript. All figures should be uploaded separately to your Editorial Manager account. Each figure must be accompanied by a title and a description after the Representative Results of the manuscript text.

Answer: We have made according changes in revised manuscript.

- Please ensure that the references appear as the following: [Lastname, F.I., LastName, F.I., LastName, F.I. Article Title. Source. Volume (Issue), FirstPage – LastPage (YEAR).] For more than 6 authors, list only the first author then et al.

Answer: We have made according changes in revised manuscript.

- Please revise the table of the essential supplies, reagents, and equipment. The table should include the name, company, and catalog number of all relevant materials in separate columns in an xls/xlsx file. Please sort the Materials Table alphabetically by the name of the material.

Answer: We have made according changes in revised manuscript.

- Please rephrase the title to be more clear. It is a bit wordy.

Answer: We have made according changes in revised manuscript.

- Please include a Summary that clearly describes the protocol and its applications in complete sentences between 10-50 words: “Here, we present a protocol to …”

Answer: We have made according changes in revised manuscript.

- For in-text formatting, corresponding reference numbers should appear as numbered superscripts after the appropriate statement(s).

Answer: We have made according changes in revised manuscript.

- Please ensure that all text in the protocol section is written in the imperative tense as if telling someone how to do the technique (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “Note.” However, notes should be concise and used sparingly. Please include all safety procedures and use of hoods, etc.

Answer: We have made according changes in revised manuscript.

- The Protocol should contain only action items that direct the reader to do something. Please move the discussion about the protocol to the Discussion.

Answer: We have made according changes in revised manuscript.

- JoVE cannot publish manuscripts containing commercial language. This includes trademark symbols (™), registered symbols (®), and company names before an instrument or reagent. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials and Reagents.

Answer: We have made according changes in revised manuscript.

- Please add more details to your protocol steps. Please ensure you answer the “how” question, i.e., how is the step performed? Alternatively, add references to published material specifying how to perform the protocol action.

Answer: We have made according changes in revised manuscript.

- 1.3: Please provide citations for the VAS and the ODI.

Answer: We have made according changes in revised manuscript.

- 2.1: Place where on the back skin? In what pattern?

Answer: We have made according changes in revised manuscript.

- 2.2: Where should the gradienter be placed?

Answer: The gradienter is on patient`s back just inferior to the fixed markers, to record patient`s body position.

- Please do not use contractions (won’t).

Answer: We have made according changes in revised manuscript.

- How many CT images are taken? Please provide all imaging parameters.

Answer: CT images (scanning layer thickness 1 mm, layer spacing 1 mm, and totally around 400 slices if done is thin slice scanning or 90 slices if done in conventional scanning.)

- 5: There are many details missing from the surgery. We need these details if any of these steps are to be filmed. Please specify all experimental parameters and instruments used.

Answer: We have made according changes in revised manuscript.

- How much is injected?

Answer：We injected 2ML bone cement via each trajectory and totally 4ML bone cement were injected to one vertebrae.

- How much of what anesthesia is used?

Answer: We inject the mixture of 1% lidocaine and 1% ropavicaine 5ml at the each puncture points to do local anesthesia.

- Please highlight 2.75 pages or less of the Protocol (including headings and spacing) that identifies the essential steps of the protocol for the video, i.e., the steps that should be visualized to tell the most cohesive story of the Protocol. Remember that non-highlighted Protocol steps will remain in the manuscript, and therefore will still be available to the reader.

Answer: We have made according changes in revised manuscript.

- Please ensure that the highlighted steps form a cohesive narrative with a logical flow from one highlighted step to the next. Please highlight complete sentences (not parts of sentences). Please ensure that the highlighted part of the step includes at least one action that is written in imperative tense.

Answer: We have made according changes in revised manuscript.

- Please include at least one paragraph of text to explain the Representative Results in the context of the technique you have described, e.g., how do these results show the technique, suggestions about how to analyze the outcome, etc. The paragraph text should refer to all of the figures. Data from both successful and sub-optimal experiments can be included.

Answer: We have made according changes in revised manuscript.

- Please discuss all figures in the Representative Results. However for figures showing the experimental set-up, please reference them in the Protocol.

Answer: We have made according changes in revised manuscript.

- Please ensure that all panels of the figures are labeled and please be consistent with the labeling. Please ensure that the panel labeling is large and high contrast with the image. Some of the red panel labels are too small and hard to see.

Answer: We have made according changes in revised manuscript.

- Can the STL file for the 3D printer be provided?

Answer：we can upload the MCS file for you, the STL file is converted by the company from the MCS file. However when we try to upload this file, the website shows that this file is over 60 MBs，so it cannot be uploaded. If you want it, maybe we can e-mail it to you.

Thank you again for your constructive comments and suggestions.

Response to Reviewer 1: Thank you very much for giving us constructive comments.

- In the introduction, the authors are listing several issues associated with the "traditional" surgical procedure. These issues should be clearly documented, including their incidence and associated clinical implication

Answer： We have mentioned the incidence and associated clinical implication of complications of traditional PVP in the introduction part.

- There is not information on the logistics aspects of the full process: how long does it take to obtain the 3D printed guides, and what is the risk for the skin markers to move during this period

Answer： We need about half an hour to reconstruct the 3-D model from the CT images for the 3-D printing, the 3-D printing company need about 6 hours to print 2 guide templates out and send them to our hospital; if the skin markers move during this period, it will reduce the accuracy of the localization of the operation and require more fluoroscope times to determine the suitable puncture points, even make the template failed to apply in the operation. However, this situation never happened in our 15 operations. We recommend draw the outline of the skin markers after patients lied down in prone position, then remove the skin markers, which could minimize this kind of concern.

- The most important concern I have relates to the lack of clinical data. While the authors state in their abstract that they have seen "some significant advantages", this should be clearly reported in the manuscript. More specifically, the authors should demonstrate that this new process bring value in terms of puncture accuracy, puncture-related complications, shorten clinical procedure, minimize radiation .... How many patients were operated according to this protocol. What are the outcomes of these surgeries. What are the potential issues encounter while implementing the protocol.

Answer： Up to now, we recruited 30 patients suffering one-level OVCF and treated in our hospital, we randomly divided them into group A (underwent “three-dimensional printing individual guide template assisted Percutaneous Vertebroplasty”) and group B (underwent traditional PVP). Fluoroscopy shot times for skin puncture points (1.71±0.83 in the group A vs 5.31±1.96 in the group B, P<0.01), total radiation doses (4.67±2.97)mGy in the group A vs（8.22±5.24）mGy in the group B, P<0.05], total fluoroscopy times (16.07±2.67 in the group A vs 26.31±3.12 in the group B, P<0.01) and operation time [(20.07±3.65)min in the group A vs (26.19±3.12)min in the group B, P<0.01] were presented statistically different in the two groups. The incidence of cement leakage occurred in the group A (2/15, 13.3%) was less than that occurred in the group B (6/15, 40.0%). Our results turn out that, compared with the traditional PVP, “three-dimensional printing individual guide template assisted Percutaneous Vertebroplasty” has some advantages: less fluoroscopy shot times during operation, less total fluoroscopy dose and shorter operation time; less incidence of cement leakage. But in this article, we are not intended to discuss about those data, we just want to introduce a new surgical method for PVP.

Thank you again for your comments and look forward to your further suggestions.

Response to Reviewer 2: Thank you very much for giving us constructive comments.

- Introduction, page 3/4, line 40 ff and M+M page 4/5, line 76-94: The manuscript should be rather focused on the technique - the diagnosis of osteoporotic fracture and in indications for vertebroplasty are not concerned by the data reported and need not to be discussed/explained in the M+M section.

Answer： We have already deleted the diagnosis of osteoporotic fracture and the indications for vertebroplasty in our revised manuscript.

- Is this a cadaver-study? How many cadavers were used, how many vertebrae/trajectories? How were the cadavers obtained (do you have consent from the corresponding ethical committee)?

Answer： This is a randomized, nonblinded, controlled clinical study, not a cadaver-study. Our project is still underway, up to now we have recruited 15 one-level OVCF patients undergoing PVP assisted by 3-D guide template (Group A) and 15 one-level OVCF patients treated by traditional PVP (Group B), there are totally 30 patients and 30 vertebrae in our research.

- What kind of markers were used (it seems that the patient has to carry them a rather long time - during the printing procedure etc.)

Answer： The markers are tiny radiopaque spacers made of stainless steel. The patients may need to carried them 1-2 days from they underwent the CT scan to the beginning of the operation. Usually the first day we take the patient to undergo the CT scan, reconstruct the 3-D model in MIMICS software and send it to the 3-D printing company to print the templates, the morning of second day, we do the operation for patient.

- What kind of CT was used? Please specify the CT-settings.

Answer： The CT machine is a 64-slice spiral CT manufactured by the General Electric Medical Group (GE, USA), tube voltage 120 KV, tube current 200 mA, scanning layer thickness 1 mm, layer spacing 1 mm.

- It is necessary to describe the applied software (what does it do, is it commercial software or a self-programmed protocol…) before describing which keypads should be selected.
What kind of printing process (company, material …) was applied? Are the implants sterile, or can they be sterilized (by which means)?

Answer： The software is a commercial software, which is called Materialise Interactive Medical Image Control System (MIMICS) v17, it is used to modeling target vertebrae DICOM imaging into 3-D data to form the STL file. The 3-D company is Beijing Huashengputian technology co., ltd., a 3-D printing company in Beijing, China. The company use Design X software to modeling the template. The guide template is made of polylactic acid, it can be sterilized under low temperature steam disinfection.

- How long did it take to obtain the printed models? Is the patient thought to keep the markers all these times on the back?

Answer： We need about half an hour to reconstruct the 3-D model from the CT images, the 3-D printing company needs about 6 hours to print 2 guide templates out and send them to our hospital. The patients do need to keep the markers on the back until the procedures begin. As the markers are tiny radiopaque spacers made of stainless steel, they will not feel discomfort seriously, which will hardly affect the quality of life of patients.

- How much and what kind of cement was infused?

Answer： We injected 2ML bone cement via each trajectory and totally 4ML bone cement were injected to one vertebra. The kind of cement is polimethylomethylacrylate cement (PMMA).

- What were the results? - How frequent did the authors observe cement leakage or other problems.

Answer： We observe the cement leakage or other problems during the whole operation procedure, and one day after the operation. The incidence of cement leakage occurred in the group A (2/15, 13.3%) was less than that occurred in the group B (6/15, 40.0%).

- The discussion has than to be focused on the data provided - the misplacement and leakage values in literature should be compared to those found in the pilot-study, as yet it is not clear that the proposed technique does improve those rates.

Answer： Up to now, we recruited 30 patients suffering one-level OVCF and treated in our hospital, we randomly divided them into group A (underwent “three-dimensional printing individual guide template assisted Percutaneous Vertebroplasty”) and group B (underwent traditional PVP). Fluoroscopy shot times for skin puncture points (1.71±0.83 in the group A vs 5.31±1.96 in the group B, P<0.01), total radiation doses (4.67±2.97)mGy in the group A vs（8.22±5.24）mGy in the group B, P<0.05], total fluoroscopy times (16.07±2.67 in the group A vs 26.31±3.12 in the group B, P<0.01) and operation time [(20.07±3.65)min in the group A vs (26.19±3.12)min in the group B, P<0.01] were presented statistically different in the two groups. The incidence of cement leakage occurred in the group A (2/15, 13.3%) was less than that occurred in the group B (6/15, 40.0%). Our results turn out that, compared with the traditional PVP, “three-dimensional printing individual guide template assisted Percutaneous Vertebroplasty” has some advantages: less fluoroscopy shot times during operation, less total fluoroscopy dose and shorter operation time; less incidence of cement leakage. But in this article, we are not intend to discuss about those data, we just want to introduce a new surgical method for PVP.

Thank you again for your comments and look forward to your further suggestions.

Response to Reviewer 3: Thank you for your thought-provoking questions and giving us constructive comments.

Question:  The technique described in this manuscript includes several steps of diagnostics, digital planning, 3D-printing, sterilization and application of a template onto the patient's skin for optimized entry points for percutaneous vertebroplasty. Vertebroplasty is quite a simple and easy to learn surgical method that must be performed under fluoroscopy, because bone cement is injected into the patient. Does the complex planning, fabrication and application of templates, which is time consuming, really justify the use of this patient-specific technique for a simple surgical procedure like vertebroplasty? A one-level vertebroplasty with a bipedicular approach can be done in no more than 15 minutes or even less. What is the time needed for the whole planning, printing, sterilization process of the template? Template-guided techniques are normally designed to apply a template on a bone for firm contact to optimize screw placement accuracy or let's say osteotomy accuracy. In the technique described here, the template is placed on soft tissue, namely the skin of the patient with underlying soft tissue (fat, muscles), which should be critically discussed as this might negatively influences the trajectory of the trocars during surgery. Even more, the template may show signs of deformation after sterilization, which further worsens its accuracy. Another concern is that you still depend on flucoroscopy to visualize the final position of the trocars and the distribution of the bone cement injected. Under these circumstances, given the presurgical effort this technique demands, and the supposed marginal reduction of radiation exposure, it should be critically discussed whether this technique is in any form beneficial in daily practice, which I personally doubt.

Answer： First, we agree with you that applying 3-D guide template into one-level OVCF patient may not save a huge amount of time and reduce radiation exposure massively, however, in our randomized, nonblinded, controlled clinical study, those data do present the statistical difference: up to now, we recruited 30 patients who suffered one segment OVCF and treated in our hospital, we randomly divided them into group A (underwent “three-dimensional printing individual guide template assisted Percutaneous Vertebroplasty”) and group B (underwent traditional PVP). Fluoroscopy shot times for skin puncture points (1.71±0.83 in the group A vs 5.31±1.96 in the group B, P<0.01), total radiation doses (4.67±2.97)mGy in the group A vs（8.22±5.24）mGy in the group B, P<0.05], total fluoroscopy times (16.07±2.67 in the group A vs 26.31±3.12 in the group B, P<0.01) and operation time [(20.07±3.65)min in the group A vs (26.19±3.12)min in the group B, P<0.01] were presented statistically different in the two groups. The incidence of cement leakage occurred in the group A (2/15, 13.3%) was less than that occurred in the group B (6/15, 40.0%).

Second, currently, our clinical task research is focus on applying 3-D guide template into one-level OVCF patient, but in the future we will apply the guiding template into complicated OVCF patients with severe osteoporosis, severe kyphosis, scoliosis or multi-segment fractured vertebra. Even the experienced surgeons take many fluoroscope times and consume a pretty long time to deal with those cases throughout the operations. In view of these difficult cases, “Three-dimensional printing guide template assisted percutaneous vertebroplasty” may offer a safer and more precise puncture approach than traditional procedure, also do better in saving operation time and reducing radiation exposure.

Third, preoperative design in MIMICS provides an intuitive and visual puncture approach of percutaneous vertebroplasty, which may help fresh surgeons learn about the anatomy of vertebra and simulate the optimal puncture trajectory, moreover it can also optimize the learning curve of this kind of procedure for young surgeons.

Additionally, you mentioned the template may show signs of deformation after sterilization, we also agree with you that after low temperature steam disinfection, the templates may be slightly deformable, however in all of our cases, this kind of deformation hardly affects the puncture accuracy. We will explore a much more suitable material to print 3-D templates.

Thank you again for your comments and look forward to your further suggestions.