I. Editorial comments:

Changes to be made by the author(s):

1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues. The JoVE editor will not copy-edit your manuscript and any errors in the submitted revision may be present in the published version.

Thank you for this amendment. All spelling or grammar issues were corrected.

2. Please revise lines 41-53 and 264-268 to avoid previously published text.

Thank you for this amendment. We revised the lines 41-53 and 264-268 to avoid previously published text. Therefore, we changed the manuscript as follows:

"Herman Snellen, a Dutch eye surgeon, used this cryolite glass to produce a lightweight hollow prosthetic eye in 1880. This lightweight prosthetic eye, the Snellen ‘reform eye’ increased the volume of prosthetic eyes resulting in better fitting into larger sockets following the introduction of enucleation procedures made possible by the development of anesthesia and asepsis.  20 years later, cryolite glass had become the most commonly used material for prosthetic eyes. Germany developed into the manufacturing centre of prosthetic eyes globally. At the start of the second world war, German cryolite glass eyes became unavailable outside the German-speaking area. Therefore (poly)methyl methacrylate (PMMA) became a substitute material for prosthetic eyes, and today PMMA is the most commonly used material for prosthetic eyes globally. Notwithstanding, in German speaking countries over 90% of the ocularists still manufacture the customized prostheses using the cryolite glass from Thuringia."

"Another aspect is the durability of both materials (PMMA vs. cryolite glass). In the first glance, the potential breakage of cryolite glass prosthetic eyes seems to be problem for anophthalmic patients. However, the mean rate of breakage is very low and amounts only one prosthesis per 26.63 wearing years.  The damages (94%) occur during removing or cleaning the glass eye  . Therefore, removing and cleaning of the prosthetic eye should be done over a filled sink.  In case of damage or loss, almost all cryolite glass prosthetic eye wearers have at least one suitable replacement prosthesis."

3. Authors and affiliations: Please provide an email address for each author.

Thank you for this important information. We provided email addresses for each author:

"Email addresses:

Marc Trester: marctrester2018@gmail.com

Joel M. Mor: joel.mor@uk-koeln.de

Niklas Loreck: niklas.loreck@uk-koeln.de

Konrad R. Koch: konrad.koch@uk-koeln.de

Ludwig M. Heindl: ludwig.heindl@uk-koeln.de"

4. Abstract: Please expand to include an overview of the method and a summary of its advantages, limitations, and applications. Please do not include references here.

Thank you for this important information. We expanded the abstract as follows:

"In Germany, Austria, and Switzerland, over 90% of the ocularists still manufacture customized prostheses using cryolite glass from Thuringia. The present manuscript demonstrates this technique - long forgotten in other countries – in detail. This manuscript shows some major advantages of manufacturing prosthetic eyes using cryolite glass in comparison to PMMA, such a lighter weight of the prosthesis, higher levels of patient satisfaction, and only one appointment for the customized manufacturing. Potential risk of breakage seems not to be a decisive disadvantage for glass prosthetic eye wearers. However, in some patients manufacturing a good fitting prosthetic eye is not possible or reasonable due to anophthalmic socket complications such as post enucleation socket syndrome, scarred fornices or an orbital implant exposure. This article gives ophthalmologists a better insight into the ocularistic care in order to improve the essential interprofessional collaboration between ocularists and ophthalmologists."

5. Introduction: Please expand to include the advantages of the presented method over alternative techniques with applicable references to previous studies, description of the context of the technique in the wider body of literature and information that can help readers to determine if the method is appropriate for their application.

Thank you for this important information. We expanded the introduction as follows:

"The purpose of the present manuscript is to comprehensively demonstrate the technique of manufacturing a customized cryolite glass prosthetic eye that is long forgotten outside the German-speaking area (Figure 1). In addition, this manuscript focuses on major advantages of this technique described in the protocol resulting in a very smooth surface of the prosthesis due to fire polishing, the light weight of the prosthesis on the basis of the hollow design, high levels of patient satisfaction and the need of only one appointment for manufacturing of the customized prosthesis. This article also gives ophthalmologists better insights into the ocularistic care in order to improve the essential interprofessional collaboration.

In 1832, the glassblower Ludwig Uri Müller from Thuringia, Germany, developed the cryolite glass prosthetic eye based on the class-leading models made in France. Benefits of cryolite glass included a better look, better tolerability, easier processing, and longer durability than previous glass eyes. Herman Snellen, a Dutch eye surgeon, used this cryolite glass to produce a lightweight hollow prosthetic eye in 1880. This lightweight prosthetic eye, the Snellen ‘reform eye’ increased the volume of prosthetic eyes resulting in better fitting into larger sockets following the introduction of enucleation procedures made possible by the development of anesthesia and asepsis.  20 years later, cryolite glass had become the most commonly used material for prosthetic eyes. Germany developed into the manufacturing centre of prosthetic eyes globally. At the start of the second world war, German cryolite glass eyes became unavailable outside the German-speaking area. Therefore, (poly)methyl methacrylate (PMMA) became a substitute material for prosthetic eyes, and today PMMA is the most commonly used material for prosthetic eyes globally. Notwithstanding, in German speaking countries over 90% of the ocularists still manufacture the customized prostheses using the cryolite glass from Thuringia. Each customized cryolite glass prosthetic eye is produced in two major steps: The first step is to produce a "half-done" cryolite glass eye what conforms nearly to a white sphere with an iris and a pupil (Figure 2). The second and decisive step is to customize the "half-done" cryolite glass prosthetic eye for the respective patient. To that end a "half-done" cryolite glass eye is selected from thousands of them (Figure 3) on the basis of the best matching iris color to the fellow eye of the patient.

The following protocol presents customizing a selected "half-done" cryolite glass eye for the specific patient. This step lasts about 25-35 minutes. All procedures performed in the following protocol involving human participants were in accordance with the ethical standards of the institutional research committee of the University of Cologne and with the 1964 Helsinki declaration and its later amendments or comparable ethical standards."

6. Please include an ethics. statement before the numbered protocol steps, indicating that the protocol follows the guidelines of your institution’s human research ethics committee.

Thank you for this important information. We included an ethics statement:

"All procedures performed in this study involving human participants were in accordance with the ethical standards of the institutional research committee of the University of Cologne and with the 1964 Helsinki declaration and its later amendments or comparable ethical standards."

7. Lines 72-79: Please move the introductory paragraph of the protocol to the Introduction, Results, or Discussion (as appropriate).

Thank you for this important information. We moved the named paragraph to the introduction. Please see point 5.

8. 1.03: What is used to remove the current prosthetic eye?

Thank you for this important amendment. We changed the manuscript as follows:

"1.03. Remove the current prosthetic eye with help of a contact lens suction cup for hard contact lenses."

9. Representative Results: Please include at least one figure or table to show the effectiveness of your technique backed up with data.

Thank you for this important amendment. We included a figure to show the effectiveness of our technique. Please see the pictures.

10. Please reference Figure 8 in the manuscript.

Thank you for this important amendment. We referenced figure 8 in the manuscript.

"Heat both slowly to 600 degrees with a Bunsen burner (Fig. 8) while continuously rotating it and melt the skewer at the open end of the "half-done" cryolite glass eye. Open the forceps and lay it down."

11. JoVE articles are focused on the methods and the protocol, thus the discussion should be similarly focused. Please revise the Discussion to explicitly cover the following in detail in 3-6 paragraphs with citations:

a) Critical steps within the protocol

b) Any modifications and troubleshooting of the technique

c) Any limitations of the technique

d) The significance with respect to existing methods

e) Any future applications of the technique

Thank you for this important amendments. We revised the discussion as follows:

"The process of customizing the half-done cryolite glass eye includes some critical steps. The first critical steps within the protocol include examination of the fitting of the current prosthesis (1.02) and examination of the anophthalmic socket without prosthesis (1.04). The ocularist has to check the current fitting of the prosthesis as well as the anophthalmic socket in detail, because on the basis of these checks the shape of the new customized prosthesis will be possibly changed compared to the old prosthesis. In addition, the ocularist has to check whether there are other points, such as extrusions of orbital implants requiring surgical interventions prior to adequate ocularistic care. Another critical step is the slow heating including a continuous rotation of the half-done cryolite glass eye. If the glass eye is not rotated continuously and evenly, the glass eye gets warped uncontrolled when heated. Heating or cooling the glass eye too quickly will result in a breakage of the glass eye. Furthermore, the temperature during shaping the cryolite glass eye has to be in the correct range (nearly 600 degrees Celsius) for an optimal result. All these steps are just a matter of practice and also a question of experience resulting in an over 6-years training period for ocularists."

"The technique described in that protocol is used for the manufacturing a hollow, double walled cryolite glass eye, a so-called 'reform eye'. These reform eyes are especially used for patients after enucleation or evisceration when they need more volume replenishment through the prosthesis. When volume replenishment is not necessary for example in phthisic eyes or in patients with microphthalmos, the technique can be modified and the cryolite glass prosthetic eye can be produced very thin and single-walled. Therefore, the backwall is melted of during the customizing process.

In the case the customized prosthesis glass eye does initially not fit, it can be heated once more, and the shape can then be modified by the ocularist. However, reheating of the glass prosthesis eye is only possible in the first days after production. After a wearing time of some days the first hydrolytic material changes and material stress occur and a reheating results mostly in a breakage of the prosthesis. Of course, in some rare cases the prosthetic eye has to be reproduced from the beginning due to a fitting error that can`t be corrected any more, for example when the prosthesis is too small, or the material is too thin.

Essentially using the technique described here allows to produce cryolite glass prosthetic eyes in nearly all shapes and sizes. The limitations of our technique depend on the anophthalmic socket. In patients with a pronounced PESS or scarring of  the fornices, it is not possible or useful to customize a prosthetic eye. These patients benefit more from an initial surgical reconstruction of the anophthalmic socket with following ocularistic care. "

"In recent decades some improvements of this technique have prevailed, especially with focus on the working conditions such as air conditioners and ergonomic workstation designs and, in a few cases, newer devices such as novel Bunsen burners. However, in the last years some potential further developments have been described in the literature. Especially, new surface coatings with antibacterial or with more hydrophilic properties are one of the main subjects of the current research regarding prosthetic eyes. However, these new coatings will probably be used especially for PMMA prosthetic eyes, because the coating procedures mostly require very high temperatures or high pressure resulting very likely in breakage of the glass prostheses. Another innovative approach to improve the prosthetic eyes is to fabricate a hollow ocular prosthesis with a functional lubricant reservoir to increase the artificial eye comfort. This technique is not established until today, and in addition further studies are needed to define the clinical outcome. A big issue of this suggested improvement could also be the growth of bacteria within the hollow prosthesis. Also with regard to this point, further studies are a high priority."

12. References: Please do not abbreviate journal titles.

According to the journal’s guidelines, we updated all references without abbreviations.

13. Table of Materials: Please also provide company and catalog number for each item in separate columns and sort the items in alphabetical order according to the name of material/equipment.

Thank you for this amendment. However, it is not possible to list the catalogue numbers or the companies. All parts of the equipment and also all parts for producing are customized or partially self constructed or handmade many years ago by firms that do not exist anymore. We sorted the equipment and material alphabetical.

Reviewers' comments:

Reviewer #1:

1. The customisation process is listed in summary at points 1.01 through to 1.16.
I was expecting that the protocol which is to be captured on video would be highlighted in yellow but as there is no highlighting, I am presuming that the protocol that is to be presented in the video is that listed after the "PROTOCOL" heading in the manuscript.
Thank you for this important information. The reviewer is right that the protocol that is to be presented in the video is that listed after the "PROTOCOL" heading in the manuscript. We highlighted the text in yellow.

2. The discussion section that follows the protocol is NOT initially related to the protocol.
Instead the discussion briefly describes the process by which a patient is handled following surgical enucleation of an eye and the surgical insertion of an orbital implant.
The discussion describes how a conformer is inserted into the inter-palpebral space to prevent scarring while the conjunctiva heals and to maintain the conjunctival fornices following during the healing process. At this point the authors describe how a progression of glass prosthetic eyes need to be fitted at intervals while swelling recedes.

We fully agree with the reviewer and revised the discussion with more focus on the protocol as follows:

"The process of customizing the half-done cryolite glass eye includes some critical steps. The first critical steps within the protocol include examination of the fitting of the current prosthesis (1.02) and examination of the anophthalmic socket without prosthesis (1.04). The ocularist has to check the current fitting of the prosthesis as well as the anophthalmic socket in detail, because on the basis of these checks the shape of the new customized prosthesis will be possibly changed compared to the old prosthesis. In addition, the ocularist has to check whether there are other points, such as extrusions of orbital implants requiring surgical intervention prior to adequate ocularistic care. Another critical step is the slow heating including a continuous rotation of the half-done cryolite glass eye. If the glass eye is not rotated continuously and evenly, the glass eye gets warped uncontrolled when heated. Heating or cooling the glass eye too quickly will result in a breakage of the glass eye. Furthermore, the temperature during shaping the cryolite glass eye has to be in the correct range (nearly 600 degrees Celsius) for an optimal result. All these steps are just a matter of practice and also a question of experience resulting in a over 6-years training period for ocularists.

The technique described in that protocol is used for the manufacturing a hollow, double walled cryolite glass eye, a so-called 'reform eye'. These reform eyes are especially used for patients after enucleation or evisceration when they need more volume replenishment through the prosthesis. When volume replenishment is not necessary for example in phthisic eyes or in patients with microphthalmos the technique can be modified and the cryolite glass prosthetic eye can be produced very thin and single-walled. Therefore the back wall is melted of during the customizing process.

In the case the customized prosthesis glass eye does initially not fit, it can be heated once more and the shape can then be modified by the ocularist. However, reheating of the glass prosthesis eye is only possible in the first days after production. After a wearing time of some days the first hydrolytic material changes and material stress occur and a reheating results mostly in a breakage of the prosthesis. Of course, in some rare cases the prosthetic eye has to be reproduced from the beginning due to a fitting error that can`t be corrected any more, for example when the prosthesis is too small, or the material is too thin.

Essentially using the technique described here allows to produce cryolite glass prosthetic eyes in nearly all shapes and sizes. The limitations of our technique depend on the anophthalmic socket. In patients with a pronounced post enucleation socket syndrome (PESS) or scarring of the fornices, it is not possible or useful to customize a prosthetic eye. These patients benefit more from an initial surgical reconstruction of the anophthalmic socket with following ocularistic care. "

"In recent decades some improvements of this technique have prevailed especially with focus on the working conditions such as air conditioners and ergonomic workstation designs and, in a few cases, newer devices such as novel Bunsen burners. However, in the last years some potential further developments have been described in the literature. Especially, new surface coatings with antibacterial or with more hydrophilic properties are one of the main subjects of the current research regarding prosthetic eyes. However, these new coatings will probably be used especially for PMMA prosthetic eyes because the coating procedures mostly require very high temperatures or high pressure resulting very likely in breakage of the glass prostheses. Another innovative approach to improve the prosthetic eyes is to fabricate a hollow ocular prosthesis with a functional lubricant reservoir to increase the artificial eye comfort. This technique is not established until today, and in addition further studies are needed to define the clinical outcome. A big issue of this suggested improvement could be the growth of bacteria within the hollow prosthesis. Also with regard to this point further studies are a high priority."

3. At 6 months post enucleation a final glass prosthesis is fitted and this is expected to last for 12 months before the level of surface etching requires a new prosthesis. This reviewer is aware that some ocularists extend the time between renewal of glass prostheses to well beyond 12 months.

Our newest data (Rokohl et al. - Cryolite Glass Prosthetic Eyes – The Response of the Anophthalmic Socket) show that most of the patients have a big inflammatory response of the anophthalmic socket. This inflammatory response seems to be stronger in patients with longer wearing time (of the same prosthesis). Most patients seem to benefit from a newly made prosthesis every 9 months. However, the reviewer is also right: in individual cases it is of course possible to extend the time between renewal of the glass prostheses to well beyond 12 months. It is very difficult to state general recommendations for all patients, but the German recommendation is in these days to renew the prosthesis at least all 12 months, also for preventing other inflammation related complications. We changed the manuscript as follows:

"Afterwards, according to the current recommendations, most adult patients in Germany get a new customized prosthesis at least once a year, primary due to hydrolytic surface changes resulting in an inflammation of the anophthalmic socket."

4. The discussion then turns to prostheses made from PMMA and the authors begin their comparison between prostheses made from each material. The authors state that PMMA prostheses are not made until 6 weeks after enucleation and that the PMMA prosthesis is then worn for the next 5 to 6 years without improvement in fit and with only annual re-polishing. The inference is that PMMA prostheses are therefore at a disadvantage. This reviewer is aware that there are ocularists who do fit PMMA prostheses earlier than 6 weeks and also that there are ocularists who willingly modify the PMMA prosthesis (e.g. by adding material or by removing material - which is not possible with glass) to correct any fitment issues.

We fully agree with the reviewer and we revised the discussion as  follows:

"PMMA prosthetic eyes have to be repolished every year due to surface changes but there is in some cases no improvement of the fitting or the design over the years."

5. The discussion continues with a comparison of the level of smoothness and the hydrophilicity of the cryolite glass and the PMMA prosthesis surface. The authors are correct that the glass surface has advantages (but this may be limited to when the prosthesis is new). They are correct when they state that more research is needed in this area.

Thank you for this information. Of course, newly made prostheses have smoother surfaces. This applies to both PMMA and glass prostheses. However, Clodius et al. (1981-Artificial Eyes: Surface Changes Following Use, as Observed by the Scanning Electron Microscope)

showed that the smoother surface of glass prostheses compared to PMMA is not limited to the new unworn prostheses.

6. At lines 274 to 276 the authors make the inference that the heavier PMMA prosthesis will irritate the conjunctiva, increase mucous discharge, increase discomfort, redden the eye lids, cause overburdening and sagging of the lower eyelid which is part of PESS. The authors cite references 5, 6, 8, 22, and 23. In my reading of these articles I did not find direct evidence that the weight of the prosthesis was a significant cause of PESS. In my view the statement that the weight of the prosthesis was a cause of lower lid sagging would be best worded as simply the opinion of the authors without direct evidence.

Thank you for this important amendment. The reviewer is right. We adjusted the manuscript as follows:

"According to the experience of the authors these heavier prostheses might cause sagging of the lower eye lid in the course of the post enucleation socket syndrome."

7. The authors make, as their final point of comparison between glass and PMMA prosthetic eyes, the time involved in the fitting process. They are correct that the number of in-office visits and time required to manufacture PMMA prostheses is much longer.
The authors do not comment on the differences in patient care requirements (how frequent is the removal and cleaning for each) for cryolite glass prosthetic eyes and PMMA prosthetic eyes. This is a significant oversight and should be included.

Thank you for this important amendment. We included patient care requirements and changed the manuscript as follows:

"So far the recommended care regime for glass prosthetic eyes included removal and cleaning on a daily basis while for PMMA prosthetic eyes removal and cleaning at an interval between monthly and six-monthly was recommended. However, the results of newest studies suggest that cryolite glass prosthetic eyes should not be removed daily for cleaning and that further research should be undertaken to develop a cleaning protocol. This protocol would recommend a minimum period of wear between cleaning sessions and should consider individual conditions, such as climatic, psychological, and environmental factors."

8. ABSTRACT

The abstract presents the aims of the manuscript succinctly

SUMMARY
The summary statement does NOT mention that the topic of this publication is the in-office protocol for customising a cryolite glass prosthetic eye for a patient.

Thank you for this important amendment. We changed the manuscript as follows:

"This manuscript shows each step of customizing a cryolite glass prosthetic eye including some major advantages of the use of cryolite glass for manufacturing an eye prosthesis compared to PMMA. In addition, this manuscript gives ophthalmologists a better insight into the ocularistic care what could improve the interprofessional collaboration."

9. Major Concerns:

It is not clear from the manuscript what the protocol is that is to be presented in video format. There is no yellow highlighting.

We fully agree with the reviewer. The protocol that should be presented in the video is that listed after the "PROTOCOL" heading in the manuscript. We now highlighted in yellow.

10. Assuming that the content for the video is ALL that is listed under the heading "PROTOCOL" then it is important that 1) the steps involved in production of the "half-dome" cryolite glass prosthetic eye initially are fully covered at the start of the video.

then 2) the steps involved in customising the half dome cryolite glass prosthetic eye to a patient would then be the second part of the video.

Thank you for this important information. Unfortunately, it is not possible to present both main steps in one and the same manuscript or video due to length restrictions. Pre-producing the "half-done" prosthetic eye is very elaborate, and therefore a separate video and manuscript would be necessary. In addition, most ocularists focus on this second step and buy the half-done prosthetic eyes from a supplier that focus on producing these half-done prosthetic eyes. Furthermore, the procedure reported here is the most critical step in respect to individual patients’ care and seems to be the most interesting issue for the ophthalmologists.

11. The major discussion topic is a comparison of prosthetic eyes made from the two materials a) cryolite glass, and b) PMMA. This comparison should include the care that needs to be exerted by the patient in maintaining the prosthesis between annual visits to the ocularist. I understand that the recommended care regimes are

(i) for glass prosthetic eyes is removal and cleaning on a daily basis.

(ii) for PMMA prosthetic eyes is removal and cleaning at an interval between monthly and six monthly.

Thank you for this important amendment. We included patient care requirements and changed the manuscript as follows:

"So far the recommended care regime for glass prosthetic eyes included removal and cleaning on a daily basis while for PMMA prosthetic eyes removal and cleaning at an interval between monthly and six-monthly was recommended. However, the results of newest studies suggest that cryolite glass prosthetic eyes should not be removed daily for cleaning and that further research should be undertaken to develop a cleaning protocol. This protocol would recommend a minimum period of wear between cleaning sessions and should consider individual conditions, such as climatic, psychological, and environmental factors."

12. Minor Concerns:

The spelling and grammar is difficult in places and needs revision.

Thank you for this amendment. All spelling or grammar issues were corrected.

Reviewer #2:

1. Manuscript Summary:

Need a revision of technique illustrations and results achieve

Thank you for this important amendment. We included various figures to show each step of the technique more detailed and to show possible results of our technique. Please see the pictures.

2. Major Concerns:

Technique defined by Author not clear, how the conforming of shape with other shapes being done

Thank you for this important amendment. The ocularists form the shape of the prosthesis based on the model of the existing prosthetic eye with sense of proportion. All these steps are just a matter of practice and also a question of experience resulting in an over 6-years training period for ocularists. We changed the manuscript as follows:

"1.10.  While continuously rotating the "half-done" glass prosthetic eye, melt the glass at the backside of the prosthesis (Figure 6) and reduce the volume of the backside by suction with help of the mouthpiece so that the backside shape is equal to the sample prosthesis or the desired shape."

3. No Pre and post picture which shows typical a custom eye prosthesis does look, and how the replica or transfer of the posterior surface is done and with what level of precision
Images should be ideally attached, as the technique described by author is already being used in dusty over 100 Years.

Thank you for this important amendment. We included various figures to show the technique and the level of precision more detailed and to show possible results of our technique. Please see the pictures.