Editorial comments:

Changes to be made by the author(s):

1. The language in the manuscript is not publication grade. Please employ professional copy-editing services.

We have employed professional copy-editing services to improve the language. A native English speaker has edited the manuscript to improve the flow and readability of the text, hopefully, the new version can fit the JoVE standard.

2. Please revise lines 57-58, 94-96, and 150-152 to avoid previously published text.

We revised these sentences according to your suggestion.

3. Keywords: Please provide at least 6 keywords or phrases.

We added more keywords as following:

Keywords: Pulmonary hypertension, hemodynamics, right ventricular pressure, pulmonary arterial pressure, mouse, catheterization, hypoxia

4. Please add a Summary section before the Abstract section to clearly describe the protocol and its applications in complete sentences between 10−50 words: “Here, we present a protocol to …”

We added a sentence before the Summary section: “Here, we present a protocol to perform an invasive hemodynamic assessment of the right ventricle and pulmonary artery in mice using the open-chest approach. “

5. Abstract: Please expand to include an overview of the method and a summary of its advantages, limitations, and applications.

We revised the Abstract section according to your suggestion.

6. Introduction: Please revise to include all of the following:

a) A clear statement of the overall goal of this method

b) The rationale behind the development and/or use of this technique

c) The advantages over alternative techniques with applicable references to previous studies

d) A description of the context of the technique in the wider body of literature

e) Information to help readers to determine whether the method is appropriate for their application

We revised the Introduction section according to your suggestion.

7. JoVE cannot publish manuscripts containing commercial language. This includes trademark symbols (™), registered symbols (®), and company names before an instrument or reagent. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials. You may use the generic term followed by “(Table of Materials)” to draw the readers’ attention to specific commercial names. Examples of commercial sounding language in your manuscript are: Beijing Vital River Laboratory Animal Technology Co., Ltd., Millar Mikro-Tip, Terg-a-zyme, PowerLab, ADInstruments, LabChart, etc.

According to your suggestion, Millar Mikro-Tip, Terg-a-zyme, PowerLab, ADInstruments and Beijing Vital River Laboratory Animal Technology Co., Ltd are all deleted in main text and removed to the Table of Materials as needed

The LabChart is a software name, so we only keep this in the main text.

8. Please revise the Protocol to contain only action items that direct the reader to do something (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “NOTE.” Please include all safety procedures and use of hoods, etc. However, notes should be used sparingly and actions should be described in the imperative tense wherever possible. Please move the discussion about the protocol to the Discussion.

We revised the Protocol section according to your suggestion.

9. Lines 80-86, 133-138: Please write the text in the imperative tense. Any text that cannot be written in the imperative tense may be added as a “NOTE”.

These sentences are related to 2 paragraphs, and they are not used to describe the surgery procedures, we checked the other published JoVE papers that the authors used a similar style for the description, so we ask for keeping current description.

10. Please add more details to your protocol steps. There should be enough detail in each step to supplement the actions seen in the video so that viewers can easily replicate the protocol. Please ensure you answer the “how” question, i.e., how is the step performed? Alternatively, add references to published material specifying how to perform the protocol action. See examples below.

We revised and added more details in the Protocol section to ensure viewers can replicate the protocol.

11. Line 89: Please specify the incubation temperature.

We revised the sentence as “Soak the pressure transducer catheter (SPR-1000, size: 1F) in 0.9% saline at room-temperature for at least 30 min before the hemodynamic experiment.”

12. Line 91: Please describe how to calibrate the pressure transducer.

The calibration details are added to the Protocol section.

13. Line 93: Please list an approximate volume to prepare.

It is about 10 mL 1.0% digestive enzyme solution, the main text has been updated.

14. Line 96: Please specify the light source used in this protocol.

A standard light source for microsurgery was used, we have updated the information in the Table of Materials.

15. Please specify surgical tools used throughout the protocol.

We described the surgical tools used in the Protocol section and updated the tools information in the Table of Materials.

16. Line 126: Please specify the euthanasia method.

We specify the euthanasia method as “After the final measurements, euthanize mice humanely through i.p. injection of an excess dose of 2,2,2-Tribromoethanol solution.”

17. Figure 6: Please define error bars in the figure legend.

The error bars in the figure legend are defined as standard deviation in the figure legend.  The figure and figure legend are all updated.

18. Please remove the titles and figure legends from the uploaded figures. Please include all the figure Legends together at the end of the Representative Results in the manuscript text.

All of titles and figure legends were all moved to the end of the Representative Results section in the main text.

19. Please upload Table of Materials to your Editorial Manager account as an .xlsx file. Please ensure that it has information on all relevant supplies, reagents, equipment and software used, especially those mentioned in the Protocol. Please sort the items in alphabetical order according to the name of material/equipment.

The Table of Materials has been updated to an .xlsx file following the JoVE standard.

20. Please ensure that the references appear as the following: [Lastname, F.I., LastName, F.I., LastName, F.I. Article Title. Source. Volume (Issue), FirstPage – LastPage (YEAR).] For more than 6 authors, list only the first author then et al. Please do not abbreviate journal titles. See the example below:

Bedford, C.D., Harris, R.N., Howd, R.A., Goff, D.A., Koolpe, G.A. Quaternary salts of 2-[(hydroxyimino)methyl]imidazole. Journal of Medicinal Chemistry. 32 (2), 493-503 (1998).

All the references are all formatted to fit the JoVE standard.

Reviewers' comments:

Reviewer #1:

Manuscript Summary:

In this paper, the authors established a right heart hemodynamic measurement methodology and validated with healthy and PAH mice. Authors measured only right ventricular pressures, without measuring volumes, and thus the method gives limited information about ventricular function. Even though, this is an interesting, brief, well-structured and useful manuscript. Authors focus also in some troubleshooting, which is very valuable for other researchers to implement the methodology. There are several major points requiring additional workout.

Major Concerns:

-It is not clear on which basis the protocol of 4 weeks of hypoxia was selected. Authors should give a precise justification or provide some reference that justifies their selection.

We added more references and justification in the Discussion section. Based on the published papers, hypoxia period is viable from 3 to 4 weeks to induce pulmonary hypertension in mice. We did pilot chronic hypoxia experiment in C57 wild type mice showed that 4 weeks of hypoxia could induce a stable pulmonary hypertension model.

We have added one more sentence in the Discussion section as following:

“ Finally, the hypoxia period is viable from 3 to 4 weeks for the hypoxia-induced hypertension model in mice. Our data showed that 4 weeks of hypoxia can induce a stable pulmonary hypertension model in C57BL mice, the PAP and RVP level are comparable with the literature, further study need to address how long the PAH model can be maintained if put back to normoxia condition for different hypoxia protocols.”

References:

- Ciuclan L., et al. A novel murine model of severe pulmonary arterial hypertension.  American Journal of Respiratory and Critical Care Medicine. 184 (10), 1171-1182 (2011).  3 weeks of hypoxia

- Shatat M. A., et al.  Endothelial Kruppel-like Factor 4 modulates pulmonary arterial hypertension. American Journal of Respiratory Cell and Molecular Biology. 50 (3), 647-653 (2014).  3 weeks of hypoxia

- Chen M. et al. Berberine attenuates hypoxia‐induced pulmonary arterial hypertension via bone morphogenetic protein and transforming growth factor‐β signaling. Journal of Cellular Physiology. https://doi.org/10.1002/jcp.28370 (2019).  4 weeks of hypoxia

-Did the authors make any anaesthetic maintenance during the hemodynamic procedure? If yes, how? If no, why?

We have added more details about maintenance of anesthesia in the Protocol section:

“Anesthetize mice with 2,2,2-Tribromoethanol, 250 mg/kg, intraperitoneal (i.p.) injection; if needed, repeat supplemental doses at 1/3 to 1/2 of the original dose during the procedure.”

-Why did the authors choose to make animal ventilation through an intubation through a "T" shaped incision on the trachea, instead of ventilation through a direct intubation trough the animal mouth? It would be less invasive.

We have revised in the first paragraph of the Discussion section to clarify as following:

“Tracheal intubation is the first important step for open chest surgeries. The classic method of tracheal intubation for small animals, such as rats or mice, involves making a "T" shaped incision on the trachea and directly inserting a "Y"-type tracheal tubing into the trachea. In practice, we find that this method is not easy during operation. The “Y”-type tracheal tubing is too large for small animals and forms an angle with the trachea; thus, it is difficult to fix the tubing in place. Additionally, once the intubation tubing accidentally comes out from airway during open chest surgery, it usually results in animal death because of loss of mechanical ventilation support. Therefore, we modified the method of endotracheal intubation   by making an incision on the skin, separating the muscle layer to expose the trachea (Figure 1C), and directly inserting the tracheal tubing into the airway through the animal mouth. The placement of tubing in the trachea can be conveniently confirmed by clamping the trachea using forceps (Figure 1D). A 22 G intravenous catheter (after removing the guide needle and only using the sheath catheter) is used as the intubation tubing, and the tubing can be easily fixed after intubation. This is a safe way for management during surgery and can significantly improve the success rate of small animal open chest surgery. However, this method requires a period of training and practice.”

-How did the authors calculate the respiration rate and tidal volume?

Calculate respiration rate and tidal volume based on body weight according the ventilator user manual ( on the page 16 and 17: www.harvardapparatus.com/media/harvard/pdf/Inspira_557058_9.pdf).

Page 16:

Page 17:

We have added a reference to the Discussion section as following: www.harvardapparatus.com/media/harvard/pdf/Inspira_557058_9.pdf

-How were the animals euthanized? Which method of euthanasia did the authors used?

We have clarified this in the Protocol section as:

“After the final measurements, euthanize mice humanely through i.p. injection of an excess dose of 2,2,2-Tribromoethanol solution.”

-Even for a methodological paper the animal number is too small to show representative results.

We did more control mice (n=10) for the right ventricular hemodynamic measurement to show the reproducibility of this method. However, we are very sorry that we did not add more hypoxia mice data because of the tight revision time and the preparation of animal models takes a long time.

-Authors should provide if there was any mortality during the hypoxia and, most importantly, during the hemodynamic protocol.

The C57BL mice are well tolerated to mild hypoxia (10% oxygen) for 3 to 4 weeks, the surgery mice rarely died during the procedure if maintain the mechanical support well and no severe bleeding during open chest procedure. We added one sentence to the Discussion section:

“Although this method is a safe procedure for a well-trained surgeon, the maintenance of tracheal intubation and prevention of bleeding during open chest is important to avoid animal death.”

- The goals of the study are not completely new (even in jove publications: https://www.jove.com/video/53335/hemodynamic-characterization-rodent-models-pulmonary-arterial & https://www.jove.com/video/55065/shunt-surgery-right-heart-catheterization-vascular-morphometry-rat ). Having this in mind, authors should better discuss and explain the important differences in their results in comparison with the previous published studies, specially concerning results in the hemodynamic evaluation and hypoxia protocol used.

We have modified the Discussion section to compare the results in RVP and PAP for PAH model.

Minor Concerns:

-Authors should provide the reference for the approval of the animal protocol.

We have added approved protocol information in the Protocol section as following:

“The animal protocol was reviewed and approved by the Institutional Animal Care and Use Committee at Fuwai Hospital, Chinese Academy of Medical Science, Peking Union Medical College (NO.0000287).”

-Authors should provide reference to the brand/model of the hypoxia chamber and oxygen sensor.

We have added O2 controller and hypoxia chamber information to the Table of Materials.

-Authors should provide the reference for the LabChart Software.

We have added the LabChart version information also a reference website of the LabChart to the Protocol section. (LabChart: https://www.adinstruments.com/products/labchart?creative=290739105773&keyword=labchart&matchtype=e&network=g&device=c&gclid=CjwKCAjwxrzoBRBBEiwAbtX1n42I2S06KmccVncUHkmExU8KKOXXREyzx8bvTrxYMSze-ooE0atcbRoCliwQAvD_BwE.)

-Authors should actualize the most recent publications references.

We updated the most recent publications in the Introduction and Discussion section about the hypoxia protocol and hemodynamics.

-The language must be checked throughout the paper by a native speaker.

According to your suggestion, a native English speaker has helped us to edit and polish the manuscript.

Reviewer #2:

Manuscript Summary:

This is a good description of the open-chest right ventricle pressure measurement but the pulmonary artery measurements are barley mention in the discussion.

Major Concerns:

1. Removing the catheter while measuring to clean the electrode. This will cause bleeding through the hole in the RV and they don't mention how long the digestion and rinsing take. The higher the pressure the more bleeding. We always make sure my probe is clean before inserting and when we finish a measurement we wipe it gently with a wet gauze before returning it to the saline beaker. This should remove any protein or debris that it has picked up while reading and pulling it out of the ventricle.

We agreed with your point. If the measurement can be done correctly so it is not necessary to remove the catheter for cleaning, at least it might cause minor bleeding although a cotton-tip can prevent the bleeding.

While we trained the beginners such as young residents or fellows to do this open-chest procedure, we noticed that they struggled for the catheterization and the maintenance of the catheter in the ventricle, it took much longer time for the measurement, so it usually causes the catheter zero-drifting. In this case, to avoid to generate artificial data, we have to take the catheter out of ventricle to clean the blood components attached to the sensor. Soak the catheter with warm 1.0% digestive enzyme solution, gently wipe the catheter with saline-soaked gauze; it is beneficial to remove blood components from the catheter. If soaking in the enzyme in room-temperature (21-24 degree), it usually takes about 15 min for digestion.

We have revised the main text in the Discussion section to clarify this concern.

2. No figure showing an actual pressure tracing when the probe is in the pulmonary artery.

We updated Figure 6 according to your suggestion. The representative pulmonary artery pressure curves for control and PAH mice are shown in Fig 6E and 6F respectively.

3. In the Protocol (Lines 117-120) they say they make a tunnel with the 25 G needle and then put the catheter in through the tunnel but in the discussion (Lines 188-193) the implication is that they leave the needle in and put the catheter through it. If they are leaving the needle in it can cause major problems such as shorting out the probe if the contacts touch the metal of the needle. We make the tunnel, remove the needle and push the probe through the tunnel left by the needle.

We agreed with your point.  We prefer the procedure you mentioned above for the safety of the catheter.

When we trained the beginners to do catheterization, it was difficult for them to insert the catheter through the pre-punctured tunnel on the beating heart, so as an option we modified the method to insert the catheter during the needle returning from the ventricular wall. It is helpful for the catheterization, but we should always use a small size needle no more than 25 G to puncture a tunnel, the bigger size needle not only increases the risk of damage catheter but also causes more bleeding.

We did minor revision in the Discussion section.

Minor Concerns:

1. Correcting the zero during a recording is something that we do not do. We zero the catheter at the beginning of the run only.

We agreed that do not correct the zero during recording.

If the catheter zero-drifting does appear during measurement, it means that the data recorded may be artificial, take out the catheter from the ventricle and recheck the zero, then try to do the measurement again.

2. How much blood flow is there once the catheter is introduced into the PA? We know they are using the 1F catheter, but it seems like it would be almost the same diameter as the PA and would impede flow through the valve. It seems that in smaller mice, the catheter would take up the diameter of the PA. We don't do the closed chest method on mice less that 16 grams and prefer them to be at least 20 grams for the 1.4F probe.

We agreed that it is difficult to hemodynamics in mice less than 16 gram; the mice we used in this experiment are all over 24 gram, so they are big enough to do the measurement.

The catheter size is 1F; it is about 0.33 mm in diameter, considering of the pulmonary artery diameter is about 1.3-1.5 mm in adult mice, the catheter should not significantly affect the pulmonary blood flow during the catheter inserted into the artery. We did not measure the pulmonary blood flow during the pulmonary artery catheterization procedure, so we don’t have the data currently to show pulmonary blood flow on this condition.

3. Using 1% Terg-a-zyme is very hard on the electrodes of the catheter. We usually just wipe ours down with a wet (water) gauze after each mouse, before returning the probe to the beaker of saline and again after the experiment before putting the catheter away to prevent salt build up.

We agreed that the Terg-a-zyme might be tricky, we noticed that soak the catheter with warm 1.0% digestive enzyme solution, gently wipe the catheter with saline-soaked gauze, it is beneficial to remove blood components from the catheter, then put the catheter into normal saline until next measurement.

4. Should mention the size of the Millar catheter in the text not just in the materials list.

We added the pressure transducer catheter size in the main text (the Protocol section); the catheter size is 1F in diameter.

5. They didn't discuss it but in Figure 6 they show the systolic pressure in the right ventricle and the systolic pressures in the pulmonary artery and they appear to be the same. They don't mention the variability in pressure between mice which is fairly high with this model and usually requires larger numbers to achieve significance. Are the error bars standard deviation or standard error of the mean?

The RVSP and sPAP are comparable in Figure 6. These data are comparable with references, we did not add more sentences to discuss.

The new Figure 6 is updated, we did more control mice to increase N number (N=10), and the error bars in Figure 6 are all changed to standard deviation.

About the variability such as in systolic pulmonary artery pressure (sPAP) between mice, it is about 17.4 to 25.8 mmHg for room air control mice, and 30.1 to 38.4 for pulmonary hypertension mice. Also, we noticed that while we trained the young residents or fellows to do chronic hypoxia, sometimes failed to maintain a consistent 10% O2 level in the hypoxia chamber, it might cause significant variation in hypoxia mice, and some mice may not develop to pulmonary hypertension only developed a mild elevation of pulmonary artery pressure. Thus, for the formal hypoxia experiment, we monitored the hypoxia chamber oxygen level in person at least twice a day, to confirm that the hypoxia process could generate a consistent 10% O2 level in the hypoxia chamber. In this way, the hypoxia mice can generate a significant increased pulmonary artery pressure and acceptable variation compare to room air control mice.

6. What is the mortality rate with this procedure? The numbers of mice are very small in this experiment. Also they don't mention the variability in pressure between mice which is fairly high with this model and usually requires larger numbers to achieve significance.

About the mortality of this open-chest procedure, it depends on the surgeon. For a well-trained surgeon, animal rarely died during the open-chest surgery if mice could tolerate anesthesia challenge. During the training period, when the young residents or fellows do such open-chest surgery in small mice, sometimes failed for ventilator support such as intubation tubing failure can cause animal die.

We did more control mice (n=10) for the right ventricular hemodynamic measurement to show the reproducibility of this method. However, we are very sorry that we did not add more hypoxia mice data because of the tight revision time and the preparation of animal models takes a long time.

About the variability, we answered the question in question 5.

7. Should mention size of the catheter in the Protocol section.

We added the pressure transducer catheter size in the Protocol section, it is a SPR-1000 catheter, 1F.