Reviewers' comments:

We would like to offer our earnest thanks to the reviewers for taking the time and effort to carefully critique our work. Because of their insights, we feel the quality of our manuscript has improved significantly. Our responses to the reviewers are bolded and italicized below. Significant changes within the manuscript are in red font color.

Reviewer #1:
The authors describe a protocol to isolate murine cardiac myocytes and perform single cell voltage or calcium imaging along with sarcomere shortening measurements. Using this technique, they confirm previously described APD restitution properties and response to 4-aminopyridine treatment. Interestingly, they also report reduction in sarcomere shortening due to reduced availability of calcium ions during calcium dye treatment and cytotoxicity due to prolonged FluoVolt dye excitation. The protocol very clearly details this procedure however there are some concerns that need to be addressed to improve the manuscript.

Major Concerns:
1. The abstract describes this manuscript as a protocol for murine cardiac cell isolation and imaging but the manuscript switches between rat and mouse protocols and rat hearts are used in figures while mouse heart protocols are described in corresponding manuscript text. For example, a rat heart is included in Figure 1A, so are the myocytes in the other Figure 1 panels also from rats? Clearly define scope of the manuscript and delineate rat and mouse sections.

We apologize for the confusion. We have used this protocol to isolate mouse, rat, and rabbit myocytes. The only thing that differs between species is the digestion time and concentration of the collagenase in the digestion buffer. In retrospect, while we started the study using mice, all the originally submitted data was recorded from rat. Therefore, to include mouse data, we added three new figures (Figure 2, 3, and 4).  The manuscript has been changed accordingly.

2. Is the dye incubation and the following washout out phase done at room temperature or is the preparation returned to the incubator for this stage? If it is at room temperature, comment on the effects of temperature change on electrophysiology. Parameters like AP take about 15-20 minutes to recover from temperature shock of going between 37 C and room temperature. Also, include how long you allow the cells to recover at 36 C before you image them.

Excellent point. To address this concern, the following was added to the manuscript on page 6 line 287 “Because excitation-contraction coupling is temperature dependent, make sure the cells have been perfused for 15 minutes before recording. This allows for myocytes to recover from the shock of going from room temperature back to 37ᵒ C as well as loosely attached cells to float away.”

3. The AP traces in Figure 3 is very noisy and this can affect the APD90 measurements. In Figure 4 the traces look much cleaner. Explain how this was accomplished and how was APD90 reliably measured in the former instance?

Another excellent point. The AP’s in the former Figure 3 were filtered but not ensembled averaged. The ones in Figure 4 are ensembled averaged traces that were filtered with a Savitzky-Golay Filter. For the voltage data presented in this manuscript, action potentials were filtered with a Savitzky-Golay Filter and ensembled averaged before being quantified. The fura-2 and sarcomere shortening data was ensembled averaged, but not filtered.

Regarding filtering, the Ionwizard has three digital filters that can be used to analyze data. A low-pass Butterworth filter, a zero-phase low-pass Butterworth filter, and a Savitzky-Golay polynomial smoothing filter. To illustrate the filtering process, a new figure was added (Figure 4). Because the digital filter settings can distort the real signal, hence alter the APD measurements, the following sentences were specifically added on page 7 at line 335: “Compared to calcium signals, single cell voltage tracings are smaller in amplitude and need post-processing to obtain a useable signal. Figure 4B shows an ensembled averaged AP made from the APs in Figure 4A. Figures 4C and 4D show an ensembled average AP after a low pass Butterworth or a Savitzky-Golay digital filter was applied. Care must be taken when filtering the signal as not to distort the real data. Notice the subtle differences in the shape of the APs in Figure 4B, 4C, and 4D.”

Minor Concerns:
1. Image post-processing is mentioned in lines 266-267 but no further details on what processing methods are applied are given. Please describe. We apologize for our carelessness. The manuscript has been changed accordingly.  
2. Label parts of the system in Figure 1A, indicate where the rat heart is in this photo. Good catch. The parts of the system are now labeled, and the legend has been changed to appropriately describe the figure.  
3. It looks like Figure 2A and 2B are at different time scales but only 1 time scale bar is included. Please correct. We double-checked the scales. While it appears to be two different time scales, it is an optical illusion. Nonetheless, to avoid confusion we put a scale on each figure.  
4. Please describe how sarcomere shortening measurements were made. The following sentences were added to the manuscript on page 7, line 317: “Shortening on the other hand is obtained by a CCD sensor. Measuring in real time up to 1000 times per second, the acquisition software performs an average of the lines within an area of interest to create a well resolved striation pattern. A fast Fourier transform (FFT) is then calculated. The peak within the power spectrum represents the average sarcomere spacing. Changes in the sarcomere spacing during pacing are then plotted and subsequently quantified.”                                                                                                                                    5. Traditional Krebs Henseleit solution uses sodium bicarbonate as a buffer, please make a note that the solution you use is a modified Krebs Henseleit solution with HEPES buffer. The following sentence has been added to page 1 line 85: “Note traditional Krebs Henseleit solution uses sodium bicarbonate as a buffer and the solution in this protocol uses Krebs Henseleit solution with HEPES buffer. Additionally, KHB-B has been changed to KHB-HB”
6. Include concentration of HEPES in culture media on line 195. 25 mM has been listed.  
7. Mention duration of dye washout phase in line 221-222. The manuscript has been changed accordingly.
8. Technically, you're not starting to pace at 0V, correct the statement in line 241-242. Another good catch. The manuscript has been changed accordingly.
9. In line 196 you state that the cells can be cultured for 1 week and in line 259 you state 4 days. Please clarify and keep consistent. It has been clarified.  


Reviewer #2:
Manuscript Summary:
The authors describe a method to isolate viable ventricular cardiomyocytes from mouse hearts using a standard Langendorff-perfusion based method; moreover, the authors comprehensively describe a simple and repeatable method to employ the largely available Ion-Optics fluorescence system (with the Fluo-3/4 filter option) to simultaneously record either sarcomere shortening and intracellular calcium (using Fluo-4) or sarcomere shortening plus membrane voltage (using FluoVolt)

Major Concerns:
1) the perfusion system is described as a constant-pressure Langendorff apparatus. Is the constant pressure provided by gravity? If so, what is the height at which the solution reservoir is positioned during heart perfusion?

We apologize for the confusion. The height of the solution reservoir is 60 cm. Figure 1 has been changed accordingly to indicate this.

In some mouse CM isolation protocols, the perfusion pressure is increased in the second half of collagenase perfusion, but the authors do not mention it here. At page 2 line 119 the authors describe a circulating water bath and suggest to set a flow rate of 12-14 mL/min. What is it used for? To refill the high reservoir? The authors should clarify these points as they make the procedure unclear.

Excellent observation. We apologize for the confusion. The text has been changed to the following on page 3 line 144: “With the solution reservoirs being set at a height of 60 cm, the circulating waterbirth needs to be set to 41ᵒ C to have the perfusate be 37ᵒ C. Unlike previously reported protocols, the height of the reservoir does not need to be changed.”

2)Page 3 line 166. After perfusion, the heart chambers are placed into separate dishes and collegenase solution is added to each dish. How long should the chambers be kept in this digesting solution before mincing them with scissors? Or should they be minced immediately? If so, the authors should suggest the time allowed for mincing, straining and centrifuging the cells before the resuspension in stopping buffer, as persistence of isolated cells in enzymatic solution might damage cells.

We apologize for the omission. The text has been changed to the following on page 4 at line 198: “After digestion, separate the chambers of the heart. Place each chamber into a separate well of a six well plate. Pour 5 mL of collagenase solution into each well.  Immediately start mincing the heart tissue using scissors. Tissue chunks should be approximately 1 cubic mm.  Using sterile transfer pipettes, gently triturate the minced heart tissue. The solution should turn cloudy.”

3)In your trial experiments, you paced the cells up to 4Hz. In my experience, good mouse cardiomyocytes can easily be stimulated up to 7-8 Hz without damaging them. Have you tried stimulating them at 7-8Hz?

Specifically, 7-8 Hz, no. The pacing protocol is a variation of two previously published pacing protocols (please see newly included references 10 and 11). Nonetheless, your point is well taken. A sinus rate of 4 Hz would be a lethal bradycardia in a mouse. Therefore, in the past we have tried 10 Hz, which corresponds to a healthy resting heart rate in mice. In the revised manuscript, we have included 3 new figures to illustrate this.

4)Although the risk of fluovolt photodamage is known, in my experience it only occurs when a very powerful illumination is used. Have you tried adjusting the illumination intensity in order to reduce fluovolt photo damage?

Yes, we have. The IonOptix illuminating source is a Xenon Short Arc Lamp. To control the intensity, the IonOptix has 5 neutral density filters preinstalled. They are ND 0.2, ND 0.3, ND 0.6, ND1/0, and ND 1.3. They correspondingly reduce the transmitted light by 37, 50, 75, 90, and 95%. Using these we can get recordings of upwards of 15 seconds without phototoxicity. We have not tried other neutral density filters. In order to address this concern, the following sentences were added to page 10 line 457 of the manuscript: “While our protocol was optimized for 5 sec recordings, if longer voltage traces are needed, additional neutral density filters will need to be purchased. The system described in the protocol comes with neutral density filters that reduce the transmitted light by 37, 50, 75, 90, and 95%.”

Minor Concerns:
1)Is fluo-4 or Fluovolt loading performed at room temperature?  The loading of all dyes was performed at room temperature. The manuscript has been changed accordingly to reflect this.  
2)Figure 1: please add the Y scale for sarcomere length. A y-scale has been added.
3)Figure 5: please describe the situation in the caption: how long was the cell paced and illuminated before developing damage? The manuscript has been changed accordingly.