# Submission ID #: 60213

# Scriptwriter Name: Anthony Iannazzi

# Project Page Link: http://www.jove.com/files_upload.php?src=18365013

# Title: Characterization of Sickling During Controlled Automated Deoxygenation with Oxygen Gradient Ektacytometry

# Authors and Affiliations:

# Minke A.E. Rab1,2, Brigitte A. van Oirschot1, Jennifer Bos1, Celeste K. Kanne3, Vivien A. Sheehan3, Eduard J. van Beers2, Richard van Wijk1

# 1Laboratory of Clinical Chemistry and Haematology, University Medical Center Utrecht, Utrecht University, The Netherlands

# 2Van Creveldkliniek, University Medical Center Utrecht, Utrecht University, The Netherlands

# 3Department of Pediatrics, Division of Hematology/Oncology, Baylor College of Medicine, Houston, Texas, USA

# Corresponding Author:

# Minke A. E. Rab		M.A.E.Rab@umcutrecht.nl

# Email Addresses for Co-authors:

# B.oirschot@umcutrecht.nl

# J.F.Bos-17@umcutrecht.nl

# Celeste.Kanne@bcm.edu

# vxsheeha@txch.org

# E.J.vanBeers-3@umcutrecht.nl

# R.vanWijk@umcutrecht.nl

Author Questionnaire:

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique? No

2. Does your protocol include software usage? Yes

If yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps.

3. Which steps from the protocol section below are the most important for viewers to see? Please list 4-6 individual steps using the step numbers listed in this document. This information is important to prepare your Videographer for your shoot.

3.2, 3.4, 4.1, 4,2, 4.3, 5.1 and 5.3

4. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1-2 individual steps using the step numbers listed in this document.

The most difficult aspect of this procedure is to work as standardized as possible. Especially in the sample preparation and cleaning of the ektacytometer. Steps 4.1, 4.2 and 5.3.

5. Will the filming need to take place in multiple locations? Yes

If yes, how far apart are the locations? The hematology analyzer is situated 20 meters away from the ektacytometer, in another room.

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

- REQUIRED Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Minke Rab: Oxygen gradient ektacytometry is able to determine the oxygen tension at which RBCs from a patient with sickle cell disease start to sickle. This oxygen tension is patient specific and can serve as a biomarker for disease severity [1].

    - INTERVIEW: Named author says the statement above in an interview-style statement while looking slightly off-camera.

  - Minke Rab: Oxygen gradient ektacytometry is a functional assay that assesses the individual sickling tendency. To date, no other functional assays are available or used [1].

    - INTERVIEW: Named author says the statement above in an interview-style statement while looking slightly off-camera.

OPTIONAL Interview Statements: (Said by you on camera)

  - Richard van Wijk: In sickle cell disease, this technique can be used to monitor treatment efficacy and it can be used for the development of new treatments [1].

    - INTERVIEW: Named author says the statement above in an interview-style statement while looking slightly off-camera.

  - Jennifer Bos Brigitte van Oirschot: Somebody who tries this for the first time may experience that the technique is very sensitive to sample preparation and sample handling. My advice is to work as standardized as possible and to develop a routine.

    - INTERVIEW: Named author says the statement above in an interview-style statement while looking slightly off-camera.

  - Jennifer Bos Brigitte van Oirschot: Visual demonstration of this technique is important because it is very sensitive and there are numerous factors that influence the measurement which you do not know when you start working with this technique [1].

    - INTERVIEW: Named author says the statement above in an interview-style statement while looking slightly off-camera.

Ethics title card: (for human subjects or animal work, does not count toward word length total)

  - All procedures were approved by the ethical committee of the University Medical Center Utrecht (UMCU) and in accordance to the Declaration of Helsinki. Patients enrolled at the Texas Children’s Hematology Center (TCHC) were approved by the local IRB and in accordance with the Declaration of Helsinki.

# Section - Protocol

- Start-up of the ektacytometer

  - To begin this procedure, switch on the computer and the ektacytometer [1] and open the software on the computer [2]. Open the nitrogen cylinder to make sure the nitrogen is available to deoxygenate the sample [3].

    - Establishing shot of the talent approaching the work area, and turning on the computer and ektacytometer.

    - Talent opens the software on the computer.

    - Talent opens the nitrogen cylinder.

  - Lower the bob in the cup, and make sure the cup can turn freely [1]. [2-TXT].

    - Talent lowers the bob in the cup, and checks if the cup turns freely.

  - When the software is running, check for the message [1] “Make sure the gas valve is open” and click OK [2].

    - Talent approaches the computer to check for the message.

    - SCREEN: *To be provided by authors: Screen capture video showing the message described above. Then, click OK. Authors, upload all screen capture videos to your project page.

  - Ensure that the ektacytometer starts the self-check process, and select Start [1]. If the self-check fails, rerun it by clicking Hardware check, pO2, Self check [2].

    - SCREEN: *To be provided by authors: Ensure that the ektacytometer starts the partial pressure of oxygen self-check process, and click Start. Authors, upload all screen capture videos to your project page.

    - SCREEN: *To be provided by authors: Show a failed self-check, then rerun it by clicking Hardware check, pO2, Self check. Authors, upload all screen capture videos to your project page.

  - Choose pO2 scan from the different tests listed on the left. Choose Settings at the right of the screen and ensure they are set as per the parameters listed in Table 1 of the text protocol, making sure to keep the same settings for every measurement. Save these settings by clicking OK, then clicking OK again [1].

    - SCREEN: *To be provided by authors: Choose pO2 scan from the different tests listed on the left. Choose Settings at the right of the screen and ensure they are set. Clicking OK, then click OK again to save the settings. Authors, upload all screen capture videos to your project page.

- Sample collection and preparation

  - First, collect blood samples by venipuncture in a tube containing EDTA [1] and store the blood at 4 degrees Celsius for a minimum of 30 minutes, but no longer than 24 hours [2].

    - Establishing shot of the talent approaching the work area with the tubes in hand.

    - Talent stores the blood samples in a refrigerator.

  - After this, mix the samples gently by inversion to homogenize it [1-TXT]. Let the sample warm up to room temperature on a roller bench before the measurement [2-TXT].

# Videographer: This step is one of the most important for viewers to see.

    - Talent inverts the samples to mix them. TEXT: Do not shake the sample.

    - Talent sets the samples on a roller bench to warm up. TEXT: See text for details on warming duration.

  - [1]. Use a hematology analyzer to measure the complete blood count [2] by placing the aspiration needling into the tube [3] and pressing the button behind the needle of the hematology analyzer to begin the measurement [4]. An analyzer will use between 20 and 200 microliters of whole blood [5].

    - Close up of the tubes.

    - Talent approaches a hematology analyzer with a sample in hand.

    - Talent places aspiration needle into the tube. (Author Comment: Also filmed a close-up of the talent placing the aspiration needle in the tube.)

    - Talent presses the button behind the needle of the hematology analyzer.

    - Added shot: Shot of the screen of the hematology analyzer with the results displayed.

  - Then, standardize the whole blood sample to an RBC count of 200 million red blood cells in 5 milliliters of PVP by calculating the volume of the sample that will be added [1].

# Videographer: This step is one of the most important for viewers to see.

    - Talent calculates how much sample volume must be added to the PVP.

- Oxygen gradient ektacytometry measurement

  - To begin, cut a pipet tip, and prewet it by resuspending the blood three times [1]. Then, pipet the calculated sample volume into PVP vial [2]. Gently mix the sample manually by inversion until it is homogenous [3].

# Videographer: This step is one of the most important for viewers to see, and one of the most difficult in the procedure.

    - Talent resuspends the blood to prewet the pipette tip (Author Comment: Also filmed a close-up of the pre-wetting of the pipette tip.)

    - Talent transfers the sample volume into PVP. [Shots 4.1.2 and 4.1.3 combined]

    - Talent inverts the sample to mix it.

    - Added shot: close-up of the injected of sample volume in PVP. (Editor: This appears to be a close up take of 4.1.2)

  - Slowly draw 2.0 milliliters of the blood and PVP mixture into a 3 milliliter syringe without the needle [1]. Push the plunger to remove any visible air bubbles and excess sample solution, until between 1.5 and 1.8 milliliters are left in the syringe [2].

# Videographer: This step is one of the most important for viewers to see, and one of the most difficult in the procedure.

    - Talent uses a syringe without the needle to draw up some of this blood/PVP mixture.

    - Talent pushes the plunger in the syringe to remove air bubbles.

  - Inject the total sample volume slowly and evenly into the bob through the connector [1-TXT]. Make sure that the level of the sample is above the oxygen sensor and the small suction hole [2].

# Videographer: This step is one of the most important for viewers to see.

    - Talent injects the total sample volume into the bob through the connector. TEXT: Do not leave any sample solution in the syringe.

    - CU or ECU shot showing that the level of the sample is above the oxygen sensor and the small suction hole.

  - In the software, click New and fill in the sample identifier, remarks, date of donation, and viscosity of PVP. Click OK and Aspirate [1]. After 60 seconds, the cup will rotate and aspirate the sample for 15 seconds [2].

    - SCREEN: *To be provided by authors: Click New and fill in the mentioned data. Click OK and Aspirate. Authors, upload all screen capture videos to your project page.

    - Show the cup rotating and the sample being aspirated.

  - When the rotation stops, click OK [1]. Close the machine lid [2]. Click Continue and Start now as oxygen gradient ektacytometry is done with a fixed gain [3-TXT].

    - SCREEN: *To be provided by authors: Click OK. Authors, upload all screen capture videos to your project page. Videographer: This may be better shown as a shot and not a screen capture. Please film a shot of the talent performing this action.

    - Talent closes the machine lid.

    - SCREEN: *To be provided by authors: Click Continue and Start now. Authors, upload all screen capture videos to your project page. TEXT: Measurement takes ~28 min.

  - After the measurement, print the report that shows the curve and parameters that are automatically calculated by the software [1]. Ensure that the raw data is automatically stored in the designated folder in the Settings [2].

    - Talent prints the report.

    - SCREEN: *To be provided by authors: Check the settings to show that the raw data is automatically stored. Authors, upload all screen capture videos to your project page.

- Cleaning of the ektacytometer

  - First, remove the sample syringe and replace it with a syringe filled with distilled water or deionized water [1]. Press Clean and slowly flush the connector during rinsing, making sure to flush in both directions [2].

# Videographer: This step is one of the most important for viewers to see.

    - Talent removes the sample syringe and replaces it with a syringe filled with water.

    - Talent presses Clean and flushes the connector. (Editor: When editing this shot, please make sure that the computer screen is only partially visible so that the drop down in the curve – which I the drop down in the y-axis – is not visible)

  - Next, remove the syringe and lift the bob [1]. Use a soft cloth to thoroughly dry the bob, and connector [2].

    - Talent removes the syringe and lifts the bob.

    - Talent dries the connector and bob with a soft cloth.

  - Then, use a large syringe to flush the connector in order to remove any water remaining in the tube and bob [1]. Keep the syringe at the inlet and block the outlet of the bob to get back pressure in the tubes, thereby removing the remaining water [2] and then, dry the cup while avoiding to touch the oxygen spot [5.3.2A]. Lower the bob in the cup, readying the machine for the next measurement [3].

# Videographer: This step is one of the most important for viewers to see, and one of the most difficult in the procedure.

    - Talent uses a large syringe to flush the connector. (Author Comment: shot in 5.2.2 take 1, 15 t/m 31 sec)

    - Talent blocks the outlet of the bob. (Author Comment: shot in 5.2.2 take 1 after 31 sec – 38 seconds)

# 5.3.2A. Added shot: Talent dries the cup. (shot in 5.2.2 take 1 after 38 seconds)

    - Talent lowers the bob in the cup.

- Shutdown of the machine

  - After ensuring that the machine is properly rinsed, check to make sure that the proper tubes relate to the cleaning solution [1].

    - Talent checks to make sure that the proper tubes relate to the cleaning solution.

  - Close the software, press Close, and press Start to begin the end-of-day cleaning program [1].

    - Talent closes the software as described.

  - Then, empty the waste bottle [1] and use a soft cloth to dry the bob and cup [2]. Close the lid of the machine [3], close the nitrogen cylinder [4], and turn off the computer and machine [5].

    - Talent empties the waste bottle.

    - Talent uses a soft cloth to dry the bob and cup.

    - Talent closes the lid of the machine.

    - Talent closes the nitrogen cylinder.

    - Talent turns off the computer and machine.

# Section – Results

- Results: Characterization of sickling by oxygen gradient ektacytometry

  - Oxygen gradient ektacytometry can be used to study the sickling behavior of red blood cells under a range of oxygen concentrations [1].

    - LAB MEDIA: Figure 3.

  - In oxygen gradient ektacytometry, continuous deoxygenation of the sample by nitrogen gas is followed by swift reoxygenation by ambient air [1]. Under these conditions, red blood cell sickling can be observed under deoxygenation [2]. This will cause a distortion of the diffraction pattern because sickled red cells will not align properly under the applied shear stress [3].

    - LAB MEDIA: Figure 3. Video Editor: Show only Figure 3A.

    - LAB MEDIA: Figure 3. Video Editor: Still show only Figure 3A.

    - LAB MEDIA: Figure 3. Video Editor: Still show only Figure 3A.

    - LAB MEDIA: Figure 3. Video Editor: Still show only Figure 3A.

  - Sickle red blood cells show a change in shape during deoxygenation that mimics conditions during oxygen gradient ektacytometry [1], but show no change in shape when oxygenated [2].

    - LAB MEDIA: Figure 3. Video Editor: Video Editor: Still show only Figure 3A. Emphasize the image labelled “Deoxygenated”.

    - LAB MEDIA: Figure 3. Video Editor: Video Editor: Still show only Figure 3A. Emphasize the image labelled “Oxygenated”.

  - This process results in distortion of the diffraction pattern during oxygen gradient ektacytometry, and thus in a decrease in the elongation index [1].

    - LAB MEDIA: Figure 3. Video Editor: Show only Figure 3B.

  - A representative curve obtained by the ektacytometer shows the 6 parameters that reflect different characteristics of the sickling behavior [1]: The maximum elongation index [2], the minimum elongation index [3], the delta EI [4], the point of sickling [5], the area under the curve [6], and the recovery [7].

    - LAB MEDIA: Figure 3. Video Editor: Show only Figure 3C.

    - LAB MEDIA: Figure 3. Video Editor: Still show only Figure 3C. Emphasize the horizontal line labelled “EImax”.

    - LAB MEDIA: Figure 3. Video Editor: Still show only Figure 3C. Emphasize the horizontal line labelled “EImin”. (This line is much shorter than EImax)

    - LAB MEDIA: Figure 3. Video Editor: Still show only Figure 3C. Emphasize the space denoted by the two-headed-arrow labelled “dEI”.

    - LAB MEDIA: Figure 3. Video Editor: Still show only Figure 3C. Emphasize the point on the curve labeled “PoS”.

    - LAB MEDIA: Figure 3. Video Editor: Still show only Figure 3C. Emphasize the area under the curve (the word “Area” denotes this space).

    - LAB MEDIA: Figure 3. Video Editor: Still show only Figure 3C. Emphasize the space denoted by the two-headed-arrow labelled “Recovery”.

  - Representative curves of red blood cells from healthy controls, patients with HbS traits, and a homozygous sickle cell disease patient is shown here [1].

    - LAB MEDIA: Figure 4. Video Editor: Emphasize Figure 4A.

  - The clear differences in the representative curves of homozygous sickle cell patients treated with hydroxyurea and transfusion highlights the usefulness of this assay [1].

    - LAB MEDIA: Figure 4. Video Editor: In Figure 4C, emphasize the curve for HbSS HU. In Figure 4D, emphasize the curve for HbSS Transfusion.

# Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Minke Rab: There are numerous factors that influence oxygen gradient ektacytometry, such as temperature, pH and osmolarity of the PVP, but also remnants of water and degree of mixing the blood/PVP sample can affect measurements, therefore standardized sample handling is essential [1].

    - INTERVIEW: Named author says the statement above in an interview-style statement while looking slightly off-camera.

  - Minke Rab: The outcome of oxygenscan measurements may be directly implemented in clinical practice, thus contributing to personalized medicine [1].

    - INTERVIEW: Named author says the statement above in an interview-style statement while looking slightly off-camera.

  - Minke Rab: This technique makes it possible to study sickling behavior, investigate the potential of new treatments for sickle cell disease and it can serve as a clinically relevant biomarker of disease severity [1].

    - INTERVIEW: Named author says the statement above in an interview-style statement while looking slightly off-camera.