BIOMECHANICAL TESTING OF MURINE TENDONS

Iden Kurtaliaj, Mikhail Golman, Adam C. Abraham, Stavros Thomopoulos

Response to reviewers

We would like to thank the reviewers and editor for their critical review of our manuscript. Our

responses to their specific comments are provided below.

Editorial comments:

General:

1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues.

Response: We have thoroughly proofread the manuscript to correct any spelling or grammar issues.

2. Please ensure that the manuscript is formatted according to JoVE guidelines–letter (8.5” x 11”) page size, 1-inch margins, 12 pt Calibri font throughout, all text aligned to the left margin, single spacing within paragraphs, and spaces between all paragraphs and protocol steps/substeps.

Response: We have carefully edited and formatted the manuscript according to JoVE guidelines.

3. JoVE cannot publish manuscripts containing commercial language. This includes trademark symbols (™), registered symbols (®), and company names before an instrument or reagent. Please limit the use of commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials and Reagents. For example: SkyScan, CTAn, SolidWorks, VeroWhitePlus, Kimwipe, ElectroForce, Eppendorf.

Response: We apologize for the use of commercial language in the manuscript. We have edited the text to read:

“1.1.2 Perform a microcomputed tomography scan of the entire bone, e.g., scan the humerus and calcaneus samples. NOTE: Depending on the scanner used, the settings will be different. For the scanner used in the current study (see Table of Materials), the recommended settings are: scan at an energy of 55 kVP, Al 0.25 filter, at a resolution of 6 μm.”

“1.1.3 Reconstruct microcomputed tomography scan projection images into cross-section images. Use recommended parameters for the experimenter’s scanner/software combination. NOTE: For the program used in the current study (see Table of Materials) it is recommended to use the following reconstruction parameters: Smoothing: 0-2, Beam Hardening Correction: 45, Ring Artefact Reduction: 4-9 and to reconstruct slices in 16-bit TIFF format.”

“1.1.4 Create a 3D surface model and save into a standard STL format compatible with most 3D printers and rapid prototyping. NOTE: For the program used in the current study (see Table of Materials) it is recommended to:”

“1.1.6.1 Manipulate the mesh to reduce the size of the STL file and make it compatible with any solid modeling computer-aided design program. NOTE: For the program used in the current study (see Table of Materials), it is recommended to follow the steps below:

1.1.6.1.1 Choose “Refine Mesh”, from the toolset Utilities.

1.1.6.1.2 Select the mesh object to edit.

1.1.6.1.3 Reduce mesh by a factor of 0.95 to lower the end size of the mesh object. NOTE: the “Remesh” tool could also be used to reduce the size of the mesh.

1.1.6.1.4 Resave the newly reduced file in STL format by choosing “Export as…”.

“1.2.1.1 Use a solid modeling computer-aided design program to create a custom-fit model of humerus gripping fixture (Fig. 1). NOTE: The program used in the current study is listed in the Table of Materials.”

“NOTE: The printer used in the current study uses PolyJet technology (see Table of Materials). It prints a 16µm layer of liquid ultraviolet curable photopolymer and then cures it by ultraviolet laser to solidify the pattern. The material used in the current study has a modulus of 2.5 GPa.”

“2.1.2 Determine the cross-sectional area of the tendon using microcomputed tomography. NOTE: For the scanner used in the current study (see Table of Materials), the recommended settings are: scan at an energy of 55 kVP, Al 0.25 filter, at a resolution of 5 μm.”

“2.1.5 Insert and glue tendon between a folded thin tissue paper (2 x 1 cm) and clamp the construct using thin film grips.”

“2.2.1 Perform tensile mechanical test on a materials testing frame. NOTE: For the testing frame used in the current study (see Table of Materials), the recommended protocol is:”

“2.2.9 After completion of tensile testing, perform a microcomputed tomography scan. NOTE: For the scanner used in the current study (see Table of Materials), the recommended settings are: scan at an energy of 55 kVP, Al 0.25 filter, at a resolution of 12.3 μm.”

“2.2.10 Repeat step 1.1.3.

2.2.10 Use a 3D visualization program compatible with your scanner to create a volume rendered 3D model of the scanned object. NOTE: The program used in the current study is listed in the Table of Materials.

2.2.11 Determine failure mode and failure site area by inspecting the 3D object.”

“First, step 1.1.2 is necessary to create a 3D model of the desired bone; however, due to the typically high resolution used for this scan, the file size may be too large to use with solid modeling programs. The software used in this protocol successfully reduced the size of the file (step 1.1.8.2) and preserved object geometry, although other softwares may also be effective to achieve this.”

“To measure the cross-sectional area of the supraspinatus tendon, we recommend microcomputed tomography scans of the bone-tendon-muscle specimen suspended in a cryotube with a flat bottom, where the bone is held upside down in the tube with agarose.”

“For the Achilles tendon, high resolution microcomputed tomography scans reveal two distinct tissues when examined: the tendon proper and the surrounding sheath, which appears as a lighter shade.”

Protocol:

1. There is a 10 page limit for the Protocol, but there is a 2.75 page limit for filmable content. If revisions cause the highlighted portion to be more than 2.75 pages, please highlight 2.75 pages or less of the Protocol (including headers and spacing) that identifies the essential steps of the protocol for the video, i.e., the steps that should be visualized to tell the most cohesive story of the Protocol.

Response: We have highlighted 2.75 pages of our protocol for filmable content, identifying the essential steps.

2. Please add more details to your protocol steps. Please ensure you answer the “how” question, i.e., how is the step performed? Alternatively, add references to published material specifying how to perform the protocol action. If revisions cause a step to have more than 2-3 actions and 4 sentences per step, please split into separate steps or substeps.

Response: We edited protocol steps 1.1.1, 2.1.1, 1.1.2-1.1.6, 1.2.1, and 2.2.2-2.2.13, and added references to steps 2.2.6-2.2.7, 2.2.13.

Specific Protocol steps:

1. 1.1.1, 2.1.1: Please provide details for how these bones (and tendons) are obtained and dissected.

Response: We thank the editors for the suggestion. We now provide a more detailed description of how to isolate muscle - tendon - bone specimens for testing (specifically, supraspinatus muscle - tendon - humerus bone and gastrocnemius muscle - Achilles tendon - calcaneus bone). The protocol was edited to read as follows:

“1.1.1 Dissect the bone of interest in preparation for 3D model creation and 3D bone grip printing; the humerus and the calcaneus are used as examples in the current protocol. NOTE: Detailed instructions to dissect bone-tendon-muscle specimens for mechanical testing are provided in step 2.1.1. The following steps should be followed to isolate bones for the purpose of creating 3D-printed bone grips.

      - Dissection of the humerus: after sacrifice, remove upper extremity skin, remove all muscles over the humerus, disarticulate the elbow and glenohumeral joints, and carefully remove all connective tissues attached to the humerus.

      - Dissection of the calcaneus: after sacrifice, remove lower extremity skin, disarticulate the Achilles tendon-calcaneus joint and joints between calcaneus and other foot bones, and carefully remove all connective tissues attached to the calcaneus.”

2.1.1 Dissect the muscle-tendon-bone of interest in preparation for tensile mechanical testing. In the current study, supraspinatus muscle - tendon - humerus bone specimens (N=10, 5 male, 5 female) and gastrocnemius muscle - Achilles tendon-calcaneus bone specimens (N=12, 6 male, 6 female) were isolated from 8 week old C57BL/6J mice.

2.1.1.1 Dissection of the supraspinatus muscle - tendon - humerus bone specimen

2.1.1.1.1 After sacrifice, position the mouse in a prone position.

2.1.1.1.2 Make an incision in the skin from above the elbow of the forepaw towards the shoulder.

2.1.1.1.3 Carefully remove the skin with blunt dissection so that the musculature of the shoulder is visible.

2.1.1.1.4 Remove the tissue surrounding the humerus until the bone is exposed and can be held securely with forceps.

2.1.1.1.5 Hold the humerus with forceps and carefully remove the deltoid and trapezius muscles to expose the coracoacromial arch.

2.1.1.1.6 Identify the acromioclavicular joint and carefully separate the clavicle from the acromion with a scalpel blade.

2.1.1.1.7 Taking care not to damage the supraspinatus tendon and its bony attachment, remove the muscle from its scapular attachment using a scalpel blade.

2.1.1.1.8 Taking care not to damage the supraspinatus tendon and its bony attachment, detach the humeral head from the glenoid; using a scalpel blade, lacerate the joint capsule and the infraspinatus, subscapularis, and teres minor tendons.

2.1.1.1.9 Disarticulate the elbow joint to separate the humerus from the ulna and radius.

2.1.1.1.10 Isolate the humerus - supraspinatus tendon - muscle specimen and clean off excess soft tissues on the humerus and humeral head.

2.1.1.2 Dissection of the Achilles tendon - calcaneus bone sample

2.1.1.2.1 After sacrifice, position the mouse in a supine position.

2.1.1.2.2 Taking care not to damage the Achilles tendon and its bony attachment, remove the skin with blunt dissection so that the musculature around the ankle and knee joints is exposed.

2.1.1.2.3 Using a scalpel blade, starting at the Achilles tendon - calcaneus attachment, carefully detach the gastrocnemius muscle from its proximal attachments.

2.1.1.2.4 Carefully disarticulate the calcaneus from the various adjacent bones.

2.1.1.2.5 Isolate the Achilles tendon - calcaneus specimen and clean off excess soft tissues.

2. 1.1.2-1.17, 1.2.1, 2.2.2-2.2.9: Please provide more detailed instructions for how to carry out these software steps, in particular for any step you intend to have filmed.

Response: We agree that more detailed instructions are needed. For each of the steps mentioned, we have edited the text to read as follows:

1.1.2-1.1.7

Response:

“1.1.2 Perform a microcomputed tomography scan of the entire bone, e.g., scan the humerus and calcaneus samples. NOTE: Depending on the scanner used, the settings will be different. For the scanner used in the current study (see Table of Materials), the recommended settings are: scan at an energy of 55 kVP, Al 0.25 filter, at a resolution of 6 μm.

1.1.2.1 Mix agarose powder in ultrapure water and microwave for 1-3 min until the agarose is completely dissolved. NOTE: It is helpful to microwave for 30-45 sec, stop and swirl, and then continue towards a boil.

1.1.2.2 Fill cryotubes up to ¾ with agarose.

1.1.2.3 Let the agarose cool for about 5-10 minutes.

1.1.2.4 Insert bone into the agarose gel (this will prevent movement artifacts during scanning).

1.1.2.5 Insert cryotube with bone into the scanner. NOTE: For the scanner used in the current study, a 16-position automatic sample changer was used for all scans. This scanner can automatically select magnification according to a sample’s size and shape.

1.1.3 Reconstruct microcomputed tomography scan projection images into cross-section images. Use the recommended parameters for the particular scanner/software combination. NOTE: For the program used in the current study (see Table of Materials) it is recommended to use the following reconstruction parameters: Smoothing: 0-2, Beam Hardening Correction: 45, Ring Artefact Reduction: 4-9 and to reconstruct slices in 16-bit TIFF format.

1.1.4 Create a 3D surface model and save into a standard STL format compatible with most 3D printers and rapid prototyping. NOTE: For the program used in the current study (see Table of Materials) it is recommended to:

1.1.4.1 Open file dataset. From the menu select the command File / Open.

1.1.4.2 Open the dialog File / Preferences and select the tab “Advanced”.

1.1.4.1 Use the double-time cubes algorithm to construct the 3D models. This algorithm minimizes the number of facet triangles and provides smoother surface detail.

1.1.4.2 Use 10 as the locality parameter; this parameter defines the distance in pixels to the neighboring point used for finding the object border.

1.1.4.3 Minimize tolerance to 0.1 to decrease file size.

1.1.4.3 NOTE: After opening the dataset, the images are shown in the “Raw Images” page.

1.1.4.4 To specify the volume of interest (VOI), manually select two images to set as the top and bottom of the selected VOI range.

1.1.4.5 Move to the second page “Region of Interest”. Manually select the region of interest on a single cross section image. NOTE: The selected region will be highlighted in red (i.e. the humerus cross-sectional area).

1.1.4.6 Repeat previous step every 10-15 cross-section images.

1.1.4.7 Move to the third page “Binary Selection”.

1.1.4.8 On the histogram menu click “From dataset”. The histogram distribution of brightness from all images of the dataset will be shown.

1.1.4.9 Also on the histogram menu, click the “Create a 3D Model” file menu.

1.1.4.10 Save 3D model of bone in STL file format.”

1.2.1

Response:

“1.2.1 Supraspinatus tendon-humeral bone

1.2.1.1 Use a solid modeling computer-aided design program to create a custom-fit model of humerus gripping fixture (Fig. 1). NOTE: The program used in the current study is listed in the Table of Materials.”

1.2.1.2 Open the STL format file of the humerus bone in a solid modeling program and save as a part file. NOTE: For the software used in the current study (see Table of Materials) the 3D bone object was saved in SLDPRT format.

1.2.1.3 Open the part file and manually create three anatomically relevant planes (i.e sagittal, coronal, transverse)

1.2.1.3.1 Manually define the sagittal plane to cut through the supraspinatus tendon attachment at the greater tuberosity.

1.2.1.3.2 Ensure that the 3D block contains the sagittal plane as a plane of symmetry. To achieve this add or cut material from the block if needed. NOTE: This plane of symmetry ensures that when the specimen is inserted into the fixtures the tendon and tendon attachment are located in the central axis of the fixture.

1.2.1.4 Measure dimensions of the bone along each of the three planes (i.e., height, width, length).

1.2.1.5 Measure the dimensions of the mechanical testing grips where the 3D printed fixture will be attached.

1.2.1.6 Begin with designing a solid block part (e.g., a solid cylinder).

1.2.1.6.1 Ensure that each dimension of the block is at least 5 mm greater than the dimensions of the humerus.

1.2.1.6.1 Account for design constraints from mechanical testing grips (i.e., ensure that the 3D printed fixture can be assembled and disassembled freely in the mechanical testing grips).

1.2.1.7 Create an assembly model with two components: the solid block and either the right or left humerus bone.

1.2.1.8 Define the orientation of the bone within the block (i.e., the angle between the tendon and bone).

1.2.1.9 Ensure that the entire bone volume fits inside the block.

1.2.1.10 Create a cavity in the block using the humerus bone as the mold. NOTE: In the current study, using the software specified in the Table of Materials, it is recommended to follow the steps below:

1.2.1.10.1 Insert the design part (humerus) and the mold base (cylinder block) into an interim assembly.

1.2.1.10.2 In the assembly window, select the block, and click Edit Component on the Assembly toolbar.

1.2.1.10.3 Click Insert → Mold → Cavity.

1.2.1.10.4 Select Uniform scaling and enter 0% as the value to scale in all directions.

1.2.1.11 Cut the assembly along the sagittal plane to create two symmetrical components that fit the bone anteriorly and posteriorly (e.g., two half cylinders, as seen in Figure. 1). NOTE: Two components are designed that fit the bone anteriorly and posteriorly. The anterior component includes a half spherical-shaped cavity extended from the anterior side of the humeral head up to the supraspinatus tendon attachment. The posterior component cavity is shaped like the posterior part of the humerus (i.e., posterior side of the humeral head, deltoid tuberosity, and medial and lateral epicondyle)

1.2.1.12 Save each component as a separate file part.

1.2.1.13 For the anterior component, ensure that the humeral head is embedded in the cavity of the part by defining appropriate tolerances. NOTE: In the current study, using the software specified in the Table of Materials, it is suggested to follow the steps below:

1.2.1.13.1 Create a revolved cut to smooth the mesh geometry of the cavity.

1.2.1.13.2 Create a sketch for the cut by emulating the cavity geometry and adding a locational clearance. NOTE: The clearance allows for free assembly and disassembly between the bone and the anterior component.

1.2.1.14 Modify the posterior component to imitate the cavity geometry to create a cut that adds clearance, as described above for the anterior component.

1.2.1.15 Make a cut in the transverse plane starting from the top of the posterior component up to the crest of the greater/lesser tubercle. NOTE: As seen in Figures 1 and 2, the posterior component includes a cut that creates an opening at the tendon attachment.

1.2.1.16 Create a snug fit between the two components to allow for free assembly and disassembly. NOTE: A hole-shaft fit with a loose running clearance was created for the fixtures in the current study.

1.2.1.17 Create 3D mirror models for each component of the fixture for the opposite limb (i.e., left or right).

1.2.1.18 Add an etch on the bottom of the fixtures to distinguish between the left and right sides.

1.2.1.19 Save all fixture parts in STL standard file format in preparation for 3D printing.”

2.2.2-2.2.9

Response:

“2.2.2.     Collect load-deformation data.

2.2.3.     Calculate strain as displacement relative to the initial gauge length of the tendon.

2.2.3.1 Use calipers to measure gauge length from attachment site to tendon grips.

2.2.4. Calculate stress as force divided by initial tendon cross-sectional area (as measured from microCT).

2.2.5.     If interested in viscoelastic behavior, a stress relaxation can be performed prior to the tension test to failure and the data can be used to calculate parameters such as A, B, C, tau1, and tau2 from the quasilinear viscoelastic model24.

2.2.6.     From the load deformation curve, calculate stiffness (i.e., the slope of the linear portion of curve), maximum force, and work to yield (i.e., the area under the curve up to yield force).

2.2.6.3 Identify the linear portion by choosing a window of points in the load-deformation curve that maximizes the R2 value for a linear least squares regression25.

2.2.6.1 Determine the stiffness as the slope of the linear portion of the load-displacement curve25,26.

2.2.7.     From the stress strain curve, calculate modulus (i.e., the slope of the linear portion of curve), strength (i.e., maximum stress), and resilience (i.e., the area under the curve up to yield stress).

2.2.8    NOTE: Using the RANSAC algorithm in Matlab (The Mathworks, Natick, MA), the yield strain (x-value) is defined as the first point when the y-fit has deviated more than 0.5% of the expected stress value (y-value). Yield stress is the corresponding y-value of yield strain.

2.2.9 NOTE: In addition to the monotonic tensile loading to failure described in the current study, cyclic loading can provide important information about tendon fatigue and/or viscoelastic properties. For example, Freedman et al reported fatigue properties of murine Achilles tendons.27

2.2.10 After completion of tensile testing, perform a microcomputed tomography scan of the entire bone, e.g., scan the humerus and calcaneus samples. NOTE: For the scanner used in the current study (see Table of Materials), the recommended settings are: scan at an energy of 55 kVP, Al 0.25 filter, at a resolution of 6 μm.

2.2.10.1 Repeat steps 1.1.2.1 - 1.1.2.5

2.2.11 Repeat step 1.1.3

2.2.12 Use a 3D visualization program compatible with the scanner to create a volume-rendered 3D model of the scanned object. NOTE: The program used in the current study is listed in the Table of Materials.”

2.2.13 Determine the failure mode and failure site area by inspecting the 3D object.”

Figures:

1. Please cite Figure 4 outside of the Figure Legends section.

Response: We thank the editors for pointing out that Figure 4 was not cited outside of the Figure Legends section. We have corrected the text to read:
“Animal sex had a significant effect on the mechanical properties of the supraspinatus and Achilles tendons (Figure. 4).”

2. Figure 3A, B: Please include more information about what is shown here in the legend.

Response: We have corrected the text to read: 
“Figure 3: Comparison of previous and current methods for mechanical testing of murine supraspinatus tendons. (A) Previous specimen preparation methods used in our laboratory prior to mechanical testing: the humerus was potted in epoxy up to the humeral head to stabilize the bone, a paper clip was placed over the humeral head to prevent growth plate fracture, and, for the epoxy to cure, the specimens were left in room temperature for 4-6 hours prior to mechanical testing. (B) Specimen preparation methods used in the current study (Steps 1.2-1.3 and 2.1.4-2.1.6): Top left shows a 3D representation of the fixtures as produced by a solid modeling program. The 3D printed fixtures are reusable and easily assembled and disassembled. The bone end of the specimen is inserted into the fixtures, securing the growth plate and exposing the tendon for gripping and testing. The tendon end is glued between a folded thin tissue paper and inserted into the grips. Preparation time for each specimen is 10-15 minutes. (C) Representative load-deformation curves for tensile testing of a supraspinatus tendon using methods described in the current study. (D) Representative load-deformation curve for tensile testing of supraspinatus tendon using prior methods, showing a growth plate failure.”

References:

1. Please do not abbreviate journal titles.

Response: We have corrected the references according to JoVE formatting requirements.

Table of Materials:

1. Please ensure the Table of Materials has information on all materials and equipment used, especially those mentioned in the Protocol.

Response: We have corrected the Table of Materials to include all materials.

Reviewers' comments:

Reviewer #1:

Manuscript Summary:

A new system is developed for murine tendon mechanical testing.

Major Concerns:

None.

Response: We thank the reviewer for the careful review of our manuscript.

Minor Concerns:

It is stated (lines 54-55) that "few studies have tested the biomechanical properties of murine tendons" and that (line 280) "mechanical properties is... rarely reported". Then the authors (line 325) state they reviewed 20 studies which did so which is a reasonable body of literature (not few or rare) for something only done relatively recently.

Response: We thank the reviewer for noting this ambiguity and agree the testing of murine tendons is becoming more common. We have replaced “rarely reported” with “uncommon”. Notably, in the last 20-30 years, major advancements in tendon research have rapidly increased this body of literature. In the last decade alone, according to PubMed, there were 1595 published peer-reviewed journal articles on tendon. 823 of these studies used murine models (~52%). However, only 6% of these murine papers reported tendon mechanical properties. As discussed in the manuscript, we believe that this small percentage is due to the difficulty in gripping these small tissues, tedious and time-consuming specimen preparation methods, and artifactual growth plate fractures.

The mechanical properties described (lines ~185-191) are just the simplest. For completeness and guidance to the reader, it should be noted that fatigue testing and/or frequency sweeps can often be beneficial.

Response: This is an excellent point. We agree that cyclic loading of tendon can provide important information about tendon fatigue and/or viscoelastic properties. We have added step 2.2.9 to read:

“NOTE: In addition to the monotonic tensile loading to failure described in the current study, cyclic loading can provide important information about tendon fatigue and/or viscoelastic properties. For example, Freedman et al reported fatigue properties of murine Achilles tendons.27”

The potential upscale to larger animal models (345) should acknowledge the potential of compliance of the 3D printed fixture material (depending on material) which is likely not an issue with the low loads of murine tendons).

Response: We agree that the modulus of the 3D printed fixture material must be high enough that it is not compliant relative to the strength of the tendon being tested. Using standard plastic materials to 3D print these fixtures is cost effective and appropriate for testing both murine and rat tendons. We have not validated the fixtures for larger animals, and agree that compliance could become an issue. Testing of tendons from larger animals may require metal 3D printing, which is considerably more expensive. The edited text now reads (line 516):

“Furthermore, these methods are not limited to the supraspinatus and Achilles tendons, as they are easily adapted to testing other murine tendons and tendons from larger animal models. To test tendons from larger animals, however, the modulus of the 3D printed fixture material must be high enough that it is not compliant relative to the strength of the tendon being tested.”

Reviewer #2:

Manuscript Summary:

The manuscript describes a new way to test ligaments and tendons of murine models. The group 3D prints a sample grip using micro CT scans of the bone that the tissue is attached to. This creates more uniform gripping at the base of the sample. This mounting procedure reduces preparation time, limit sample slippage, prevent tearing at the grips, and reduce variations between samples.

Response: We thank the reviewer for the careful review of our manuscript.

Major Concerns:

Does a new grip need to be made for every sample. If you want the best possible grip it is assumed that every sample would have its own custom grip. Printing them may not be expensive, but scanning the samples for the printing requirements could get expensive.

Response: We thank the reviewer for raising these important points. Our point by point responses to the reviewer are as follows:

Does a new grip need to be made for every sample.

Response: It is not necessary to make a new grip for every sample. The grips are reusable (there is no glue or epoxy in the fixture; the bone is held in a press fit). Furthermore, small variations from sample to sample do not affect the effectiveness of the fixtures. For the design of the fixtures in the current study, we scanned each bone once (i.e., left humerus, right humerus, and calcaneus) and created one 3D model for each bone. We then used the three fixtures to conduct all mechanical testing for this manuscript.

If you want the best possible grip it is assumed that every sample would have its own custom grip.

Response: For animals of the same age, the bone geometry is nearly identical, thus the same fixture can be used for testing of all specimens. In this manuscript, we used 8-week old mice (skeletally mature adult mice) and 3D printed fixtures specific to this age group to test tendons. It was not necessary to create separate male and female fixtures. For other age groups (e.g., 4-week mice) or mice with unique bone phenotypes it is recommended that fixtures that fit the particular geometries of the bones are manufactured.

Printing them may not be expensive, but scanning the samples for the printing requirements could get expensive.

Response: As the fixtures are reusable, only one scan is needed for any particular group. Additional fixtures can be printed based on the scan, as needed.

We have clarified the above points in the revised paper (steps: 2.1.4, 475-483).

Minor Concerns:

Perhaps the protocol would be better suited for a chart?

Response: Figure 3 summarizes the protocol for specimen preparation prior to mechanical testing, comparing previous methods to the new method. Although this is not a chart, we now provide more detail in the figure’s legend. For the detailed protocol, we followed the required format of JoVE, and note that the final publication will also include the key parts of the protocol in an instructional video.

How were the COV and average calculated?

Response: We summarized mean +/- standard deviation for each of the mechanical properties (i.e., maximum force, stiffness, maximum stress, and modulus), as reported from all cited studies (Tables 1 and 2). COV was calculated by dividing the standard deviation by the mean. The mean of the COVs was then determined to provide a measure of typical variation in published methods.

It could be said that the first 3 studies are driving up the average COV for the other studies.

Response: Thank you for identifying this. We conducted a literature research and included all studies that determined the mechanical properties of the murine supraspinatus tendon. Although the first three studies have higher COVs that some of the other reports, we note that they were conducted by three different laboratories. Furthermore, all COVs are included in the table, so the reader can compare the current results to individual prior studies as well as to the mean COV from prior studies.

Would it further improve the outcome if the bone was epoxied into the printed part or if it was roughened on the inside?

Response: We thank the reviewer for the suggestion. Because the fixtures are printed at a relatively high resolution accurately reflecting the bone geometry, epoxy is not necessary. Use of epoxy may indeed marginally improve fixation, but this would be at the cost of substantial additional preparation time and the inability to reuse the fixture. Roughening the inner surface of the fixture may also improve fixation, but would be difficult to implement and the expected benefit would be marginal.

This work seems like a methods paper.

Response:  Although we present data on the effect of sex on murine tendon mechanical properties, the paper is indeed primarily a methods paper. This is consistent with the goals of JoVE: “JoVE publishes the leading peer-reviewed, PubMed-indexed video methods journal. Articles consist of high-quality video demonstrations and detailed text protocols which facilitate scientific reproducibility and productivity. The scope of the journal includes novel techniques, innovative applications of existing techniques, and gold standard protocols in the physical and life sciences.”

It would be more convincing if the group tested samples using the epoxy method as well as the custom grip method that they designed.

Response: We thank the reviewer for bringing up this valuable point. The last three studies from our laboratory (Table 1) used the epoxy method to determine the mechanical properties of murine supraspinatus tendons. Other groups use similar approaches and none of them use 3D printed fixtures. Therefore, a direct comparison can be made with the results using the current methods and the prior methods from our group and others, as presented in Table 1.