Dear Dr. Dsouza, Senior Review Editor, JoVE,

Thank you for the time and effort spent evaluating the manuscript we submitted to JoVE entitled: “Manufacturing and Calibrating Soil Surface Temperature and Moisture Sensors”. We appreciate the opportunity to revise our paper to address the constructive comments and suggestions from the reviewers. Below we respond, first to the editorial comments and second to the comments from each peer-reviewer. Our responses are in bold font following each comment and we include the line numbers where changes can be found in the manuscript.

We feel the manuscript is improved following revision and each suggestion and comment has been addressed. We hope that this version will be suitable for publication in JoVE.

Sincerely,

Armin Howell, on behalf of the co-authors.

Editorial Comments:

Protocol Numbering: Please adjust the numbering of your protocol section to follow JoVE’s instructions for authors, 1. should be followed by 1.1. and then 1.1.1. if necessary and all steps should be lined up at the left margin with no indentations. There must also be a one-line space between each protocol step.

-We have updated the numbering in the protocol to properly reflect JoVE’s instructions for authors. We have also placed a one-line space between each step.

Protocol Highlight: After you have made all of the recommended changes to your protocol (listed above), please re-evaluate the length of your protocol section. Please highlight ~2.5 pages or less of text (which includes headings and spaces) in yellow, to identify which steps should be visualized to tell the most cohesive story of your protocol steps.
1) The highlighted steps should form a cohesive narrative, that is, there must be a logical flow from one highlighted step to the next.
2) Please highlight complete sentences (not parts of sentences). Include sub-headings and spaces when calculating the final highlighted length.
3) Notes cannot be filmed and should be excluded from highlighting.

-We have highlighted 2.5 pages of the protocol that tells the most cohesive story.

Discussion: JoVE articles are focused on the methods and the protocol, thus the discussion should be similarly focused. Please ensure that the discussion covers the following in detail and in paragraph form (3-6 paragraphs): 1) modifications and troubleshooting, 2) limitations of the technique, 3) significance with respect to existing methods, 4) future applications and 5) critical steps within the protocol.

-We believe our manuscript addresses all sections: modification options, limitations of the sensors, why the methods are significant, critical considerations when using the sensors, and potential future applications of the sensors in the Discussions section. The text currently fits within the 3-6 paragraph length constraints. However, please let us know if you think any changes are necessary.

Figures: Please remove the figure/table legends from the figure files and place them directly below the Representative Results text.

-We removed all figure legends from the figure pdf files and placed them below the Representative Results text.

References:
1) Line 453, 486: Please use superscripted citation format

-We updated the citation format in line 453 (now line 639) and line 486 (now line 678) to be superscripted.

If your figures and tables are original and not published previously or you have already obtained figure permissions, please ignore this comment. If you are re-using figures from a previous publication, you must obtain explicit permission to re-use the figure from the previous publisher (this can be in the form of a letter from an editor or a link to the editorial policies that allows you to re-publish the figure). Please upload the text of the re-print permission (may be copied and pasted from an email/website) as a Word document to the Editorial Manager site in the "Supplemental files (as requested by JoVE)" section. Please also cite the figure appropriately in the figure legend, i.e. "This figure has been modified from [citation]."

-All figures and tables in this manuscript are original and not previously published.

Comments from Peer-Reviewers:

Reviewer #1:
Manuscript Summary:
The paper describes efficiently the procedure to build and calibrate a low-cost resistivity soil moisture sensor to be used in shallow depths. Many experimental sensors under this technology can be found in the bibliography. However, novelty is not a constraining factor in this journal, so that the technology limitations should be also accepted under this consideration.
The methods described are enough accurate to reproduce the experiments, but some theoretical concepts should be considered in the manuscript (Minor Concerns)

Minor Concerns:
1) Sometimes in the manuscript (Lines 429 or 442 for instance) the sensor is considered as a "capacitance sensor". The sensor proposed in this study is a resistivity sensor, since the measurements are performed from the Wenner method. Capacitance or capacitive sensors are based on the measurement of the dielectric constant of the media.

-Thank you for this comment. You are correct that we incorrectly referred to the sensors as “capacitance sensors” and we have changed each occurrence of this to properly reflect that we are describing resistivity sensors. Now corrected on lines 21, 559, 574, 647, 682.

2) Sometimes in the manuscript, as in Line 393 among others, Siemens are considered to be a unit to express electrical resistance. This is not a purely correct assumption. Siemens is the unit of the Conductance, that is the inverse magnitude of Resistance, so try to change the different sentences to express it in a scientific correct way.

-Thank you for this clarification. We updated all the references that incorrectly described Siemens to properly show that Siemens are a measure of conductance. Lines 523-524, 526, 535, 572, 606, 609, 610, 612, 684, 686.

3) Line 405 resistance is not the magnitude presented in Figure 6. It is Conductance.

-Thank you again for catching this misunderstanding. We updated the line you referred to as well as the caption for the figure referenced in this line. The new text can be found on lines 535 and 606.

4) Line 429 the influence of temperature in water content measurement is discussed, but it is not marked that the soil resistance is temperature dependent, so that lower water content measurement registered under warm conditions are not only due to a quicker drying but as a consequence of calibration temperature dependence as well. In other words, there would be a different calibration curve for each soil at every temperature.

-This is a very good point and thank you for ensuring that we note that temperature changes affect the sensors’ measurement of conductance. We now make this point in lines 559-561, saying: “It should be noted that the electrical conductivity of soils is influenced by soil temperature as increases in temperature lead to higher conductivity of soils15.” Also, we now describe in the method section future plans for conducting temperature calibrations to determine temperature effects on substrate conductivity. We do not believe that the temperature effect on soil moisture values from warmed-plot soil moisture data explains the soil moisture patterns observed, as at the temperatures we measured we would expect the increased soil temperature to make the soils look wetter, not drier. However, we agree temperature effects are important to consider and discuss and we now do so.

5) Line 481 Sensor salinity influence is highlighted. This is a very well-known drawback of resistivity sensors for water content measurement. A citation reporting this problem could be added to emphasise this condition.

-Thank you for this suggestion. We used Rhoades and Ingvalson’s manuscript: “Determining Salinity in Field Soils with Soil Resistance Measurements” to demonstrate how salinity and resistance measurements are related.

Reviewer #2:
The manuscript described methods for manufacturing soil surface sensors that simultaneously measure temperature and moisture of the top 5 mm. This is a very well written and structure paper. I have only minor comments:
1. The protocol must be a numbered list: step 1 followed by 1.1, followed by 1.1.1, etc.

-We are grateful for the Reviewer’s supportive comments and suggestions for improvement. We have updated the protocol numbering to properly reflect JoVE’s requirements.

In line 189 it is said that from step 1.2.6, where？

-We hope the updated protocol numbering will make these step reference easier to find and understand. In Step 1.5.3., we wanted to make it clear that the reader is using the 0.13 inch piece of heat shrink tube that was cut and placed on to the cable jacket in step 1.2.6. Prior to our protocol numbering update these step numbers may have been more difficult navigate and hopefully are much clearer now, but please let us know if we have missed the point of this suggestion and other changes are being proposed. Step 1.5.3 can be found in lines 217-221 and step 1.2.6 can be found in lines 158-161.

In line 365 it is said that from step 4.4.25, where？

-Here too, we are referencing prior steps to clarify linked steps for the reader. Specifically, in step 4.5.1 we are instructing readers to subtract the dry calibration container weights determined in step 4.4.1 from the weight of the dry calibration container with substrate weight determined in step 4.4.34. Step 4.4.1 can be found in lines 396-398, step 4.4.34 in line 483, and step 4.5.1 can be found in lines 487-489.

In line 368 it is said that from step 4.5.1, where？

-Here too, we are referencing prior steps to clarify linked steps for the reader. Specifically, in step 4.5.3 we are instructing readers to divide the water weights determined in step 4.5.2 by the soil weight determined in step 4.5.1. These steps can be found lines 487-496.

In line 370 it is said that from step 4.5.3, where？

-As above, we are referencing prior steps to clarify linked steps for the reader. Specifically, in step 4.5.4 we are instructing readers to match the measurement times to the gravimetric water content measurements determined in step 4.5.3. These steps can be found in lines 495-499.

2.Description of control circuit can be added.

-Due to article length constraints and because the datalogger and multiplexer were purchased and are not part of the design described here, we did not include detailed information about the control circuit or other aspects of those purchased parts of the system. However, the part numbers and descriptions can be found in the supplemental Materials List spreadsheet. Based on the Reviewer comment, we have now also added the following text to line 243-245 of the manuscript: “At each measurement time, the datalogger opens ports to the multiplexer which then acts as a relay and allows current to run to the resistivity sensor.”

3.Software systems should also be introduced

- Per the journal’s instructions, we did not include any company information in the manuscript. The two types of software used were (1) Campbell software that programs the Campbell datalogger and (2) excel software to enter, organize, and assess the calibration curves. We are happy to include this information if the Editor wishes.

Reviewer #3:
Manuscript Summary:
and major concerns
The paper describes a soil temperature and moisture sensor. It targets at the .5cm surface of soil and seems that the technology is useful for desert area. It may be a good technology article and well written. However, it lacks scientific contribution. It has a lot detailed making procedure which may not be valuable to general readers. The technology may be suitable for patent application. For this reason, I am not fully convinced by the paper except the journal is focused on the techniques.

We appreciate this comment, but respectfully disagree that the paper lacks scientific contribution. We do agree that the format is elaborate, but this is the format as outlined by the journal and it allows for a reader to build the probes themselves, which is a key goal of this paper. Based on this comment, we have added text to further contextualize the utility of the probes, including outside of desert environments, and we have read through to ensure that the scientific contribution is clear. We feel this manuscript is well aligned with the journal’s goals and instructions for formatting, and thus hope the reviewer will agree this paper will make a strong contribution to the literature.