## Table 2
Problem | Possible reason | Solution
The cardiomyocyte detaches during maximal activation | Insufficient gluing time; The glue is old and has dried | Increase the time of the gluing step
| There is Triton® in the cell suspension solution, which can no longer be removed | Repeat the extraction procedure with one or two additional Triton® wash out steps
The cardiomyocyte has low force under control conditions | The extraction went wrong and delivered low-quality cells | Increase the sample size and do a new extraction. If the problem persists is probably due to improper sample collection - discard this sample
The cell is visibly contracting but no force is recorded; The cell has unusual force values | The force transducer is off | Turn it on
| The force transducer is not well calibrated | Calibrate the force transducer using a set of known weights (check the manufacturer’s instruction manual).
| The force transducer needle is loose | Glue the needle again using crystal bond 509 or jewelers wax.
The striation pattern is not good enough to determine the sarcomere length | Insufficient light | Increase microscope light or move the cell back to the coverslip and assess sarcomere length again (the wells has lower light intensity)
| The extraction went wrong and delivered low-quality cells | Increase the sample size and do a new extraction
| Needles’ tips are not in the same plane | Using micromanipulators, adjust the needles’ tips up or down until finding a focused sarcomeres
No length and/or force variation during acquisition | The motor or the force transducer are off | Turn them on
| The motor is broken and not producing cell shortening | Replace it or try to calibrate it using a function generator
Too much noise on the acquisition recordings | Too much air flow around the equipment | Protect the equipment from the direct air flow
| Too many vibrations around the equipment | A stabilization table is advisable. Even then, it is recommended to remove any equipment that might have a compressor or emit vibrations (freezer, fridges)
Ca2+-sensitivity curve has strange values and the force values do not increase with [Ca2+]. | The mixture of activating and relaxing solution was not done properly (check 3.10 to 3.14 of the methods section, possibly due to insufficient mixing) | Defrost the vials with the same concentration, collect all vial’s content in the same beaker, mix with a stirrer and divide them again. Test these solutions again in a new cell. If this does solve the problem, prepare a new batch of Ca2+-containing solutions