# Submission ID #: 60584

# Scriptwriter Name: Qingyun Ping

# Project Page Link: https://www.jove.com/account/file-uploader?src=18478168

# Title: Sample Preparation in Quartz Crystal Microbalance Measurements of Protein Adsorption and Polymer Mechanics

# Authors and Affiliations:  Gwen E. dePolo1*, Emily Schafer2*, Kazi Sadman1, Jonathan Rivnay2, Kenneth R. Shull1

1Department of Materials Science and Engineering, Northwestern University, Evanston, IL, USA

2Department of Biomedical Engineering, Northwestern University, Evanston, IL, USA

*These authors contributed equally.

# Corresponding Author:

# Kenneth Shull

# k-shull@northwestern.edu

# Email addresses for Co-authors:

# gwendepolo2023@u.northwestern.edu

# emilyschafer2023@u.northwestern.edu

# kazisadman2014@u.northwestern.edu

# jrivnay@northwestern.edu

Author Questionnaire:

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique? No

2. Does your protocol include software usage? Yes

If yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps.

3. Which steps from the protocol section below are the most important for viewers to see? Please list 4-6 individual steps using the step numbers listed in this document. This information is important to prepare your Videographer for your shoot. (You do not need to include steps that will be screen captured. Please do not list entire sections.)

2.3, 2.6, 3.4, 3.5, 3.7

4. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1-2 individual steps using the step numbers listed in this document. (Please do not list entire sections.)

Step 3.7 is a difficult step to coordinate. Practicing moving a sensor from the spin coater to the potassium bromide solution can help with this issue.

5. Will the filming need to take place in multiple locations? Yes. The labs are on the third floor of one building and the second floor of a second building. The labs are a five minute walk apart and the buildings are connected by a walkway.

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

Authors, these headshots will be used for the JoVE Dedicated Author Webpage. Here is one example if you wish to take a look.

- REQUIRED Interview Statements: (Said by you on camera)  - All interview statements may be edited for length and clarity.

  - Gwen dePolo: The quartz crystal microbalance is useful for sensing small concentrations of another compound and can be applied to questions regarding the mechanics of soft matter and biomaterial systems [1].

    - INTERVIEW: Named author says the statement above in an interview-style shot while looking slightly off-camera.

  - Gwen dePolo: The main advantage is that the quartz crystal microbalance is able to obtain highly accurate information for small sample sizes [1].

    - INTERVIEW: Named author says the statement above in an interview-style shot while looking slightly off-camera.

OPTIONAL Interview Statements: (Said by you on camera)  - All interview statements may be edited for length and clarity.

  - Emily Schafer:  Studying the properties of model biopolymer films can be extremely useful in understanding biological processes, such as enzymatic degradation pathways and kinetics in complex tissues [1].

    - INTERVIEW: Named author says the statement above in an interview-style shot while looking slightly off-camera.

  - Gwen dePolo: The technique is well suited for studying a wide range of materials systems and for understanding how mechanical properties of polymeric materials respond to the environment [1].

    - INTERVIEW: Named author says the statement above in an interview-style shot while looking slightly off-camera.

  - Emily Schafer: When extracting mechanical property information, it is necessary to work with films of an appropriate thickness [1]. Researchers must be careful to apply the correct modeling for the film during analysis [2-LM].

    - INTERVIEW: Named author says the statement above in an interview-style shot while looking slightly off-camera.

    - LAB MEDIA: Figure 1

# Section - Protocol

- QCM-D (Quartz Crystal Microbalance with Dissipation) Collagen Adsorption: Data Acquisition

  - After turning on all necessary equipment, remove the flow module from the chamber platform [1] and unscrew the large thumb screws to open the module [2].

    - Talent removes the flow module.

    - CU: Talent unscrews the screws.

  - Dry the O-ring on the flow module with a stream of nitrogen gas, and check the O-ring is lying flat [1]. Mount the sensor on the O-ring by placing the sensor with the active surface side down and anchor-shaped electrode oriented toward the marker in the flow module [2].

    - Talent applies nitrogen gas to dry the O-ring and flow module.

    - Talent mounts the sensor.

  - Find the initial resonance frequencies of the sensor [1]. Place the inlet pump tubing in the 1x PBS [2]. Start the external pump flow at 25 microliters per minute [3] and visually inspect the tubing to be sure that the fluid is flowing through the tube. Allow the fluid flow for at least 15 minutes to properly equilibrate [4].

    - Talent operates on the computer to find the frequencies. Author NOTE: We took shots of this step but can also attach a screen capture. For the screen capture (which will be named 60584_screenshot_pf1), start at 0:05.

    - Talent places the inlet pump tubing in the buffer.

    - Talent starts the pump flow. Show the flow rate setting. Author NOTE: The firsts two takes of this step were using the wrong flow rate (50 microliters/minute). The correct flow rate is 25 microliters/minute.

    - Shot of the tubing showing fluid is flowing through. Important Step

  - In the software, press “Start Measurement” to start the measurement to begin data acquisition. Monitor the frequency and dissipation values for at least 5 minutes to ensure a stable baseline [1].

    - SCREEN: Talent starts the measurement. 60584_screenshot_1; 0:00-0:15

  - Stop the pump and move the inlet tubing to the collagen-acetate buffer solution, and resume fluid flow [1]. Note the time of this event for later analysis [2].

    - Talent stops the pump and moves the inlet tubing to another buffer, and then, starts the pump.

    - Talent records the time.

  - Allow the new frequency and dissipation values to equilibrate [1] for 8-12 hours to a stable value [2]. Then, stop the pump, move the inlet tubing back to the 1x PBS solution, and resume fluid flow [3]. Note the time of this event for later analysis [4].

    - SCREEN: Frequency and dissipation values changing. 60584_screenshot_2; 0:00-0:20

    - SCREEN: Show a stable value. Important Step 60584_screenshot_3; 0:00-0:10

    - Talent stops the pump, moves the tubing into PBS buffer, and restarts the pump. Important Step

    - Talent records the time.

  - Again, allow the new frequency and dissipation values to equilibrate [1] for 30 minutes to a stable value [2]. End the data acquisition of the measurement and save the data [3].

    - SCREEN: Frequency and dissipation values changing. 60584_screenshot_4; 0:00-0:20

    - SCREEN: Show a stable value. 60584_screenshot_5; 0:00-0:25

    - SCREEN: Talent ends data acquisition, and saves. 60584_screenshot_6; 0:00-0:23

- QCM Polyelectrolyte Complex Swelling: Sample Preparation

  - First, position a bare quartz crystal sensor in a sample holder connected to the vector network analyzer and computer [1]. Turn on the analyzer to apply an oscillating voltage to the sensor [2], and collect a reference conductance spectrum for the sensor in air [3].

    - Talent positions a sensor.

    - Talent turns on the analyzer.

    - Talent shows the spectrum. Author NOTE: This particular step was interacting with the data collection program. We tried some wide shots (where you could see the sample and the screen) as well as some videos of the screen. I also have screen captures of this step that can be included.

  - Then, submerge the sample holder in a lipless 100-milliliter beaker filled with distilled water and collect a reference conductance spectrum for the bare sensor in water [1].

    - A) Talent submerges the sample holder in water,

    - B) collects spectrum. Author NOTE: This step became too difficult to film smoothly, so we labeled the steps as 3.2.1.A and 3.2.1.B. There are also screen captures for step 3.2.1.B if this would be better.

  - Insert a small silicon wafer into a 0.5 molar potassium bromide solution at an angle to create a slide for the quartz sensor during the annealing step to prevent the film from coming off the sensor [1].

    - Talent inserts a silicon wafer into the solution at an angle.

  - After cleaning the sensor surface, add the PEC (pronounce as P-E-C) to the surface of the sensor [1-TXT]. If the complex is in two phases, polymer rich and polymer poor [2-TXT], slowly insert the pipet into the solution [3]. Evacuate the pipet by blowing bubbles while moving the pipet into the denser polymer rich phase [4].

    - Talent adds solution to the surface of the sensor. TEXT: PEC is PSS (polystyrene sulfonate) : PDADMA poly(diallyldimethylammonium). Author NOTE: This step was not recorded. Please include the text with the shot of 3.4.2, which was filmed in such a way to accommodate having text next to the vials.

    - CU: Shot of the two phases. Important Step TEXT: PEC is PSS (polystyrene sulfonate) : PDADMA poly(diallyldimethylammonium).

    - CU: Talent inserts the pipet into the solution.

    - CU: Talent blows bubbles and moves the pipet into the denser phase. Important Step

  - After releasing a couple bubbles in the polymer rich phase, draw up 0.5-0.75 milliliters of the polymer rich solution into the pipet [1]. Maintaining pressure on the pipet bulb to not allow the polymer poor phase to enter the pipet, draw the pipet out of the solution [2].

    - Talent draws up solution. Important Step

    - CU: Focus on the talent’s hand.

  - Wipe the outside of the pipet using a Kimwipe [1]. Add enough solution dropwise onto the surface of the quartz sensor to completely cover the surface. Make sure there are no visible bubbles in the solution on the sensor surface [2].

    - Talent wipes the pipet.

    - CU: Talent adds solution onto the surface.

  - Spin coat the PEC sample [1] and immediately submerge the sensor in the 0.5 molar potassium bromide solution to prevent salt crystallization on the film. Allow the film to anneal for at least 12 hours [2]. NOTE: 3.7.1. and 3.7.2 combined into one shot.

    - Talent spin coats the sample. Important Step

# Talent submerges the sensor in solution. Important Step

- QCM Polyelectrolyte Complex Swelling: Measurement of the Film in Air and Water

  - Transfer the sensor to a beaker filled with distilled water to remove the excess potassium bromide from the film and the back side of the sensor. Leave the sensor in the solution for 30-60 minutes [1].

    - Talent transfers the sensor to a beaker filled with water.

  - Take a measurement of the film applied to the surface of the sensor in air [1]. Reference to the bare sensor in air. Allow the film data to equilibrate [2].

    - Talent positions the sensor

# 4.2.1.b. Talent turns on the analyzer. NOTE: 4.2.1. was split into 4.2.1.a and 4.2.1.b

    - SCREEN: Film data changing, with comparison to the reference. 60584_screenshot_6; 0:00-0:22, 3:35-4:42 Video editor: This portion can be shortened by cutting down the waiting times between data collection.

  - Next, insert dried calcium sulfate into a 100-milliliter lipless beaker [1] and measure the completely dry film thickness [2]. Remove calcium sulfate from the beaker and rinse the beaker with distilled water [3].

    - Talent inserts dry film into a beaker.

    - SCREEN: Talent shows the thickness in the analysis system. 60584_screenshot_8; 00:00-00:25

    - Talent takes out the dry film, and rinses with water.

  - Fill the 100-milliliter lipless beaker with 30 milliliters of distilled water. Insert a stir bar to ensure the water is circulating around the film [1]. With reference to the bare sensor in water, measure the film in water [2] for about 30-45 minutes or until the film data are equilibrated [3].

    - Talent fills the beaker with water, and adds a stir bar.

    - Shot of the sensor in water.

    - SCREEN: Film data stops changing. 60584_screenshot_9; Video editor: One of the first stills from this clip will show the data not changing.

  - Prepare a 15 milliliters solution of 5 molar potassium bromide by measuring 5.35 grams of potassium bromide into a graduated cylinder and fill to 15 milliliters with distilled water. Swirl [1] until dissolved [2].

    - Talent adds compound and water into a cylinder, and swirls.

    - Talent shows the compound dissolved.

  - Face the film away from where the potassium bromide solution is being added to the water so that the film does not dissolve [1]. Add the potassium bromide solution to the beaker with distilled water [3] in 0.1 molar increments [2-LM].

    - Talent faces the film away from where the solution is being added.

    - LAB MEDIA: Table 1

    - Talent adds solution into the beaker.

  - Make sure the system has equilibrated [1] before adding another addition [2] of the potassium bromide solution [3].

    - SCREEN: A screenshot of an equilibrated segment. 60584_screenshot_10.png

    - Talent adds more solution into the beaker.

    - SCREEN: The system shows a change in frequency and creates a new step in the frequency data. 60584_screenshot_10.png

  - After all the data have been acquired, remove the film from the holder and place in a beaker of distilled water. Allow the salt to leave the film for 30-60 minutes [1] and air dry the film [2].

    - Talent transfers the film from the holder into a beaker.

    - Talent takes out the film and places it on a surface.

# Section – Results

- Results: Collagen Adsorption

  - In this study, the introduction of collagen solution caused protein adsorption to begin, which was observed as a steady decrease in frequency and increase in dissipation over time [1], until the density of adhered collagen plateaus at a stable baseline [2].

    - LAB MEDIA: Figure 3A&B – Video editor: Emphasize the decrease trend in Figure 3A, and increase trend in Figure 3B.

    - LAB MEDIA: Figure 3A&B – Video editor: Emphasize the stable baseline.

  - The frequency changes are correlated with the mass of the film [1] and the dissipation is correlated with energy dissipation from the film [2]. A theoretical model is needed to obtain the viscoelastic properties quantitatively [3].

    - LAB MEDIA: Figure 3A

    - LAB MEDIA: Figure 3B

    - LAB MEDIA: Figure 3A&B

  - The viscoelastic analysis of collagen using a power law model [1] shows the areal mass [2], complex shear modulus [3], and viscoelastic phase angle respectively [4].

    - LAB MEDIA: Figure 4

    - LAB MEDIA: Figure 4 - Video editor: emphasize Figure 4A

    - LAB MEDIA: Figure 4 - Video editor: emphasize Figure 4B

    - LAB MEDIA: Figure 4 - Video editor: emphasize Figure 4C

  - The first 10 hours indicate the main adsorption stage of the collagen to the sensor surface [1], with the period between 10 and 20 hours showing the equilibration stage [2] before the buffer wash performed at 20 hours [3].

    - LAB MEDIA: Figure 4 - Video editor: emphasize 0-10 h in each figure.

    - LAB MEDIA: Figure 4 - Video editor: emphasize 10-20 h in each figure.

    - LAB MEDIA: Figure 4 - Video editor: emphasize from 20 h and further in each figure.

  - Here is the plot of the viscoelastic phase angle and the complex shear modulus over the general range of samples measured using the QCM (pronounce as Q-C-M) [1]. The green line indicates the linear relationship between the two properties [2].

    - LAB MEDIA: Figure 5

    - LAB MEDIA: Figure 5 - Video editor: emphasize the green line.

  - This plot shows how a PEC can demonstrate a wide range of mechanical properties based on the ratio of polymer, water, and salt in the system [1].

    - LAB MEDIA: Figure 5

# Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Gwen dePolo: The sample preparation directly affects the results from a QCM experiment. Understanding what data you wish to learn from a sample informs how thick the sample should be [1].

    - INTERVIEW: Named author says the statement above in an interview-style shot while looking slightly off-camera.

  - Emily Schafer: The QCM is very useful for sensing environmental and biological processes. It can also probe rheological behavior in the high frequency regime for characterizing the viscoelasticity of materials [1].

    - INTERVIEW: Named author says the statement above in an interview-style shot while looking slightly off-camera.