# Submission ID #: 60673

# Scriptwriter Name: Bridget Colvin

Project Page Link: http://www.jove.com/files_upload.php?src=18505313

### Title: Assessing Early Stage Open-Angle Glaucoma in Patients by Isolated-Check Visual Evoked Potential

Authors and Affiliations: Xiang Fan1,2, Lingling Wu1,2, and Aihua Ding1,2

1Department of Ophthalmology, Peking University Third Hospital

2Beijing Key Laboratory of Restoration of Damaged Ocular Nerve, Peking University Third Hospital

# Corresponding Author:

# Lingling Wu

# wullc@hotmail.com

Email addresses for Co-authors:

fanxiang@126.com

thyrsa_ding@163.com

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

- REQUIRED Interview Statements (Said by you on camera): All interview statements may be edited for length and clarity.

NOTE to Video Editor: Authors accidentally filmed using the Draft and not the final script. The protocol section didn’t change much from the draft to the final, so the only parts affected were the interview statements. For the most part, they’re usable, I just crossed out the ones that we probably shouldn’t use.

  - Xiang Fan: This method can help answer how this unique objective electrophysiological visual field functional measure work in the diagnosis of early stage of open-angle glaucoma.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Xiang Fan: This technique can selectively stimulate M-ON pathway which is initially damaged in glaucoma, especially for early stage.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Xiang Fan: This technique mainly extends toward the diagnosis of glaucoma. It creates low spatial frequency / high temporal frequency bright stimuli and records cortical activity initiated primarily by afferents in the M-ON pathway which is initially damaged in glaucoma.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Xiang Fan: This method can be applied to diagnose disease in optic nerve, optic chiasma, optic radiation or optic cortex by different setting of stimuli.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Xiang Fan: To whom is trying this technique for the first time, I recommend that please note step 1.2, 1.3, 2.2 and 4.1-6 which are key points for success.

    - INTERVIEW: Above Talent speaking the statement above in an interview-style shot, looking slightly off-camera

  - Xiang Fan: Because this method is a very unique objective electrophysiological visual field functional measure which needs experienced doctors or technicians to perform. So visual demonstrations will help a lot.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Introduction of Demonstrator (Said by you on camera):

  - Author Name: Demonstrating the procedure will be _Aihua Ding__ a _technician__ from my laboratory. NOTE: I’m not sure who actually delivered this statement, but the final draft said that Aihua Ding would be delivering the statement and Wei Ni would be demonstrating.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

    - The named technician, post doc, student looks up from workbench or desk or microscope and acknowledges the camera

Ethics title card: (for human subjects or animal work, does not count toward word length total)

  - Procedures involving human subjects have been approved by the Ethics Committee Review Board of Peking University Third Hospital and conformed to the Declaration of Helsinki.

# Section - Protocol

- Isolated-Check Visual Evoked Potential (icVEP) System Setup

  - To set up the icVEP (pronounce ‘I-C-V-E-P’) system, click the Setting button [1] and confirm that the frame rate is set to 6 hertz [2], the luminance of the static background of the display is set to 51 candelas/square meter [3], and the total cycles are set to 20 [4].

    - WIDE: Talent clicking button, with monitor visible in frame

    - SCREEN: screenshot_1_for_1.3 Video Editor: please emphasize Frames per cycle 6

    - SCREEN: screenshot_1_for_1.3 Video Editor: please emphasize Bkgrnd Intensity 51.00

    - SCREEN: screenshot_1_for_1.3 Video Editor: please emphasize Total Cycles 20

  - To differentiate the OAG (pronounce ‘O-A-G’) patients and control subjects, confirm that the sinusoidal temporal signals are set to 10 hertz at six frames per cycle and with a 15% positive-contrast [1].

    - SCREEN: To be provided by Authors: Shot of sinusoidal signals set to 10 hz ag 6 frames/cycle with a 15% positive-contrast. TEXT: OAG: open angle glaucoma NOTE: Authors are in the process of uploading all SC

- Results Reliability Optimization

  - To ensure the acquisition of reliable results, confirm that the refractive error is corrected to adapt for a distance of 114 centimeters [1], that the intraocular pressure is no more than 30 millimeters of mercury on the day of examination [2], that the pupil diameters are at least 2 millimeters and without mydriasis [3], and that each Subject is at rest and is quiet at least 30 minutes before the examination [4].

    - WIDE: Talent checking refractive error

    - Shot of intraocular pressure

    - Shot of pupil diameter/diameter being measured

    - Talent checking that Subject is at rest

  - To avoid the influence of a study curve, check the right eye [1] and the left eye [2].

    - Talent checking right eye

    - Talent checking left eye

  - Then, check the right [1] and left eyes again [2] and record this second result [3].

    - Right eye being checked

    - Left eye being checked

    - Talent recording result

  - Initiate a retest after at least a 30-minute rest when the R-value between both eyes shows a difference of greater than 0.2, indicating that the result is unreliable [1].

    - Shot of R-value > 0.2

- Examination

  - To cue facilitated and careful fixation on the center of the screen, click the IC button [1] to confirm that the spatial pattern is a 24 x 24 array of isolated checks to subtend an 11-degree visual field with a 2 x 2 array fixation cross without sinusoidal temporal signals [2].

    - WIDE: Talent clicking button, with monitor visible in frame

    - SCREEN: To be provided by Authors: Shot of 24 x 24 array of isolated with 2 x 2 array fixation cross w/o sinusoidal temporal signals

  - Place the gold cup electrodes filled with electrolytic water-soluble paste [1] at the midline sites on the scalp based on the international 10-20 system [2], taking care that the testing distance is 114 centimeters [3].

    - Talent placing electrode onto scalp

    - LAB MEDIA: Figure 3

    - Shot of testing distance between subject and system Video Editor: please emphasize 114-cm distance as necessary/appropriate

  - When all of the electrodes have been placed, click the Start button [1]. One run will last for 2 seconds, with the first second presenting half of the test contrast level as an adaptation condition [2] and the second second presenting the full test contrast [3].

    - Talent clicking start, with monitor visible in frame

    - SCREEN: screenshot_1_for 2.3

    - SCREEN: To be provided by Authors: Shot of second half presenting full test contrast

  - Note the prompt from the program and repeat the run when noise is detected and the EEG (E-E-G) epoch is rejected [1-TXT].

    - SCREEN: screenshot_1_for 2.4 TEXT: EEG: electroencephalography

  - The EEG data will be displayed on the monitor when the run is determined to be valid after prompting to either accept or reject the data based on the reliability [1].

    - SCREEN: screenshot_1_for 2.5

- Automatic Data Processing

  - Once the data has been accepted, the program will indicate what the next run should be until a set of eight valid runs are accumulated [1].

    - WIDE: Talent at computer, initiating next run, with monitor visible in frame

  - Each run will produce a fundamental frequency component [1-TXT].

    - SCREEN: screenshot_1_for 3.1 TEXT: If one FFC is outlier, program will discard outlier and prompt for additional run

  - When all 8 runs have been completed [1], the program will calculate the mean fundamental frequency component [2] and the radius of a 95% confidence circle using the T-squared circular statistic that is automatically produced from the eight fundamental frequency components within less than 1 minute after the end of the test [3].

    - SCREEN: screenshot_1_for 3.2

    - SCREEN: screenshot_1_for 3.3 Video Editor: please emphasize Freq Component Calculated: 10.00 Hz

    - SCREEN: screenshot_1_for 3.3 Video Editor: please emphasize SNR = 1.93

# Section – Results

- Results: Representative Early Stage Open Angle Glaucoma Patient Analyses

  - The changing of retinal nerve fiber layer thickness in the temporal superior [1] and inferior quadrants [2] were calculated by subtracting the retinal nerve fiber layer thickness value from the standard value obtained from a database of healthy individuals [3].

    - LAB MEDIA: Figure 4B Video Editor: please emphasize TS quadrant

    - LAB MEDIA: Figure 4B Video Editor: please emphasize TI quadrant

    - LAB MEDIA: Figure 4B

  - In this representative analysis, a total of 44 OAG patients [1] and 39 control subjects were included [2].

    - LAB MEDIA: Table 1Video Editor: please emphasize OAG patients column

    - LAB MEDIA: Table 1Video Editor: please emphasize Control subjects column

  - No statistical differences in age, sex, right or left eye, best-corrected visual acuity, spherical equivalent, or pupil diameter were measured between patients and controls [1] but the signal-to-noise ratio was significantly lower in patients [2].

    - LAB MEDIA: Table 1Video Editor: please emphasize age, sex, right eyes/left eyes, BCVA, spherical equivalent, and pupil diameter rows

    - LAB MEDIA: Table 1Video Editor: please emphasize icVEP-SNR row

  - Using an a priori signal-to-noise ratio criterion of 0.93, the specificity of the receiver-operating-characteristic anlaysis reached 100% with a sensitivity of 65.9% [1].

    - LAB MEDIA: Figure 5 1Video Editor: please emphasize blue data line

  - For the patients, abnormalities in the central 11-degree visual field test were calculated by the numbers of abnormal points with different possibility criteria [1].

    - LAB MEDIA: Figure 4C Video Editor: please outline/emphasize red square

  - With a criterion level of p < 0.5, the amount of abnormal test points in the central 11-degree visual field was significantly negatively correlated with signal-to-noise ratio [1].

    - LAB MEDIA: Table 2 Video Editor: please emphasize When P<5% data row

  - Thickness changing of the retinal nerve fiber layer in the temporal superior quadrant was significantly positively correlated with signal-to-noise ratio [1], while the standard automated perimetry mean deviation, standard automated perimetry mean deviation of the other eye, thickness changing of the retinal nerve fiber layer in the temporal inferior quadrant, and the baseline intraocular pressure and central corneal thickness were not correlated [2].

    - LAB MEDIA: Table 3 Video Editor: please emphasize Temporal Superior quadrant data row

    - LAB MEDIA: Table 3 Video Editor: please emphasize SAP-MD, SAP-MD of the other eye, Temporal Inferior quadrant, Baseline IOP, and CCT data rows

# Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Xiang Fan: If you find a person may be diagnosed as glaucoma suspect or open-angle glaucoma, you can use this procedure. And make sure that step 1.2 & 1.3 setting for isolated-check.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Xiang Fan: Ocular hypertension or subjects with extension of cup-to-disk ratio can also be performed with this procedure to differentiate open-angle glaucoma.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Xiang Fan: Fortunately, the second generation of icVEP have been designed. It was developed for smaller screen and HD display, also shorter exam distance. The machine is much smaller than first generation, which is much more convenient.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Xiang Fan: No, it’s very safe.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera