Answers to the editorial and reviews’ comments;

Editorial comments:
1. The language in the manuscript is not publication grade. Please use professional copy-editing services. The language is severely affecting the review.

Yes. I ordered such a professional copy-editing services after I made the re-revised papers, then I submitted this revised manuscript.

2. The following lines need to be revised to avoid plagiarism: 26-29, 42-47, 289-292, 299-303,

Yes, the author have exchanged them to some new sentences according to this comment by the editor.

Changed line;

Before (Line 26-29 in the original manuscript);

Probing during arthroscopy surgery, which is executed by pulling (or hooking) or pushing soft tissues in and/or around joint with a metallic probe, allows us to assess conditions of the soft tissues. However, the current assessment of the probing by surgeons is only subjective and qualitative, in other words “surgeons’ feeling”.

After (Line 27-29 in the revised manuscript);

Probing in arthroscopic surgery is performed by pulling or pushing the soft tissue, which provides feedback for understanding the condition of the soft tissue. However, the output is only qualitative based on the “surgeon’s feeling”.

Before (42-47 in the original manuscript);

“surgeons’ feeling”. If the resistance of the soft tissues (e.g. capsule or labrum in the hip joint, or tendon or ligament in the knee joint) during the pulling can be measured quantitatively, it must be useful for surgeons to assess to judge the necessity of a repair of the soft tissues and an indication of additional surgical intervention is necessary even after the primary repair is done3-5. Furthermore, some criteria with key quantitative variables for indication of surgical intervention must be established for many surgeons.

After (Line 41-46 in the revised manuscript);

On the basis of this context, if the resistance of the soft tissue (e.g., capsule or labrum in the hip joint, meniscus or ligament in the knee joint) during pulling could be measured quantitatively, it could be useful for surgeons to judge the necessity of a repair for the soft tissue and an indication of whether additional surgical intervention is necessary even after the primary repair is completed3-5. Furthermore, criteria for key quantitative variables to indicate necessary surgical intervention must be established for surgeons.

Before (Line 289-292 in the original manuscript);

According to the current report, the quantitatively measured values can be useful to evaluate whether the soft tissue is intact or not. The highest resistance levels of the acetabular labrum went down when the labrum was cut from the labrum intact situation. The highest resistance levels of the labrum recovered from the situation of the labrum cut when the labrum was repaired. Therefore, the measured value by this probing device can be useful to evaluate whether a surgical intervention is enough

After (Line 289-296 in the revised manuscript);

According to this study, the quantitatively measured values by pull-probing with this device can be useful for evaluating the condition of the joint soft tissue. The highest resistance levels of the acetabular labrum decreased when the labrum was cut. Furthermore, the high resistance levels were recovered when the labrum was repaired. Thus, the probing force can also be useful for evaluating whether surgical intervention is sufficient. Furthermore, this pull-probing can be utilized for assessing other soft tissues as well, such as anterior and posterior cruciate ligaments for instability, medial and lateral collateral ligaments for valgus or varus balance in knee surgeries, labrum and rotator cuff in shoulder surgeries, as well as for other arthroscopic surgeries.

Before (Line 299-303 in the original manuscript);

The 95% confidence interval of the highest resistance level at the labrum intact (6.5 to 9.9) was not overlapped to that at the labrum cut (3.0 to 5.0). Furthermore, the highest resistance levels of the labrum recovered significantly (From the labrum cut 4.0N to 7.8N) from the situation of the labrum cut when the labrum was repaired. Therefore, since a threshold of the highest resistance level of the labrum to detect some lesions of the labrum might be determined, the author consider that the measurement by the probing device can be useful.

After (Line 298-306 in the original manuscript);

I previously reported similar results using 10 fresh cadaver hip specimens with a similar probing device3. The highest resistance levels of the labrum were significantly reduced when the labrum was cut (intact labrum, 8.2 N; cut labrum, 4.0 N). Furthermore, the high resistance level of the labrum was significantly recovered when the labrum was repaired (cut, 4.0 N; repaired, 7.8N). Furthermore, resistance level for the cut labrum (3.0-5.0 N) was statistically separated with 95% confidence from those of the intact (6.5-9.9 N) and repaired labrum (6.7-9.1 N). Therefore, a threshold for detecting lesions in the labrum might be determined, which is approximately 5 N (4-6 N on cadavers) of the highest resistance level of the labrum. According to the current study, such a threshold on the phantom hip might be around 2-3 N.

3. Please avoid overly commercial language. The manuscript cannot become an advertisement for your commercial product. Please be neutral in tone and remove all company names to the Table of Materials.

According to the reviewer’s comment, the author would have avoided overly commercial language.

However, the author does not still understand about underlined “remove all company names to the Table of Materials.” According to the editorial comments to the originally submitted manuscript, the editor mentioned “The table should include the name, company, and catalog number of all relevant materials in separate columns in an xls/xlsx file.” Therefore, I have still remained company names in the excel file.

In addition, this manuscript is about how to use this probing device as your Journal.

Therefore, the author consider that the information of this product should be left.

4. Please specify that ethics approval has been obtained.

The author did not use any human tissues but just artificial materials. Therefore, the author thinks that it is not necessary to include in the ethical approval.

5. Please specify inclusion/exclusion criteria.

The author has already mentioned that the range of the force by this device (“The measurement range of this device was 50 N in the direction of the probe axis, and 10 N in the two remaining directions.” in the previous manuscript).

Therefore, the author consider that the values are alternated as the inclusion/exclusion criteria.

6. As a clinical device is presented, please present results of its clinical efficacy.

Unfortunately, the author has not experienced the probing device as the clinical device in the manuscript. Therefore, the author can’t present any results.

Reviewers' comments:
Reviewer #2:
Manuscript Summary:
The paper describes an instrumented arthroscopic probe which allows to measure the resistance of the tissues in surgical operations. The sensor located in the probe allows to measure quantitatively the tri-axial forces generated between the probe and the tissue. The topic is very interesting and it can be very useful for surgeons.

Major Concerns:
The use of strain gauges could provide transverse sensitivity. That means that when a force is applied in one direction it produces a response not only in that direction but also in transversal directions. The response in transversal directions reduce the accuracy of the probe. The author should provide information about transverse sensitivity of the device. Information about the resolution (i.e. smallest change that can be detected) of the device is also desirable.

To address the reviewer’s comment, the author re-measured the information of the transverse value while probing in the representative result as one example. The maximum transverse value was just 2.8 % to the resultant force of z and y directions (Line 239). In addition, the author confirmed the resolution of this probing device (0.005N), then added it in the revised manuscripts (Line61).

Minor Concerns:
Line 220, where it says "pf" it should say "of".
According to the reviewer’s comment, the author revised it.

Reviewer #3:
Manuscript Summary:
The study describes a device and presents a protocol to quantitatively probe joint soft tissues.

Major Concerns:
1. Another potential application of the presented probing device could be pre-implant assessing of allografts (e.g. tendons for Anterior Cruciate Ligament reconstruction or menisci)

The author agrees with the reviewer’s comment. The author added this potential as a comment in the discussion part.

Added lines (293-296);

Furthermore, this pull-probing can be utilized for assessing other soft tissues as well, such as anterior and posterior cruciate ligaments for instability, medial and lateral collateral ligaments for valgus or varus balance in knee surgeries, labrum and rotator cuff in shoulder surgeries, as well as for other arthroscopic surgeries.

2. Controls seem appropriate, but something more should be stated about the mechanical values of cartilage samples calculated by the "customized device for classical indentation", validating the read values against literature or manufacturer information

The author considers that results of this kind of experiments depends on some conditions such as the size, shape, compression speed of the probe. The author did not validate the data by the device in the manuscript. However, the author confirmed that the mechanical values of cartilage samples by this customized device in another publication (Ref. 6) are within those in several previous literatures (Ref. 11，16，19). In such, the author added this statement in the supplement of the manuscript.

Changed lines (364-366);

I did not validate the data by this device, but I confirmed the mechanical values of cartilage samples by this device previously; the elastic modulus was 0.46 MPa (0.27 MPa SD)6, which is consistent with that found in several previous literature studies11,16,19.

3. Lines 64-69: it is not clear what "distance" stands for. Distance of what? Moreover, it seems that pulling force and distance are operator-dependent: how to fix them, or how to compare different situations? That is, how to assure "repeatability", which together with "accuracy" and "precision" makes a measurement actually "quantitative"?

Although the reviewer might not look well, there is a gap between the sliding part and grip component. The sliding part can be shifted with the index finger to the grip component. Also the distance of this sliding can be arranged with the posterior nut. On the other hand, if pulling probing, the above gap is firstly made, then the sliding part is slid until it is touched to the grip component.  If holding the grip component by hand, the value can be controlled.

According to the reviewer’s comment, the author confirmed the repeatability by measuring three times. The value was described in the representative result of the revised manuscript.

Changed lines (61-63);

The precision and accuracy were also measured by a commercialized weight with known weight (50 g).  The precision was 0.013 N and the accuracy is 0.0035 N.

4. In this context, surgical common practice is qualitative probing. Therefore, quantitative sensing represents a promising upgrade. In the DISCUSSION paragraph, lines 315-331, the author lists the quantitative solutions presented in literature with their limits, which should be overcome by the probing device of the manuscript. As regards cartilage, at least a comparison, with "Arthro-BST" device, is lacking and should be added. "Arthro-BST" appears in REFERENCES (ref. 23), but is not cited in the text (see also Quantitative Arthroscopic Assessment of Articular Cartilage Quality by Means of Cartilage Electromechanical Properties, 2018)

The author mistook the numbering of the references. The author cited it as ref.22, and also cited the paper which the reviewer recommend to read as ref.23.

The author also indicated the comparison between the author’s device and “Arthro-BST”.

Minor Concerns:
1. Revise the text at lines 42-45

According to the reviewer’s comment, the author revised it.

2. Line 50, remove "evaluation"

According to the reviewer’s comment, the author revised it.

3. Line 66, "measured value" in place of "measuring value to evaluate the resistance"

According to the reviewer’s comment, the author revised it.

4. Revise paragraph at lines 71-79, in particular the verb tense

According to the reviewer’s comment, the author revised it.

5. It seems from Figure 2 that auto-scale could be useful for better showing the measures

The author asked the software company, but its answer was not easy, and additional cost was necessary. Therefore, the author selected alternative suggestion, the range of the scale had become shorten (z-axis; 50N to 10N, x and y- axes: 8N). The figure 2 was revised by the above change.

6. Revise line 165-166, e.g. "… when some manual checking is done, for example by touching …"

According to the reviewer’s comment, the author revised it.

7. There is a NOTE opening paragraph 2, a sort of summary. For symmetry, a NOTE also at the beginning of paragraph 1 should be added.

Although the reviewer suggest to add the NOTE for the paragraph1, the author already have written the NOTE for it.

8. Line 211, remove "to make"

According to the reviewer’s comment, the author revised it.

9. Line 215-216, rewrite as "… with a stopper opposite to the side of the push-probing"

According to the reviewer’s comment, the author revised it.

10. Line 218, rewrite as "2.2. Measurement of cartilage resistance while push-probing"

According to the reviewer’s comment, the author revised it.

11. Line 230, rewrite as "2.3. Measurement of cartilage resistance by a classical indentation device"

According to the reviewer’s comment, the author revised it.

12. Paragraph 2.3.2 deals with correlation between probe and indentation device measures. Given the small number of points, Spearman correlation (r and p) should be indicated and calculated

Although the reviewer suggest to add the Spearman correlation’ result, the author already have written it in Line 245-.

13. Line 239, add a number to REPRESENTATIVE RESULTS (i.e. 3). Lines 240 and 245 become respectively 3.1. and 3.2 subparagraphs.

According to the reviewer’s comment, the author revised it.

14. Line 247, use "significant" instead of "good"

According to the reviewer’s comment, the author revised it.
15. Line 257, remove "force sensor while probing"

According to the reviewer’s comment, the author removed it.

16. Line 283, rewrite as "The current report shows that the probing device is able to measure…"

According to the reviewer’s comment, the author rewrote it.

17. Line 284, use "Specifically" in place of "Then"

According to the reviewer’s comment, the author changed it.

18. Line 302-304, the author comments about a possible threshold definition. The procedure to define it, on cadaver or phantom hip, should be indicated in the text

According to the reviewer’s comment, the author added it .

19. Line 306, use "significant" instead of "good"

According to the reviewer’s comment, the author changed it.

20. Restate concepts at lines 308-310 in a clearer manner

According to the reviewer’s comment, the author revised the concepts, explained more.

21. Line 325, remove "is only"

According to the reviewer’s comment, the author removed it.

22. Line 353, reference is 24 and not 23.

Actually, the author mistook the numbering of the references. The author revised it. (The referenced paper of the number 7 was separated as 7 and 8 in the previous manuscript. Finally the reference became 25 due to the addition of another ref.)

Miscellaneous

- The author added the company name of some elastic thin cover (CIV-Flex Transducer cover, Civco) at the line 82 in the revise manuscript.

- The author revised the sentences of the Figure 6.