Submission ID #: 60733

Scriptwriter Name: Anastasia Gomez

Project Page Link: https://www.jove.com/account/file-uploader?src=18523613

Title: Cooling or Warming the Esophagus to Reduce Esophageal Injury During Left Atrial Ablation in the Treatment of Atrial Fibrillation

Authors and Affiliations:

Jason Zagrodzky1, Mark M. Gallagher2, Lisa W.M. Leung2, Tiffany Sharkoski3, Pasquale Santangeli4, Cory Tschabrunn4, Jose M Guerra5, Bieito Campos5, John MacGregor6, Jamal Hayat2, Brad Clark7, Alex Mazur8, Marcel Feher9, Martin Arnold9, Mark Metzl10, Jose Nazari10, Erik Kulstad11

- St. David's South Austin Medical Center, Austin, TX

- St George's University Hospitals NHS Foundation Trust, St. George’s, University of London, London, UK

- Hospital of the University of Pennsylvania, Perelman Center for Advanced Medicine, Philadelphia, PA

- University of Pennsylvania Perelman School of Medicine, Philadelphia, PA

- Hospital de la Santa Creu I Sant Pau, Universitat Autònoma de Barcelona, CIBERCV, Carrer de Sant Quintí, Barcelona, Spain

- PeaceHealth Medical Group, St. Joseph Medical Center, Bellingham, WA

- St. Vincent Hospital, Indianapolis, IN

- University of Iowa, Iowa City, IA

- University Hospital Erlangen, Erlangen, Germany

- NorthShore University Health System, Evanston, IL

- Department of Emergency Medicine, University of Texas, Southwestern Medical Center, Dallas, TX

Corresponding Author:

Jason Zagrodzky (jzagrodzkyep@gmail.com)

Email addresses for Co-authors:

Mark Gallagher 	(Mark.Gallagher@stgeorges.nhs.uk)

Lisa Leung 		(Lisa.Leung@stgeorges.nhs.uk)

John MacGregor 	(JMacGregor@peacehealth.org)

Tiffany Sharkoski 	(Tiffany.Sharkoski@uphs.upenn.edu)

Pasquale Santangeli 	(pasquale.santangeli@uphs.upenn.edu)

Cory Tschabrunn 	(cory.tschabrunn@uphs.upenn.edu)

Jose M Guerra 	(jguerra@secardiologia.es)

Bieito Campos 	(bcamposg@santpau.cat)

Jamal Hayat 		(jamal.Hayat@stgeorges.nhs.uk)

Brad Clark 		(bradley.clark@ascension.org)

Alex Mazur 		(alexander-mazur@uiowa.edu)

Marcel Feher 		(Marcel.Feher@uk-erlangen.de)

Martin Arnold 		(martin.arnold@uk-erlangen.de)

Mark Metzl 		(MMetzl@northshore.org)

Jose Nazari 		(JNazari@northshore.org)

Erik Kulstad 		(Erik.Kulstad@UTSouthwestern.edu)

Author Questionnaire:

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique? N

2. Does your protocol include software usage? N

3. Will the filming need to take place in multiple locations? N

If yes, how far apart are the locations?

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

Authors, these headshots will be used for the JoVE Dedicated Author Webpage. Here is one example if you wish to take a look.

- REQUIRED Interview Statements: (Said by you on camera)  - All interview statements may be edited for length and clarity.

NOTE: Some interview statements were filmed by authors and should be uploaded to the project page.

  - Mark Gallagher: This protocol is significant for demonstrating an approach to esophageal cooling or warming during left atrial ablation for the treatment of atrial fibrillation.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Jason Zagrodzky: The main advantage of this technique is the ability to provide proactive heat transfer rather than detecting thermal energy after the fact using standard temperature monitoring.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Lisa Leung: This technique also avoids the need for procedural pauses and device manipulation required with other methods used to reduce esophageal injury.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Mark Metzl: This technique is easy to implement, and can be used for both radiofrequency (RF) ablation and cryoablation.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

Introduction of Demonstrator: (Said by you on camera)

  - Mark Metzl: Demonstrating the procedure will be Natalie Zoltek, RN, MSN.

    - INTERVIEW: Author saying the above

    - The named technician, post doc, student looks up from workbench or desk or microscope and acknowledges the camera.

Ethics title card: (for human subjects or animal work, does not count toward word length total)

  - This protocol follows the guidelines of local institution’s human research ethics committee.

# Section - Protocol

- Placement of Device

  - Begin by attaching the esophageal temperature measurement device to the heat exchanger via the device connectors [1]. Power on the unit, place it in manual mode [2], and ensure that the water is flowing through the device with no leaks [3].

    - WIDE: Establishing shot of talent connecting the device to the heat exchanger.

    - Talent powering on the unit and placing it in manual mode.

    - Talent inspecting the device.

  - Determine the appropriate insertion depth by measuring from the patient’s lips to the earlobe [1] and from the earlobe to the xiphoid process [2], noting this depth on the device [3].

    - Talent measuring from lips to earlobe.

    - Talent measuring from earlobe to xiphoid process.

    - Talent noting the depth on the device.

  - Next, use a water-soluble lubricant to generously lubricate between 15 and 25 centimeters of the distal end of the device [1]. If possible, extend the patient’s head to further facilitate insertion [2] and apply gentle pressure posteriorly downwards to insert the device past the oropharynx and into the esophagus [3].

    - Talent lubricating the device.

    - Talent extending the patient’s head.

    - Talent gently inserting the device. Video Editor: Use Figure 1 as an inset or split screen here, or show it immediately after this shot.

  - To increase ease of placement, consider pinching the water outflow tract to increase device stiffness [1]. If the endotracheal tube cuff is overinflated, consider temporary pressure reduction via air removal to avoid compressing the esophageal pathway [2].

    - Talent pinching the water outflow tract.

    - Talent removing air.

  - If necessary, determine the placement location via fluoroscopy to check if the tip of the device is below the diaphragm [1]. Secure the water hoses and device to avoid accidental dislodgement [2] and, if stomach decompression is desired, connect the central lumen to low-intermittent suction [3-TXT]. The device can be visualized with intracardiac echocardiography, or ICE (pronounce ‘ice’) [4].

    - Talent determining placement of device with fluoroscopy. Videographer: Film the fluoroscopy images on the screen as well. NOTE: Author will upload image, if not use Figure 5.

    - Talent securing the water hoses and device.

    - Talent connecting the central lumen to suction.

    - LAB MEDIA: Screen shot of device visualized with ICE. NOTE: Author should upload image

- Temperature Modulation and Monitoring

  - For radiofrequency, or RF, ablation, ensure that the heat exchanger is set to manual mode and that the appropriate water temperature is set [1]. On a typical heat exchanger, press the Temp Control button and use the up and down arrows to select the target water temperature [2].

    - WIDE: Talent adjusting settings on heat exchanger. NOTE to video editor: 3.4.1. can be used here

    - Close up shot of talent pressing Temp Control and selecting the temperature.

  - Once the digital display shows the desired target temperature, initiate water flow by pressing the Manual Control button [1]. Use a water temperature setpoint of 14 degrees Celsius for initial insertion [1].

    - Talent pressing the Manual Control button after target temperature is reached.

    - Device set to 14 degrees Celsius. NOTE to video editor: 3.4.1. can be used here

  - Approximately 7 to 10 minutes before application of RF energy, change the setpoint to 4 degrees Celsius [3].

    - Talent adjusting the setpoint to 4 degrees

  - For additional anti-inflammatory effects of cooling after RF ablation, maintain the water temperature setpoint at 4 °C for up to 20 minutes after the completion of posterior wall ablation [1].

    - WIDE: Talent adjusting the temperature setpoint and/or timing an additional 20 minutes for treatment.

  - For cryoablation, set the water temperature to 42 degrees Celsius shortly after placement [1] and provide additional patient warming if needed to counter the systemic cooling effect of cryoablation [2].

    - Talent setting the temperature to 42 degrees.

    - Talent covering the patient with a blanket.

- Temperature Monitoring and Troubleshooting

  - If needed to maintain patient temperature while using the esophageal cooling, supplemental warming modalities such as warming blankets or head covers can be used [1]. NOTE: Move 4.2 before 4.1.

    - Talent covering the patient with a blanket or head cover.  NOTE: Author will upload picture showing patient temperature on standard monitor. Add this as an inset.

  - During ablation, the patient temperature will usually stay in a normothermic range. To measure patient temperature, use any approved thermometer other than esophageal, since device interference will prevent accuracy [1]. NOTE: Move 4.2 before 4.1.

    - Talent measuring patient’s temperature.

  - Confirm water flow at correct temperature by checking the setpoint and touching the device, which should be firm [1a], to ensure adequate pressure and appropriate temperature [1b-added shot].

4.3.1bTalent touching the device.

4.3.1.a Added shot: Talent confirming “water at setpoint”

  - When troubleshooting, ensure that there is no blockage of water flow [1]. The water paddle wheel, if present, should be continuously spinning and the low-flow alarm should not be activated [2]. If blockage occurs, stop treatment and determine the cause and location of obstruction [3].

    - Talent inspecting device for blockage.

    - Water paddle wheel spinning.

    - Talent stopping treatment.

- Removal of Device

  - To remove the device, press the appropriate button to pause water flow [1], close clamps on the hose set and device tubing if present and withdraw the device by gently pulling anteriorly in a similar manner to standard orogastric tube removal [3].

    - Talent pressing the button to pause water flow.

    - Talent withdrawing the rube.

  - Power down the heat exchange unit via the power switch [1] prior to unplugging from wall power [2]. Disconnect the water hose connectors from the device [3] and dispose according to institutional guidelines [4].

    - Talent flipping the power switch.

    - Talent unplugging device.

    - Talent disconnecting the water hose.

    - Talent disposing of the hose.

# Section – Results

- Results: Outcomes of Esophageal Cooling during Left Atrial Ablation

  - A meta-analysis of existing studies found that esophageal cooling via direct instillation of cold liquid into the esophagus during RF ablation [1] results in a 61% reduction in high-grade lesion formation [2].

    - LAB MEDIA: Figure 6. Video Editor: If the resolution of Figure 6 is poor, please use the figure from Meta-analysis summary.xlsx

    - LAB MEDIA: Figure 6. Video Editor: Emphasize both bars of the Treatment data.

  - Data from a randomized-controlled clinical trial evaluating a dedicated cooling device also show a statistically significant 74% reduction in endoscopically identified lesions compared to standard LET (pronounce ‘L-E-T’) monitoring. Sample case studies of RF and cryoablation in combination with an esophageal heat transfer device are outlined in the text manuscript [1-TXT].

    - LAB MEDIA: Table 1. Video Editor: Add the following information in text overlay underneath the table. The numbers indicate ablation parameters for the “protected” versus “control” groups, respectively. Position the text in a way that makes it obvious which numbers correspond to which group. TEXT: Ablation Parameters  ;  RF duration: 14.5 versus 14.1 minutes  ;  Average Force: 17.8 versus 19.1 grams  ;   Maximum RF Power: 34.1 versus 33.9 W, and Average Ablation Index: 384 versus 394

# Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Natalie Zoltek: When attempting this procedure, it is important to ensure that water is flowing continuously prior to placement, to provide sufficient stiffness, and when the device is being used, to ensure continuous heat transfer.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B-roll: 2.1.3. for water flow, 4.3.1. for stiffness and heat transfer.