## Sheet1
Problem | Cause | Solution
Failed to correct the reference image | Low contrast | Acquire a higher contrast image if possible. If a bright-field reference image is used, reacquire the image in a water-based solution to obtain higher contrast of the cell. Alternatively, try applying computational noise reduction (e.g. Gaussian filtering). Turn off local alignment, which is more sensitive to noise.
| Contamination of unrelated images | Remove the source of the unrelated images in the sample if possible. For crosstalk reference images, check the excitation spectra of the dyes used for the target images. If the dyes are excited during acquisition of a crosstalk image (e.g. Alexa Fluor 568 or 594), consider other dyes (e.g. Alexa Fluor 555). If dusts on the camera chip creates an obvious channel difference, clean the camera chip or use a computational flat-fielding method.
| An extremely bright spot made by a cosmic ray | Acquire the image again if possible. Alternatively, try applying computational noise reduction (e.g. Median or Gaussian filtering).
| Deconvolution artifacts (artificial signals at the axial and lateral edges) | Trim the edge pixels or Z sections after deconvolution. If one side is trimmed, the other side should also be trimmed to maintain the image center.
| The Z step size too sparse | A Z stack should be acquired to fulfill the Nyquist criterion as written in protocol 2.1.3.
| Optical aberration | Spherical aberration is the major aberration caused by users. Choose the right objective lens for the sample and use a coverslip thickness of 170 µm. If the objective lens is equipped with a correction ring, adjust it to find the position where the highest fluorescence count is obtained from the focus. In the case of an oil-immersion objective without a correction ring, adjust the refractive index of the immersion oil that increases the fluorescence count at the focus.
| Field of view is unfilled (Fig. 7) | In the case of biological calibration reference images, average many images. In the case of crosstalk or bright-field reference images, use local alignment.
| An unidentified software bug | Report the issue through GitHub (https://github.com/macronucleus/Chromagnon/issues)
Failed to correct the target image | Metadata of the image file is lost | Use the original microscope file format which contains complete metadata, and avoid converting to a multipage tiff file before processing. Use the same ordering of channels as written in protocol 3.3.
| Wrong alignment methods for the given microscopy | Do not apply the local alignment method when measuring from biological calibration reference images to target images. Do not use crosstalk reference images other than wide-field microscopy.
| Differences in imaging conditions | Keep the imaging conditions constant between the reference and the target images as written in protocol 2.3.3.
| Differences in sample (including coverslip) | Always use the same mounting medium, coverslip (e.g. No. 1.5H) and a similar depth of focus.
| Microscope drift since the calibration was last made | Make a calibration as often as every two weeks. Keep the temperature constant, and use a floating table to avoid hardware drift of the microscope.