60811 Time codes

60811_screen_capture_1:

1.3.12 Graph the deflection of the silicone post (x, on the x-axis) against the gravitational force of each test weight (mg, on the y-axis). This should yield a linear relationship between force and deflection. 00:02 - 00:40

1.3.13 Plot a linear regression function passing through (0,0) and the acquired data (see Figure 6A for examples). The slope of this function (mg=kx) is the stiffness k of the magnetically responsive silicone post at the tested magnet spacing. 00:41 - 01:24

60811_screen_capture_2:

1.3.17 Plot the resulting k values against the respective magnet distances. This should yield a negative exponential relationship.  00:02 - 00:42

1.3.18 Plot a regression function through these values. Here, the “Non-linear fit  One-phase decay”-function in GraphPad Prism (version 6.00 for Windows, GraphPad Software, La Jolla California USA) was used. This regression function describes the relationship between magnet spacing and afterload (see Figure 6B for an example). 00:43 - 01:53

60811_screen_capture_3: (“DUMMY”)

3.2.2 Connect the software to the piezo stage motor by selecting the port designated as the stage port during installation of the motion control software and then click the “open port”-button.

00:02 - 00:19

3.2.3 Go to the “System”-panel. Select “Open Loop” in the “Loop”-dropdown menu. 00:20 - 00:26

3.2.5 Go to the “Motion”-panel. Click the “Zero”-button to reset the current position of the piezo stage to 0 mm. 00:34 - 00:39

3.2.6 Manually move the magnet plate to its lowest possible position. Write down the encoder position (indicated in the “Motion”-panel by “Enc”) to determine the range of motion for the piezoelectric stage motor. 00:40 - 00:55

3.2.7 Set the “Travel Limits” in the “System”-panel to values within the range of motion determined in the previous step. This prevents the magnet plate from bumping into the culture plate or the bottom of the afterload tuning device. 00:56 - 01:10

3.2.9. Go to the system panel and change the feedback loop mode to “Closed Loop”. Doing this ensures that the stage will correct for any errors in its positioning. 01:22 - 01:28

3.2.10 Click the “Save”-button in the “Save Parameters”-box to store these settings in the system.

Not shown…

60811_screen_capture_4: (“DUMMY”)

3.2.14 Type this value into the “Target Position 1”-input field in the “Motion”-panel and click “Go” to adjust the EHTs’ afterload to the calculated value. 00:05 - 00:40