# Submission ID #: 60811

# Scriptwriter Name: Brigid Stadinski

# Project Page Link: http://www.jove.com/files_upload.php?src=18547403

# Title: Magnetic Adjustment of Afterload in Engineered Heart Tissues

# Authors and Affiliations:

Benjamin Becker1,2*, Marita L. Rodriguez1,2*, Tessa R. Werner1,2, Justus Stenzig1,2, Thomas Eschenhagen1,2, Marc N. Hirt1,2

1Institute of Experimental Pharmacology and Toxicology, University Medical Center Hamburg-Eppendorf, Hamburg, Germany

2DZHK (German Centre for Cardiovascular Research), partner site Hamburg/Kiel/Lübeck, Germany

*Equal author contribution to this work.

Authors: Please ensure that all authors’ names are spelled correctly and that the affiliations listed here are correct. This is how your names and affiliations will appear in your video.

# Corresponding Author:

Marc N. Hirt (m.hirt@uke.de)

# Email addresses for Co-authors:

Benjamin Becker (6693850@stud.uke.uni-hamburg.de)

Marita L. Rodriguez (m.rodriguez@uke.de)

Tessa R. Werner (tewerner@uke.de)

Justus Stenzig (j.stenzig@uke.de)

Thomas Eschenhagen (t.eschenhagen@uke.de)

# PLEASE READ THE INSTRUCTIONS IN THE GRAY BOXES CAREFULLY AND USE TRACK CHANGES WHILE MAKING ANY EDITS TO THE DOCUMENT.

# This document has several sections on separate pages, so take care to view each page.

Author Questionnaire:

Authors, please fill out the unanswered questions below.

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique? (Y/N)  No

Can you record movies/images using your own microscope camera? (Y/N)

If no, JoVE will need to record the microscope images using our scope kit (through a camera port or one of the oculars). Please list the make and model of your microscope.

2. Does your protocol include software usage? (Y/N)  Yes

If yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps.

3. Which steps from the protocol section below are the most important for viewers to see? Please list 4-6 individual steps using the step numbers listed in this document. This information is important to prepare your Videographer for your shoot. (You do not need to include steps that will be screen captured. Please do not list entire sections.)

The most important steps for the viewer to see are steps 2.2.1, 2.4.1, 2.6.2, 2.12.1 and 3.12.2.

Authors, please answer this question with the steps listed here in the Protocol section below for use by the videographer.

4. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1-2 individual steps using the step numbers listed in this document. (Please do not list entire sections.)

Authors, please answer this question with the steps listed here in the Protocol section below for use by the videographer.

The braces manufactured in steps 2.4.1 and 2.4.2 have to be bent rather precisely. To ensure a good fit, we use a jig as a template for bending and also test the fit of each brace into a silicone rack. We discard braces that do not slide easily into the hollow silicone posts.

5. Will the filming need to take place in multiple locations? (Y/N) No

If yes, how far apart are the locations?

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

Authors, these headshots will be used for the JoVE Dedicated Author Webpage. Here is one example if you wish to take a look.

- REQUIRED Interview Statements: (Said by you on camera)  - All interview statements may be edited for length and clarity.

- The total introduction length (i.e., Required and Optional Interview Statements) cannot exceed 150 words.

- Restrict the length of each statement to no more than 30 words.

- Please answer the questions below in full sentences to highlight the significance of your protocol. You will be expected to memorize and deliver these sentences as spoken interview statements during filming.

- Indicate the full name of each author who will give each statement. If only one author is giving the REQUIRED statements, the same author may speak both statements.

# Why is your protocol significant? OR What key questions can this method help answer?

  - Benjamin Becker: This system can be used to expose engineered heart tissues to varying regimens of afterload in order to study their response regarding force development, remodeling and maturation.

# What is the main advantage of this technique?

  - Benjamin Becker: Our technique allows for precisely subjecting beating tissues to a broad range of afterload in customizable patterns during long term culture, all without even opening the cell culture plate.

OPTIONAL Interview Statements: (Said by you on camera)  - All interview statements may be edited for length and clarity.

- If you would like to have additional speakers, the following OPTIONAL questions may be answered to provide additional introductory information about your protocol.

- These statements must be spoken by different authors than those who gave the Required Interview Statements, and are limited to one statement per additional author.

- The length of each statement is restricted to no more than 30 words and contributes to the total introduction length, which cannot exceed 150 words.

- Indicate the full name of each author who will give each OPTIONAL statement.

# Do the implications of this technique extend toward the therapy (or diagnosis) of a particular disease, disability, or challenge? How so?

  - Author Name: ___________(Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera).

# Are there any specific areas of research that this method could provide insight into?

# Can this method be applied to any other systems?

  - Marc Hirt: Our method could be modified to control afterload in different muscle tissue culture systems, which could include skeletal muscle, smooth muscle or excised papillary muscle.

# How would you expect an individual who has never performed this technique to struggle?

# Do you have any advice to offer to somebody who is trying this technique for the first time?

  - Author Name: ___________ (Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

# Why is visual demonstration of this method critical?

  - Justus Stenzig: Since our system’s approach is rather unique and features some technical aspects that many biomedical scientists are not familiar with, visual demonstration may aid in recreating our setup.

# Introduction of Demonstrator: (Said by you on camera)

- Please use this statement ONLY if any of the individuals who will be demonstrating the procedure on camera have not given a required or optional Introduction interview statement already.

- Include the full name(s) of the person(s) demonstrating the experiment followed by their title (e.g., technician, post doc, grad student, clinician, etc.)

- Also indicate the full name of the author who will introduce the demonstrator(s).

  - Author Name: Demonstrating the procedure will be _________ (name of the person or persons), a _________ (technician, post doc, grad student) from my laboratory. (Add additional mention of demonstrators as necessary).

    - Interview style: Author saying the above

    - The named technician, post doc, student looks up from workbench or desk or microscope and acknowledges the camera.

# Section - Protocol

- Read through the entire protocol carefully to understand what you will need on the filming day and prepare accordingly.

- The two-digit numbers (e.g. 2.1., 2.2.) represent the “steps” of you protocol and will be read by a professional voiceover talent.

- The three-digit numbers (e.g. 2.1.1., 2.2.2.) represent the “shots” that our videographer will capture at your lab.

- To ensure that your protocol can be filmed in a single work day, the protocol is restricted to 30 steps and/or 60 shots.

- It is critical for a smooth and organized shoot that all materials and work spaces are prepared and labeled (if applicable) in advance.

- Any specimens/samples that require long or overnight incubation steps should be prepared in advance. (i.e. day 0 sample preparation will be filmed on the day of the shoot; day 1 samples should be prepared the day before the shoot so their processing can be filmed on the day of the shoot/after their overnight culture/treatment/etc.)

- Preparation of the Afterload Tuning Platform

  - To manufacture the magnetically responsive silicone racks, acquire the 24-well plate compatible racks of silicone posts described in the text protocol [1].

    - Talent positions the silicone racks.

  - Keeping a fixed polarity, lubricate the magnets with water and insert them, one at a time, into the outermost posts of the silicone racks [1].

    - Talent lubricates the magnets with water and inserts one of them into the outmost posts of the silicone racks.

  - Use a blunted piece of stainless-steel dental wire to carefully push them to the bottom of the hollow post cavity [1-TXT]. You can stack up to five magnets in each post [2].

    - Talent uses a blunted piece of stainless-steel dental wire to carefully push the magnets to the bottom of the hollow post cavity. TEXT: See text for all materials

    - Post as talent continues to stack the magnets.

  - Use round-nose pliers to bend stainless steel dental wire into braces 11.25 millimeters wide and 15 millimeters long. To ensure the correct dimensions are achieved, one can use a self-made jig to aid in the wire bending [1]. Then, use wire cutters to cut the braces and a file to smooth the cutting surface [2].

    - Talent bends the stainless steel dental wire into braces 11.25 mm wide and 15 mm long using a self-made jig.

    - Braces as talent uses wire cutters to cut the braces and a file to smooth the cutting surface.

  - Lubricate the braces with water and insert them into the silicone rack, fixing the second- and third- to outermost post in the process [1].

    - Talent lubricates the braces with water and inserts them into the silicone rack, fixing the second- and third- to outermost post in the process.

  - Begin preparing the afterload tuning device as described in the text protocol. Attach the magnet holder to the piezoelectric stage using a non-magnetic material [1]. This can be achieved using an L-shaped piece of aluminum [2].

    - Talent begins attaching the magnet holder to the piezoelectric stage using an L-shaped piece of aluminum. Continue action in next shot.

    - L-shaped piece of aluminum as talent uses it to attach the magnet holder to the piezoelectric stage.

  - To enable visual analysis of the tissues, install a light source within the afterload tuning device [1]. Here, an array of LEDs was employed to illuminate the engineered heart tissues from below [2].

    - Talent points out the installed light source within the afterload tuning device.

    - Afterload tuning device as talent turns them on to demonstrate how they illuminate the EHTs from below.

  - To calibrate the afterload tuning system, mount one of the silicone racks vertically using non-magnetic materials, such that the magnetically responsive silicone posts are oriented horizontally [1].

    - Talent mounts one of the silicone racks vertically, such that the magnetically responsive silicone posts are oriented horizontally.

  - Now, mount one of the plate magnets on a horizontally traveling linear stage such that it is axially aligned with the magnetically responsive post [1].

    - Plate magnet as talent mounts it on a horizontally traveling linear stage such that it is axially aligned with the magnetically responsive post.

  - Position the calibration magnet a defined distance from the magnetically responsive silicone post using the horizontal stage [1].

    - Talent positions the calibration magnet a defined distance from the magnetically responsive silicone post using the horizontal stage.

  - Place a camera to the side of this set-up in order to be able to optically record the post’s deflection under the influence of the test loads [1]. Take a picture of the post in the absence of any weights to use as a reference for the post’s “neutral” position [2].

    - Talent places a camera to the side of this set-up to optically record the post’s deflection under the influence of the test loads.

    - Talent takes a picture of the post in the absence of any weights.

  - Without changing the camera’s perspective, attach one of the loads to the very end of the silicone post [1]. Then, take a picture of the post bending under the influence of the weight [2].

    - Talent attaches one of the loads to the very end of the silicone post.

    - Post bending under the influence of the weight as talent takes a picture – Authors is this something that we can film? The post deflection is usually so small (about 1 mm) that I’m not sure whether you can show it properly in a regular video recording. The deflection is easily visible in the picture taken with the magnifying USB-camera during the step itself though, if the camera shot doesn’t turn out well.

  - Now, graph the deflection of the silicone post on the x-axis [1] against the gravitational force of each test weight on the y-axis [2]. This should yield a linear relationship between force and deflection [3].

    - 60811_screen_capture_1 – Video editors, please show from 0:22-0:28

    - 60811_screen_capture_1 – Video editors, please show from 0:10-0:21

    - 60811_screen_capture_1 – Video editors, please show from 0:29-0:40

  - Plot a linear regression function passing through (0,0) and the acquired data [1]. The slope of this function is the stiffness k of the magnetically responsive silicone post at the tested magnet spacing [2].

    - 60811_screen_capture_1 – Video editors, please show from 0:41-1:00

    - 60811_screen_capture_1 – Video editors, please show from 1:01-1:24

  - Plot the resulting k values against the respective magnet distances. This should yield a negative exponential relationship [1].

    - 60811_screen_capture_2 – Video editors, please show from 0:22-0:38

  - Plot a regression function through these values. For example, use Non-linear fit, One-phase decay function in the analysis software. This regression function describes the relationship between magnet spacing and afterload [1].

    - 60811_screen_capture_2 – Video editors, please show from 0:41-1:24

- Afterload Modification Experiments

  - To prepare the afterload tuning device for experiments, connect the piezoelectric stage motor to the motion controller, and connect the motion controller to the computer [1]. Make sure the motion controller is also connected to a power source [2].

    - Talent connects the piezoelectric stage motor to the motion controller and connects the motion controller to the computer.

    - Talent secures the motion controller plug into a power source.

  - Then, start the motion controller platform software. Connect the software to the piezo stage motor by selecting the port designated as the stage port during installation of the motion control software and then click the open port button [1].

    - 60811_screen_capture_3_UPDATED_031219 – Video editors, please show from 0:02-0:15

  - Go to the System panel. Select Open Loop in the Loop dropdown menu [1].

    - 60811_screen_capture_3_UPDATED_031219 – Video editors, please show from 0:16-0:23

  - Manually move the magnet plate to its highest position, the closest possible magnet spacing dmin. The magnet plate should make contact with the culture plate mount [1]. Authors, how should the narrators pronounce dmin in the narration of the script? We would suggest pronouncing it “demin”, very similar to “demon”.

    - Magnet plate as talent manually moves it to its highest position and the magnet plate makes contact with the culture plate mount.

  - Now, go to the Motion panel. Click the Zero button to reset the current position of the piezo stage to 0 millimeters [1].

    - 60811_screen_capture_3_UPDATED_031219 – Video editors, please show from 0:24-0:35.

  - Manually move the magnet plate to its lowest possible position. Write down the encoder position to determine the range of motion for the piezoelectric stage motor [1].

    - Magnet plate as talent manually moves it down to its lowest possible position. Editors, maybe it would make sense to show shots 3.6.1 and 3.6.2 as split screen?

    - 60811_screen_capture_3_UPDATED_031219 – Video editors, please show from 0:36-0:54. Please use a zoom bubble to highlight the values moving next to “Enc.”

  - Set the Travel Limits in the System panel to values within the range of motion determined in the previous step. This prevents the magnet plate from bumping into the culture plate or the bottom of the afterload tuning device [1].

    - 60811_screen_capture_3_UPDATED_031219 – Video editors, please show from 1:00-1:12.

  - Once again, move the magnet plate to its highest position and click the Zero button. Go to the system panel and change the feedback loop mode to Closed Loop. Doing this ensures that the stage will correct for any errors in its positioning [1].

    - Magnet plate as talent manually moves it to its highest position. Editors, this is the exact same shot as 3.4.1, so maybe it could be recorded once an used twice?

    - 60811_screen_capture_3_UPDATED_031219 – Video editors, please show from 1:13-1:34.

  - Click the Save button in the Save Parameters box to store these settings in the system [1].

    - 60811_screen_capture_3_UPDATED_031219 – Video editors, please show from 1:37-1:44.

  - Place a 24-well culture plate containing engineered heart tissues on magnetically responsive silicone racks on the culture plate mount [1].

    - Talent places a 24-well culture plate containing EHTs on magnetically responsive silicone racks on the culture plate mount.

    - Beating EHTs in a 24-well culture plate filmed from above –Editors, could we add this as an “optional” shot if we are able to provide a set of beating EHTs for filming (which we probably will)?

  - To calculate the magnet spacing necessary to achieve a desired afterload, solve the nonlinear regression function as described in the text protocol. Subtract dmin from the calculated magnet spacing d. The result is the distance the magnet plate has to travel from its Zero position to achieve the desired afterload. –Authors, are these calculations generally done in a lab notebook or on a computer through a spreadsheet program or some other program? If the latter, could you provide a screen capture movie of this process?                                                                        So far, we have solved the equations by hand in a lab notebook. Afterwards, we used a calculator to calculate the magnet spacing d. Maybe a still picture of the original function (as given by GraphPad) and the solved function could be shown during the narration of this step?

  - Type this value into the Target Position 1 input field in the Motion panel and click Go to adjust the engineered heart tissues’ afterload to the calculated value [1].

    - 60811_screen_capture_3_UPDATED_031219 – Video editors, please show from 1:52-2:14.

    - Magnet plate as it is moving to the target position.

OPTIONAL – Critical Step Statement:

- An OPTIONAL brief statement may be submitted for further elaboration of the best way to perform the required technique for the single most critical step of this procedure.

- If there is no single critical step, then there is no need to fill out this statement.

- This will be an interview style shot interjected after the relevant step within the Protocol section of the video.

- This statement is limited to 30 words or less.

- Please indicate the full name of the Author who will give this statement and the step of the protocol to which the statement pertains using the step numbers from the Protocol section (above).

# Fill in the details below based on the instructions above for the “Critical Step Statement”

# Author name, Step            :           (Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

# Section – Results

# The Results section is restricted to 200 words of narrative. Please read through the results as presented to make sure that it accurately represents your findings. If you would like to highlight other data, please revise this section accordingly, keeping in mind the word count restriction. Please note that we cannot include narrative without an accompanying visual.

- Results: Stepwise Afterload Increase and Interval Afterload Regimen Authors, feel free to change this title if you have something else in mind (please just keep it to <2 lines of text)

  - Control and MR-EHTs produced from rat hearts were cultured in the absence of magnetic afterload until a plateau in contractile force was reached [1-TXT].

    - Figure 7A Authors: Please provide a separate version of figure 7A without the “A” label. If you have flattened this figure, please submit it in its original file format as unflattened, layered images so that different components can be emphasized in the video. TEXT: MR-EHT = magnetically responsive-engineered heart tissues Editors: All demanded separate figures have been uploaded, named “Figure 7A unlabeled” etc.

  - On this day, MR-EHTs and control EHTs had similar mean forces [1]. Over the next week, the afterload exerted on MR-EHTs was incrementally increased from 0.91 to 6.85 milliNewtons per millimeter, while afterload for control EHTs remained constant [3].

    - Figure 7A – Video editor, please emphasize the point on the graph at 24 days of EHT culture on the x-axis.

    - Figure 7A – Video editor, please emphasize the red bars plotted from day 25 on.

  - Mean contractile force increased with increasing afterload up to 0.95 milliNewton [1], which marks more than a 3-fold increase in force compared to the average value measured for control EHTs [2].

    - Figure 7A – Video editor, please emphasize the blue plotted line from day 25 on.

    - Figure 7A – Video editor, please emphasize the black plotted line from day 25 on.

  - Post deflection, on the other hand, decreased compared to control tissues [1]. On the last day of culture, the mean deflection measured for MR-EHTs was only 0.11 millimeters… [2] compared to 0.48 millimeters for control EHTs [3].

    - Figure 7B Authors: Please provide a separate version of figure 7B without the “B” label. If you have flattened this figure, please submit it in its original file format as unflattened, layered images so that different components can be emphasized in the video. Editors: All demanded separate figures have been uploaded, named “Figure 7A unlabeled” etc.

    - Figure 7B – Video editor, please emphasize the blue plotted line from day 25 on.

    - Figure 7B – Video editor, please emphasize the black plotted line from day 25 on.

  - Rat EHTs on magnetically responsive silicone posts were cultured at a minimal afterload of 0.91 milliNewton per millimeter until a plateau in contractile force was reached [1].

    - Figure 8A Authors: Please provide a separate version of figure 8A without the “A” label. Editors: All demanded separate figures have been uploaded, named “Figure 7A unlabeled” etc.

  - From this day onward, MR-EHTs underwent a 7-day afterload regimen which exposed the EHTs to cycles of afterloads alternating between 0.91 and 6.85 milliNewton per millimeter [1].

    - Figure 8A Authors: Please provide a separate version of figure 8B without the “B” label. If you have flattened this figure, please submit it in its original file format as unflattened, layered images so that different components can be emphasized in the video. Editors: All demanded separate figures have been uploaded, named “Figure 7A unlabeled” etc.  Video editor, please emphasize the black curve.

  - The afterload of control EHTs was kept constant at 0.60 milliNewton per millimeter over the entire duration of culture[1]. The observed differences were not statistically significant [2].

    - Figure 8B - Video editor, please emphasize the blue bars.

    - Figure 8B - Video editor, please emphasize the black bars.

# Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

- Below are questions for statements that can be used to further emphasize the significance of your protocol. At least one statement is required.

- Each statement is limited to 30 words.

- Answer the questions in full sentences, as you will be expected to memorize and deliver the sentences as spoken interview statements during filming.

- Indicate the full name of the author who will give each Conclusion Interview statement.

# What is most important thing to remember when attempting this procedure? Please indicate the steps (e.g., 2.4., 2.5.) in the Protocol section this advice correlates to.

  - Author Name:

# Following this procedure, what other methods can be performed? What questions would these additional methods answer?

  - Benjamin Becker: Combined with optical contractility analysis, this method could be used to measure short-term contractile response to changing afterload in real time, which could be helpful for investigating physiological muscle properties.

# After its development, did this technique pave the way for researchers to explore new questions within a specific scientific field? If so, how?

  - Author Name: ____ (Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

# Are any of the reagents or instruments hazardous? If so, please use this interview statement to remind viewers of what precautions they should take.

  - Benjamin Becker: Strong magnets may suddenly cling to each other, potentially causing injury or damaging the magnets themselves. To avoid this, keep the magnets separated at a safe distance.

Thank you for following the instructions and addressing our questions. We will incorporate your answers/suggestions and send you the finalized script before your shoot. You will also receive detailed shoot preparation instructions in the email accompanying the finalized script.