# Submission ID #: 60861

# Scriptwriter Name: Bridget Colvin postshoot by Anastasia

Project Page Link: http://www.jove.com/files_upload.php?src=18562653

Title: Validated LC-MS/MS panel for Quantifying 11 Drug-Resistant TB Medications in Small Hair Samples

Authors and Affiliations: Andrew Reckers1, Anita Wen1, David Aguilar1, Peter Bacchetti2, Monica Gandhi3, John Metcalfe4, and Roy Gerona1

1Department of Obstetrics, Gynecology and Reproductive Sciences, TB Hair Analysis Laboratory, University of California San Francisco

2Department of Epidemiology and Biostatistics, University of California San Francisco

3Department of Medicine, Division of HIV, Infectious Diseases, and Global Medicine, University of California San Francisco

4Department of Medicine, Division of Pulmonary and Critical Care Medicine, University of California San Francisco

# Corresponding Author:

# Roy Gerona

# Roy.Gerona@ucsf.edu

Email addresses for Co-authors:

Andrew.Reckers@ucsf.edu

anitawensci@gmail.com

davidaguilartx@gmail.com

Peter.Bacchetti@ucsf.edu

Monica.Gandhi@ucsf.edu

John.Metcalfe@ucsf.edu

Author Questionnaire:

1. Microscopy: Does your protocol involve video microscopy? N

2. Does your protocol demonstrate software usage? Y

3. Which steps from the protocol section below are the most visually important?

2.1., 2.2., 3.2., 3.5.

4. What is the single most difficult aspect of this procedure and what do you do to ensure success?

2.1., 2.2. deserve extra highlighting because incorrectly collected samples are difficult to analyze.

5. Will the filming need to take place in multiple locations (greater than walking distance)? N

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

- REQUIRED Interview Statements (Said by you on camera): All interview statements may be edited for length and clarity.

  - John Metcalfe: As patient drug compliance and metabolism can be irregular and

# drug resistance is common, the objective measurement of drug levels in patients with MDR TB is important [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Monica Gandhi: If adherence to drug regimens can be more easily and accurately quantified, healthcare providers may be able to better design treatment plans to limit the development of drug resistance [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL Interview Statements (Said by you on camera): All interview statements may be edited for length and clarity.

  - Roy Gerona: We have developed the first multi-analyte mass spectrometry panel for analyzing multiple MDR TB medications within small hair samples in a single assay [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Roy Gerona: Our multi-drug panel requires only one, 2-milligram sample per patient and uses neither a solid phase extraction nor filtering step to prevent any loss of sample volume [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Monica Gandhi: Objective measures of drug exposure measured early in the treatment course could potentially alert clinicians to problems with individual drugs in the treatment regimen [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Andrew Reckers: Hair samples must be pulverized on the day of analysis for an accurate drug concentration level measurement [1].

    - INTERVIEW: Above Talent speaking the statement above in an interview-style shot, looking slightly off-camera

Ethics title card: (for human subjects or animal work, does not count toward word length total)

  - Procedures involving human subjects have been approved by the Institutional Review Board (IRB) at the University of Cape Town and the University of California, San Francisco.

# Section - Protocol

- Hair Sampling

  - To acquire hair samples, after obtaining written informed consent, get materials ready, including a thin label, two participant labels, a pair of scissors, an unfolded piece of aluminum foil, an alcohol wipe, and a plastic bag [1].

# 2.1.1 WIDE: Talent selecting scissors Videographer: Important/difficult step

  - Wipe the blades of the scissors clean with an alcohol wipe [1].

    - Blade being clean with an alcohol wipe.

  - Go to the back of the head and lift the top part of the hair, using a hairclip if desired. Gather 30 strands of hair from the back of the head and isolate the thatch with your hand. Take the scissors down to the scalp and cut the thatch of hair [1].

    - Hair being cut Videographer: Important/difficult step

  - Carefully put the thatch of hair down on the aluminum foil [1].

    - Thatch of hair being placed on the aluminum foil

  - Put the thin label over the distal end of the hair thatch [1], taping it securely to the aluminum foil. The distal end is the side furthest from the scalp [2].

    - Distal end being labeled

    - Tape being placed Videographer: Important/difficult step

  - Fold over the piece of aluminum foil on top of the hair thatch and put a participant label on top of the aluminum foil [1].

    - Hair being folded onto the aluminum foil Videographer: Important/difficult step

  - Put the folded-over piece of aluminum foil into the plastic bag [1] and put another participant label on top of the plastic bag	[2].

    - Aluminum foil being placed into the plastic bag

    - Plastic bag being labeled

- Drug Extraction

  - For drug extraction from the hair samples, weigh out 2 milligrams of each hair sample into individually labeled bead tubes [1-TXT] and 2 milligrams of the blank hair sample into each of 16 additional bead tubes [2].

    - WIDE: Talent adding bead tube to balance TEXT: Trim hair strands to 3 cm from proximal end as necessary

    - Talent adding blank hair to one of 16 tubes, with all 16 tubes visible in frame

  - After weighing, place all of the bead tubes into a bead mill homogenizer [1] and homogenize the samples at a speed of 6.95 meters/second for two 30-second cycles with a 15-second rest period between each cycle [2].

    - Talent placing tube(s) into homogenizer Videographer: Important step

    - Samples being homogenized OR Homogenizer settings be set Videographer: Important step

  - After the second homogenization, add 500 microliters of the appropriate internal standard mix to each sample [1-TXT] and add 500 microliters of methanol to the matrix blank tube [2].

    - Talent adding standard mix to tube(s), with internal standard mix containers visible in frame TEXT: See text for all standard mix preparation details

    - Methanol being added to tube, with methanol container visible in frame

  - Spike the calibration curve tubes with the appropriate volume of reference standard mix according to the Table [1] and place all of the tubes in a 37-degree water bath with gentle shaking for 2 hours [2].

    - LAB MEDIA: Table 3 Video Editor: please emphasize Volume added column

    - Talent placing tube(s) into water bath

  - At the end of the incubation, transfer the supernatant from each bead tube into new, labeled microcentrifuge tubes [1] and add 500 microliters of methanol to the bead tubes [2].

    - Supernatant being added to tube(s), with new tube labels visible in frame Videographer: Important step

    - Methanol being added to bead tubes Videographer: Important step

  - Vortex the bead tubes [1] and transfer the methanol wash to the corresponding microcentrifuge tubes [2-TXT].

    - Tube(s) being centrifuged

    - Wash being added to tube(s) TEXT: Ok to transfer pulverized hair

  - Centrifuge the samples [1-TXT] and carefully transfer the supernatants into new tubes without disturbing the hair pellets [2].

    - Talent placing tube(s) into centrifuge TEXT: 10 min, 2800 x g, RT

    - Talent adding supernatant(s) to tube

  - Evaporate the liquid in the tubes to dryness at 32 degrees Celsius [1] and reconstitute the samples with 200 microliters of high-pressure liquid chromatography-grade water with 1% formic acid [2].

    - Talent evaporating tubes/placing tubes to evaporate

    - Mobile phase A being added to tube(s), with mobile phase A container visible in frame

  - After vortexing, transfer the liquid samples to amber vials with 250-microliter inserts [1].

    - Talent adding sample(s) amber vial(s)

- Liquid Chromatography Tandem Mass Spectrometry (LC-MS/MS) Preparation

  - For liquid chromatography-tandem mass spectrometry, first install a 2- x 100-millimeter column with a 2.5-micrometer-particle size and a 100-angstrom-pore size with polar endcapped, ether-linked, phenyl beads fully made of porous silica into the column compartment [1].

    - WIDE: 15-20 s Talent loading column onto instrument

  - Confirm that the column also has a manufacturer recommended guard cartridge installed [1].

    - Shot of guard cartridge installed in column

  - Copy and paste the compound information included in the MDR-TB LCMS (M-D-R-T-B L-C-M-S) method transitions.csv file [1-TXT] into the box in the Mass Spec window [2].

    - Talent at computer, copying information, with monitor visible in frame

    - SCREEN: screenshot_1: 00:04-00:15 TEXT: MDR-TB LCMS method transitions.csv included in Supplemental Materials

  - In a spreadsheet, create a batch sequence of the calibration curve, quality controls, patient samples, calibration curve, quality controls, patient samples, calibration curve, and quality controls [1].

    - SCREEN: screenshot_2: 00:03-00:04

  - Add solvent blank injections at the start and end of the run as well as before and after the calibration curve, quality controls, and patient samples [1].

    - SCREEN: screenshot_2: 00:05-00:10 Video Editor: please emphasize Solvent Blank-Inj data cells (from 39-63 in spreadsheet)

  - Then place at least eight solvent blank injections after the injections of the calibration curve [1] and a high-quality control vial to reduce the analyte carryover [2].

    - SCREEN: SCREEN: screenshot_2: 00:10-00:12 Video Editor: please emphasize Solvent Blank data cells (from 78-85 in spreadsheet)

    - Talent placing vial

- Data Analysis

  - When the batch is complete, open the quantitation software [1].

    - WIDE: Talent opening software, with monitor visible in frame

  - Scroll through the transitions to ensure that the automatically selected retention time is accurate and check that the Gaussian Smoothing is set to 1.5 [1-TXT].

    - SCREEN: screenshot_4: 00:03-00:26 Video Editor: can speed up or not show entire clip; please emphasize Gaussian Smoothing value when mentioned TEXT: Integrate all injections using same parameters

  - Next, click Displays the calibration curve and Regression and set the Weighting Type to 1/x [1].

    - SCREEN: screenshot_5: 00:02-00:08

  - Click OK and validate the calibration curve and quality control samples to assure that the batch ran successfully. For each quantifier reference analyte, at least two-thirds of the calibration points must have an accuracy within 80-120% [1-TXT].

    - SCREEN: screenshot_5: 00:09-00:40 Video Editor: can speed up or not show entire section; Video Editor: please emphasize Accuracy column when mentioned TEXT: Click set peak to not found to exclude outliers

  - Check that the R-value displayed above the calibration curve is >0.975 [1].

    - SCREEN: screenshot_5: 00:40-00:45

  - If all of the above conditions are satisfied, the batch has passed, and the samples can be quantified. Click Edit in the toolbar and click Copy entire table [1].

    - SCREEN: screenshot_5: 00:45-00:53

  - Then paste the table into a spreadsheet and use the average of the calculated concentration of the two sample injections to determine the reported concentration of each sample [1].

    - SCREEN: screenshot_6: 00:04-00:11

# Section – Results

- Results: Representative Drug-Resistant (DR) TB Medication Quantification

  - Here an illustration of a chromatogram with confirmed levels of all tested 11 drug resistant-TB drugs is shown [1].

    - LAB MEDIA: Figure 1 Video Editor: please sequentially add/emphasize data peaks

  - In these figures, the extracted ion chromatograms for one particular drug [1] in one of the calibrators can be observed [2].

    - LAB MEDIA: Figure 2 Video Editor: please emphasize top graph

    - LAB MEDIA: Figure 2 Video Editor: please emphasize bottom graph

  - The quantifier [1] and qualifier transitions [2] are used to qualitatively confirm the presence of the drug, as the ratio between the area of the quantifier and the area of qualifier remains constant across samples [3].

    - LAB MEDIA: Figure 2 Video Editor: please emphasize blue line in top graph

    - LAB MEDIA: Figure 2 Video Editor: please emphasize red line in top graph

    - LAB MEDIA: Figure 2

  - The internal standard is also monitored to ensure that each sample injection is normalized [1].

    - LAB MEDIA: Figure 2 Video Editor: please emphasize blue line in bottom graph

  - In this analysis, a convenience sample of 15 hair samples among a total study population of 96 patients taking drug resistant-TB drugs under directly observed therapy conditions from Western Cape, South Africa, was evaluated [1] to determine representative levels of drug resistant-TB drugs across the lowest and highest levels for each analyte [2].

    - LAB MEDIA: Table 6

    - LAB MEDIA: Table 6 Video Editor: please emphasize Sample values column

# Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Andrew Reckers: In the absence of a solid phase extraction or filtering step, take care when transferring the supernatant into evaporation tubes, as anything transferred may end up injected into the column [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera (Step: 3.7.)

  - Roy Gerona: As the rotor of the bead disruptor can fly off if it is not tightly screwed and locked, be sure to always double check the rotor before operating [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - John Metcalfe: In contrast to HIV drugs, the relationship between TB medication pharmacokinetics, adherence, and clinical outcomes remains largely unstudied [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Roy Gerona: We have hypothesized that some TB drugs do not accumulate well in hair. Targeting the metabolites of these drugs may provide a more complete understanding of patient adherence [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera