## Sheet1
Filename | Step(s) | Content | Time code
61002_screenshot_1 | 3.1. | Using the micromanipulators, camera and system controller software (Figure 7A, see Table of Materials) to measure sarcomere length. Move the piezo and/or force probe to set the initial sarcomere length of the myofibril to 2.5 µm. | 00:03 - 00:54
| NOTE | A sarcomere length of 2.5 µm ensures optimal overlap between myosin heads and actin.
| 3.2. | Measure myofibril dimensions: Using the vessel function of the system controller software measure length and width (Figure 7B, C).
| NOTE | When rotating the camera, it may tilt horizontally and/or vertically. To check the alignment of the camera, a spirit level can be used to verify that the camera is rotated, and not tilted.
| 3.2.1. | Position the myofibril in the center of the video image using the microscope stage.
61002_screenshot_2 | 3.2.2. | Draw a square from one side of the myofibril to the other side of the myofibril. For the length, make sure to include the dark edge of the glue droplets (Figure 2A) in the square since the image processing is based on contrast. | 00:02 - 00:06
61002_screenshot_2 | 3.2.3. | Start recording the data in the system controller software (see Table of Materials) by pressing ‘Start’ and after 5 s pause the system controller software data recording by pressing the ‘Pause’ button. The length is now recorded in the data. | 00:10 - 0:29
61002_screenshot_2 | 3.2.4. | For the width first rotate the camera (see Table of Materials) 90˚ and then use the contrast of the edge of the myofibril itself. | 00:35 - 01:24
61002_screenshot_2 | 3.2.5. | Start recording the data in the system controller software (see Table of Materials) by pressing ‘Start’ and after 5 s pause the system controller software data recording by pressing the ‘pause’ button. The width is now recorded in the data. | 01:26 - 01:44
61002_screenshot_3 | 3.4.5. | Align the top channel of the theta glass with the myofibril using the manipulator and check the position by performing a fast-step (signal generator settings can be found in Table 1) with the system controller software (Figure 2B-C, see Table of Materials). | 00:05 - 00:24
| NOTE | Make sure that the bottom channel will be aligned with the myofibril during the activation phase of the fast-step (Figure 2B-C).
| 3.5. | Turn on background flow of Rx (Figure 4A) to create a laminar background flow in the flow chamber:
| NOTE | The background flow is necessary to prevent turbulent flow as a result from the pCa solution flow from the Ɵ-glass.
| 3.5.1. | Turn on inflow of flow chamber with Luer valve lever.
61002_screenshot_4 | 3.5.1.1 | Send the following parameters to outflow pump to start drain the flow chamber and prevent overflowing of the flow chamber (Figure 9): Valve: Bath valve (2); Microstep mode: Micro; Plunger target: 48000; Plunger speed: 38-40 (arbitrary). | 00:05 - 00:22
61002_screenshot_5 | 4.1.3.1. | Calculate the piezo movement necessary to slacken the myofibril 15 % and enter this value in the signal generator (Figure 5D, Table 1). | 00:03 - 00:10
| 4.1.3.2. | Start recording the data in the system controller software by pressing ‘Start’.
61002_screenshot_7 | 4.1.3.3. | Open valve 1 and 6 with the data acquisition panel (Figure 6A), software to start flow of relaxing solution and pCa 4.5 through the theta glass. | 00:08 - 00:12
| 4.1.3.4. | Reset range of the interferometer so that the baseline force will be at 0 V, by selecting and pressing ‘Reset range’ on the interferometer.
61002_screenshot_6 | 4.1.3.5. | When the force trace is stable, perform the theta glass fast-step (step size: 100 µm). Signal generator settings can be found in Table 1. | 01:02 - 01:17
61002_screenshot_6 | 4.1.3.6. | When the force plateau is reached, perform the shortening-restretch with the piezo. Signal generator settings can be found in (Figure 5D, Table 1). An activation-relaxation trace similar to Figure 4E will be recorded and visible in the system controller software. | 01:08 - 01:11
| NOTE | A custom protocol can be made to automate the steps above.
61002_screenshot_6 | 4.1.3.7. | Pause the system controller software by pressing the ‘pause’ button.
| 4.2.2.1. | Start recording the data in the system controller software by pressing ‘Start’.
| 4.2.2.2. | Reset range of the interferometer so that the baseline force will be at 0 V, by selecting and pressing ‘Reset range’ on the interferometer.
61002_screenshot_6 | 4.2.2.3. | Perform a stepwise stretch with the signal generator in the system controller software to operate the piezo. Example signal generator settings can be found in Table 1 (Figure 5E). | 02:46 - 03:18
61002_screenshot_6 | 4.2.3. | Shorten the myofibril to slack length with the piezo after the stretch is finished. Example signal generator settings can be found in Table 1. | 03:18 - 03:35
61002_screenshot_6 | 4.3. | Pause the system controller software by pressing the ‘pause’ button. | 03:36 - 03:38
61002_screenshot_6 | 4.4. | Stop the recording of data by pressing the ‘Stop’ button in the system controller software. | 03:38 - 03:40
61002_screenshot_6 | 4.5. | Save the data by pressing ‘File’ and ‘Save data’ in the system controller software. | 03:40 - 03:54