# Submission ID #: 61007

# Scriptwriter Name: Bridget Colvin

Project Page Link: http://www.jove.com/files_upload.php?src=18607183

Title: Terahertz Imaging and Characterization Protocol for Freshly Excised Breast Cancer Tumors

Authors and Affiliations: Nagma Vohra1,*, Tyler Bowman1,*, Keith Bailey2, and Magda El-Shenawee1

*These authors contributed equally to the work

1Department of Electrical Engineering, University of Arkansas

2Oklahoma Animal Disease Diagnostic Laboratory, Oklahoma State University

# Corresponding Author:

# Nagma Vohra

# nvohra@email.uark.edu

Email addresses for Co-authors:

tcbowman@uark.edu

kbailey1@illinois.edu

magda@uark.edu

Author Questionnaire:

1. Microscopy: Does your protocol involve video microscopy? N

2. Does your protocol demonstrate software usage? Y

3. Which steps from the protocol section below are the most visually important?

2.3., 3.10., 4.2.

4. What is the single most difficult aspect of this procedure and what do you do to ensure success?

2.3., 4.2. The single most difficult aspect of this procedure is tissue handling. Poor tissue contacts with the imaging window leading to presence of air bubbles at the tissue-polystyrene interface, and that creates artifacts of low reflection in the THz image. Extra effort should be taken to ensure good tissue contact with the polystyrene window by possibly repositioning the tissue.

5. Will the filming need to take place in multiple locations (greater than walking distance)? N

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

- REQUIRED Interview Statements (Said by you on camera): All interview statements may be edited for length and clarity.

  - Magda El-Shenawee: This protocol uses terahertz imaging and spectroscopy to evaluate the surgical margins of breast tumors with the aim of substantially decreasing the need for a second surgery [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Magda El-Shenawee: Terahertz imaging can be used to distinguish between cancerous and non-cancerous tissues based on their optical properties in a nonionizing and biologically safe manner [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Nagma Vohra: This technique was developed for breast cancer margin assessment but is applicable to other solid tumors that require surgical excision and margin assessment [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera Videographer: Can cut for time

  - Nagma Vohra: This method allows imaging of the inside of FFPE tissue blocks without the need for physical slicing, facilitating the identification of tumor margins at any depth [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera Videographer: Can cut for time

  - Tyler Bowman: This protocol requires the performance of critical steps, such as tissue handling and stage balancing, that can be difficult to understand without visual demonstration [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera Videographer: Can cut for time

Ethics title card: (for human subjects or animal work, does not count toward word length total)

  - Procedures involving animal subjects have been approved by the Institutional Animal Care and Use Committee (IACUC) at the University of Arkansas. Human tissues were used with approval from the Institutional Biosafety Committee (IBC) at the University of Arkansas.

# Section - Protocol

- Tissue Handling

  - Before handling the tissue, cover a stainless-steel metal tray with a biohazard bag [1] and place all of the appropriate lab materials within easy reach of the covered tray [2].

    - WIDE: Talent covering tray with bag

    - Talent placing tools next to tray TEXT: Keep any used tissues/wipes/gloves on biohazard material surface until disposal

  - Next, transfer the fresh tumor sample into a Petri dish on the tray [1] and use gross inspection to select a side of the tumor by that is sufficiently flat and has little blood and few blood vessels [2].

    - Talent placing tissue into dish

    - Tumor being manipulated/examined

  - Then place the side to be imaged on grade 1 filter paper to remove any excess medium and to clear the tissue of fluid or secretions [1], repositioning the tumor to a dry spot as the paper saturates [2].

    - Tumor being placed onto filter paper Videographer: Important/difficult step

    - Tumor being moved to dry spot

- Terahertz (THz) Imager Setup

  - After drying the tumor for 5 minutes, place a 1.2-millimeter thick polystyrene plate with an approximately 37-millimeter diameter onto the scanning window [1] and place the scanning window and the polystyrene plate onto the sample stage of the terahertz imager [2].

    - WIDE: Talent placing plate into scanning window

    - Window being placed onto sample stage

  - In the main window, click the Fixed-point Scan icon to activate the terahertz antennas and to start sending and receiving the reflected terahertz signal from a single point on the polystyrene plate [1].

    - SCREEN: screenshot_1

  - Click the Motor Stage Dialog icon. The motor control window will open up [1].

    - SCREEN: screenshot_2: 00:00-00:04

  - To center the reflected pulse from the polystyrene in the main window, click the Optical Density Axis arrows to adjust the optical delay axis [1].

    - SCREEN: screenshot_2: 00:07-00:15

  - Click the Data Acquisition Settings button to open the data acquisition settings dialog window and change the optical delay value from 5 to 4 volts [1].

    - SCREEN: screenshot_3: 00:00-00:09

  - Adjust the vertical position of the scanning stage with the micrometer scale [1] until the minima of the secondary pulse is the strongest [2] and adjust the optical delay of the axis in the Motor Control Window to put the primary reflection outside of the range of the reflected signal being measured [3].

    - Stage being adjusted

    - SCREEN: screenshot_3: 00:30-40

    - SCREEN: screenshot_3: 00:47-01:51 Video Editor: please speed up

  - To level the sample stage and record the reference signal, click the Motor Control Dialog button to open the motor control window and reposition the motor control and main software windows so that the time domain signal is visible while adjusting the motor positions [1].

    - SCREEN: screenshot_4: 00:00-00:04

  - To level the A-axis using following steps, in the Motor Control Window, change the value of the A-axis from 0 to -10 and hit Enter [1].

    - SCREEN: screenshot_4: 00:04-00:09

  - The stage will move to the minus 10-millimeter position on the A-axis [1] and a shift in the signal position will be observed in the main window [2].

    - Stage moving

    - SCREEN: screenshot_4: 00:09-00:15

  - Use the adjustable micrometer scale to move the minimum peak of the signal back to the previous vertical position [1] and change the A-axis value to plus 10. Click Enter [2].

    - Stage being moved Videographer: Important step

    - SCREEN: screenshot_4: 00:16-00:24

  - The stage will move to the plus 10-millimeter position on the A-axis [1] and a shift in the signal will be observed again [2].

    - Stage moving

    - SCREEN: screenshot_4: 00:25-00:27

  - Note the direction and distance that the signal shifted from its previous position and change the A-axis value back to minus 10. The signal will return to the original vertical position [1].

    - SCREEN: screenshot_4: 00:27-00:38

  - Rotate the leveling screw on the A-axis of the scanning stage [1] and shift the signal to double the distance in the same direction the stage moved from the original position [2].

    - Screw being rotated

    - SCREEN: screenshot_4: 00:39-00:45

  - Use the micrometer on the scanning stage to shift the signal back to the original position [1] and repeat the adjustment until the signal at plus 10 and minus 10 are equal and the peak for both positions is focused at the original position [2].

    - SCREEN: screenshot_4: 00:45-00:51

    - SCREEN: screenshot_4: 03:13-00:32

  - Once the leveling of the A-axis has been achieved, change the A-axis value to 0 [1] and repeat the same procedure for the B-axis [2].

    - SCREEN: screenshot_5: 00:00-00:02

    - SCREEN: screenshot_5: 00:03-00:13

  - Once both axes have been leveled, return both the A- and B-axes to 0 millimeters and close the Motor Control Window [1]. Then verify that the signal is in its original position in case it has shifted a little [2].

    - SCREEN: screenshot_5: 01:15-01:28

    - SCREEN: screenshot_5: 01:28-01:30

  - To record this signal as the reference, in the Data Acquisition Properties window, change the averaging value to 5 and leave all of the other parameters at their default settings [1].

    - SCREEN: screenshot_6: 00:00-00:09

  - Then click New Reference. The averaging counter will count from 0-20. Once the counter reaches 20, change the averaging value to 1 and click OK. The reflected signal from the polystyrene will be saved as the reference for any subsequent scans [1].

    - SCREEN: screenshot_6: 00:10-00:38 Video Editor: can speed up

- THz Tumor Imaging

  - When all of the parameters have been set, transfer the imaging window from the scanning stage to the tissue handling area [1] and mount the tumor onto the polystyrene plate [2].

    - WIDE: Talent placing plate in tissue handling area

    - Tumor being placed onto plate

  - Remove any air bubbles within the tumor with tweezers [1] or gently roll the sample onto the polystyrene until the air gaps are minimized [2].

    - Bubbles being removed Videographer: Important/difficult step

    - Sample being rolled Videographer: Important/difficult step

  - Place absorptive spacers at regular intervals around the test sample [1] and place another polystyrene plate above the tumor [2].

    - Spacer(s) being placed

    - Plate being placed

  - Gently press the tumor surface as flat as possible [1] and tape down the polystyrene-tumor-polystyrene arrangement on the sample window [2].

    - Plate being pressed

    - Setup being taped

  - Tyler Bowman: A flat tissue surface and a good contact with the polystyrene plate are vital, as air bubbles, excess fluid, and an uneven surface can ruin the tumor tissue imaging [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Flip the sample window [1] and take photos of the tumor to keep a record of its orientation [2].

    - Talent flipping window

    - Talent photographing tumor

  - Return the sample window to the scanning stage [1] and click the Image Parameter Dialog button to open the Image Acquisition Parameters window [2].

    - Talent placing window onto stage

    - SCREEN: screenshot_7: 00:00-00:04

  - Set the values of Axis1 minimum, Axis1 maximum, Axis2 minimum, and Axis2 maximum to fully enclose the position of the tumor in the imaging window [1].

    - SCREEN: screenshot_7: 00:04-00:19

  - Set Axis1 step and Axis2 step to 0.2 millimeters for the imaging scan [1].

    - SCREEN: screenshot_7: 00:19-00:24

  - Then, under the Measure menu, select Flyback 2D Scan and create a folder and file name under which to save the scan data in the pop-up window [1].

    - SCREEN: screenshot_7: 00:34-01:10 Video Editor: please speed up

# Section – Results

- Results: Representative Tumor Sample Imaging and Characterization

  - In this analysis, a grade 1-2 infiltrating ductal carcinoma obtained from a 49-year-old woman via a left breast lumpectomy surgery procedure was assessed [1].

    - LAB MEDIA: Figures 9A and 9B

  - Upon correlating the terahertz image with the pathology image [1], it was clear that the cancer region [2] exhibited a higher reflection than the fat region [3].

    - LAB MEDIA: Figure 9C

    - LAB MEDIA: Figure 9C Video Editor: pleasee emphasize red region

    - LAB MEDIA: Figure 9C Video Editor: pleasee emphasize blue region

  - Tomographic imaging revealed that, as the frequency increased [1], the calculated absorption coefficient values for the cancer and fat pixels increased [2], with the cancer pixels showing higher values than the fat at both frequencies [3].

    - LAB MEDIA: Figures 9D and 9E Video Editor: please emphasize Absorption coefficient legend

    - LAB MEDIA: Figures 9D and 9E Video Editor: please emphasize green area in Figure 9D and red/orange area in Figure 9E

  - In contrast, the refractive index of both tissues decreased [1] as the frequency increased [2].

    - LAB MEDIA: Figures 9F and 9G Video Editor: please emphasize images in Figures 9F and 9G

    - LAB MEDIA: Figures 9F and 9G Video Editor: please emphasize Refractive index legend

  - Transmission spectroscopy analysis of the same tumor [1] reveals a good agreement for the frequency range of both the extracted absorption coefficient and refractive index [2] for both sections extracted from the tumor [3].

    - LAB MEDIA: Figure 12

    - LAB MEDIA: Figure 12 Video Editor: please emphasize graphs

    - LAB MEDIA: Figure 12 Video Editor: please emphasize 1 and 2 circles in Figure 12A

  - Note that insufficient handling of the tissue can lead to misleading imaging results [1], resulting, for example, in a larger presence of cancer in the tumor being suggested [2] or in the prescence of excess liquid within the tissue in undried samples [3].

    - LAB MEDIA: Figures 10A and 10B

    - LAB MEDIA: Figures 10A and 10B Video Editor: please emphasize green/yellow cancer signal in Figure 10B

    - LAB MEDIA: Figures 10A and 10B Video Editor: please emphasize red DMEM signal in Figure 10B

# Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Nagma Vohra: Additional characterization can be performed on the THz image data to obtain information about the frequency-dependent properties of the tissue and statistical analysis can be performed for automated image segmentation [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Tyler Bowman: This technique has led to the development of imaging and spectroscopy algorithms for investigating the signatures of ductal carcinoma in situ to differentiate low-grade tumors from deadly high-grade tumors [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera Videographer: Can cut for time

  - Magda El-Shenawee: Remember to handle cancerous tissues and formalin solution with caution and that any system elements or tools that contact the tumor tissue must be cleaned or discarded appropriately [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera Videographer: Can cut for time