JoVE61031R1 "Simultaneous Quantification of Selected Kynurenines Analyzed by Liquid Chromatography-Mass Spectrometry in Medium Collected from Cancer Cell Cultures "

We thank all Reviewers for their constructive comments. Our changes within the revised manuscript are marked in red.

Editorial comments:

1. Please copy-edit the manuscript as the language is not publication grade.

The manuscript has been again reviewed very carefully to address all grammatical errors. In addition, an experienced scholarly writer edited the manuscript. The major changes are marked in red.

2. Please revise the highlighting to be 2.75 pages or less. Please highlight the entire steps to ensure that a complete and coherent story is told. We need to know the actions and the details required to perform that action in order to film.

The action has been highlighted in yellow and constitutes no more than 2.75 pages.

3. Please address all the specific comments marked in the manuscript. Please use the attached version for revision. This is formatted to match the journal's style.

All comments have been addressed.

4. Please obtain explicit copyright permission to reuse any figures from a previous publication. Explicit permission can be expressed in the form of a letter from the editor or a link to the editorial policy that allows re-prints. Please upload this information as a .doc or .docx file to your Editorial Manager account. The Figure must be cited appropriately in Figure Legend, i.e. “This figure has been modified from [citation].”

We did not copy the same results that are published in our JPBA paper, but used the chromatogram of the other repeat, despite that the chromatogram looks very similar. In addition we have added citation for the original JPBA paper.

Reviewers' comments:

Reviewer #3:

Minor Concerns:

There is no comparison of this method with the previous established methods, in particular liquid chromatography with ultraviolet or fluorescent detection. A comparison of peaks obtained via the current described methods from MDA-MB-231 cells and SK-OV-3 cells, with peaks obtained using these prior methods, and a brief description of the observations (what are discrepancies? how does the peak quality compare?) would help strengthen the manuscript.

Answering this question, we would like to emphasize that it was not our goal in this paper to compare different methodologies that can be employed for kynurenine (Kyn) quantification. Similarly, the fluorescence-based quantification was not a subject of our work. However, in our other studies, we have employed liquid chromatography coupled with ultraviolet detection (DAD as detector) for Kyn determination. Unfortunately, comparison of the results obtained using LC-MS and HPLC-DAD is hampered by several methodological discrepancies like gradient programs, applied mobile phases, retention times of the same metabolites, as well as difficulties to separate some interfering co-eluting compounds present in the material derived from living cells. In addition to other studies, our recently published results (Sadok et al., Simultaneous voltammetric analysis of tryptophan and kynurenine in culture medium from human cancer cells, Talanta 209 (2020) 120574, DOI: 10.1016/j.talanta.2019.120574) show that LC-MS allows for detection and quantification of lower amounts of Kyn in comparison to LC-DAD method. Determination of Kyn generated by MDA-MB-231 cells was difficult by LC-DAD (detected, but under LOQ), whereas LC-MS was suitable for quantification of the same metabolite in medium from cultured cells. Other metabolites (3-hydroxykynurenine, 3-hydroxyantrahilic acid, xanthurenic acid) were not thoroughly studied in cell cultures using LC-DAD thus we do not have enough data to comment on this topic.

Reviewer #4:

It seems that this version of the manuscript is a revised version of a previous one (?) but it was difficult to understand which changes and improvements have been made.

Correct, this is a revised version of the manuscript. The changes introduced in response to previous reviewer’s comments have been marked in red in revised version.

Major Concerns:

1. the internal standard used, 3-nitrotyrosine (3NT), is -as stated by the authors- the main limitation of this protocol. Tyrosine nitration is a chemical modification that occurrs during oxidative stress conditions. Certain types of cancers induce oxidative stress and thus 3-NT formation, although it is not clear to what extent. The use of 3-NT as the internal standard limits the flexibility of the presented protocol and potentially introduces a bias in the quantification of the analytes. I agree that the use of ILIS is not always necessary, although it represents the best internal standard in MS detection. However, the authors could (and should) have selected other commercially available molecules that have similar chromatographic properties and are not produced endogenously by mammalian cells.

During method development we have examined the tested samples (medium collected after culturing cancer cells (MDA-MB-231 and SK-OV-3) for presence of the endogenous 3-NT. There was no detected trace of this compound. Furthermore, endogenous 3-NT is present in cell culture medium at negligible amount (Teixeira et al., Development of a new HPLC-based method for 3-nitrotyrosine quantification in different biological matrices, Journal of Chromatography B 1046 (2017) 48-57, DOI: 10.1016/j.jchromb.2017.01.035)

compared to the exogenously added 3-NT thus it will have no impact on quantification. We emphasized in line 611 of the manuscript that endogenous 3-NT level should be initially verified. In our future path, we will consider testing other types of internal standards.

2. This manuscript represents the "step-by-step" version of a protocol published elsewhere by the same authors. However, the protocol here seems to be a simplified version of the published one, which is odd considering the scope of JoVE. I would strongly encourage the authors to describe the entire protocol providing more details. For general interest, it would be better to include a short explanations of why a certain reagent or procedure was chosen (e.g. different extraction protocols, protein precipitation procedures, LC buffers and so on). Moreover, there is limited information regarding the cell culture procedures, including duration of incubation, need for LC-SQ data normalization (e.g. determination of protein content), etc.

During the method development each step of the LC-MS analysis was carefully optimized and is described in our previous paper (Sadok et al., Application of the optimized and validated LC-MS method for simultaneous quantification of tryptophan metabolites in culture medium from cancer cell, Journal of Pharmaceutical and Biomedical Analysis 176 (2019) 112805). The rationale for selecting the critical conditions of analysis (LC buffers, reagent for protein precipitation) has been also explained there. Instead of copying the previous article, our goal was to describe handy protocol of kynurenines determination in culture medium from cancer cells. Thus, inhere we focused on the practical aspects providing some tips and recommendations how to perform the analysis. Considering the Reviewer’s suggestion, we decided to add to the protocol a description of cell culture conditions, protein determination and data normalization (point 7 and 9 in the protocol). In addition, a brief summary on selection of solvent for protein removal and LC mobile phase composition have been included (DISCUSSION, paragraph 3).

3. There is no explanation of why charcoal treatment is performed. I assume it is done to remove tryptophan (Trp) and its catabolites from the culture medium used (please, specify). Removing Trp from the culture media represents a big challenge for any kind of mammalian cell, and, among other things, will modify the whole kynurenine pathway activity. It is likely that this protocol will be used to study the activity of the kynurenine pathway in different cells types or how it is modulated by different challenges/treatments. To this end, it would be more suitable to use a Trp-containing media. Charcoal treatment could be applied to FBS, which potentially introduces variable amounts of kynurenines, or FBS could be replaced with BSA (mostly for short incubations).

The charcoal treatment was used only to prepare the calibration standards and quality control samples. The experimental samples were not treated with charcoal. Charcoal treatment was applied to remove the tryptophan catabolite - kynurenine (Kyn) as it was found in the initial culture medium (DMEM-HG supplemented with glucose and FBS) and could impact proper calibration. The second paragraph of the 'REPRESENTATIVE RESULTS' section discuses this issue. We have also provided Figure 1, which compares results obtained for the complete and charcoal treated culture medium. Definitely, during charcoal treatment the amount of other compounds like tryptophan are significantly reduced or removed. Nevertheless, as it was highlighted, charcoal purification was used only for preparation of calibration curves. In case of experimental samples, the complete medium (without purification) used for cell culturing was also analyzed, and the determined Kyn was subtracted from the amount measured in medium collected from the cultured cells. This information can be found in the 'REPRESENTATIVE RESULTS' section (at the end of the second paragraph).

The authors detect an unknown peak that is probably Trp (fig 3) but it is unclear how they concluded that this peak corresponds to Trp? The retention time of compound X (Fig 3) and Trp (Fig 4) seem different enough to suggest that these two are different molecules. In addition, the retention time of XA, present in both figures, is more or less the same, suggesting that the retention shift between compound X and Trp is not due to an LC issue. It is also surprising that the concentration of Trp is so much higher in the medium of one cell line vs the other? Is the assumption that this is all due to protein catabolism occurring in MDA-MB-231 cells? In figure 1 you show the analysis of the medium used before and after charcoal treatment, but there is no trace of Trp (or compound X, if they correspond). How can this be explained? A standard DMEM medium contains more than 10 mg/L of Trp. This point should be clarified. I would recommend including Trp in this analysis since its standard is easily accessible and inexpensive. Moreover, this would give the possibility to draw some conclusions regarding IDO/TDO activity (as suggested by the authors in the conclusions section).

This is correct that MDA-MB-231 and SK-OV-3 cells utilize Trp at different pace. It was confirmed in our other work (Sadok et al., Simultaneous voltammetric analysis of tryptophan and kynurenine in culture medium from human cancer cells, Talanta 209 (2020) 120574, https://doi.org/10.1016/j.talanta.2019.120574), where we attempted to determine Trp and Kyn content in medium from MDA-MB-231 and SK-OV-3 cultured cells. We found that MDA-MB-231 utilizes Trp in smaller extent compared to SK-OV-3 cells resulting in lower concentration of Kyn. Furthermore, our observations of standard DMEM using LC-DAD analysis (unpublished data) always showed the signal from Trp. It is not surprising, since Trp is one of the culture medium ingredients. In the present paper employing LC-MS to study tryptophan metabolism via kynurenine pathway the LC-MS system was set to monitor ions selected for individual metabolites at the defined timeframe. Thus, Trp could not be observed on chromatogram since we did not include in LC-MS settings the ions unique for Trp. We verified once again the LC-MS results obtained during DMEM analysis from different experiments. Below we present in Fig. R1 another example of chromatogram of DMEM prepared according to herein described LC-MS protocol (untreated with charcoal, spiked with 3NT). We do not see a significant peak that might correspond to Trp (probably it is caused by apparatus settings that in this case did not allow for recording of the characteristic for Trp ions). We agree with Reviewer that we jumped into conclusions about compound X too fast. We removed that assumption from the text since it is not clear what the compound X is. This issue will be interesting to dwell on in the future, however it is not a subject of this paper.

Fig.R1. Example of LC-MS result obtained during analysis of complete DMEM used for MDA-MB-231 and SK-OV-3 culture spiked with 3NT (untreated with charcoal).

Since Trp was not included in our protocol it is not possible to add this compound to our analysis without extensive methodological adjustments. Simultaneous analysis of several compounds is more complicated than analysis of just one analyte. The addition of this analyte will require re-adjustment of the method settings including evaluation of Trp ionization in the selected mobile phase or assessment of different mobile phase for simultaneous analysis of 3HKyn, Kyn, 3HAA, XA, Trp and 3NT (as internal standard). Also sample preparation step might differ, e.g. significant loss of Trp might be observed during protein precipitation using the selected solvent for 3HKyn, Kyn, 3HAA, XA. Moreover, the final protocol (after Trp addition) should be validated in terms of linearity, precision, accuracy, and recovery to assess whether the method provides reliable results. It might be the goal of a new project to expand on additional analytes.

4. Figure 2A looks rather similar to figure 3 in ref. 22. Is this in line with JoVE's publication policies? Otherwise permission from the ref 22 journal should be obtained.

All figures prepared for the JoVE were not used elsewhere, despite a fact they might look very similar. However, to comply with the JoVE policy we included the citation in the figure legend along with a statement that the figure has been modified from Sadok et al [27].

Minor Concerns:

1. There are some steps of the protocol that are too specific to a given instrument or lab setting (e.g. automatic pipette, evaporator, use of an autosampler, …). Unless these points/instruments are important (in this case it should be specified), this should be re-write it in a more general way. On the other side, it is very useful to have your UHPLC-SQ set up, calibration curve concentrations and so on.

The details were provided in response to the previous review comments in order to describe the action and show how critical steps are performed. It is indicated when a reader should follow the specific instrument manual.

The other mentioned details on calibration and so on are provided in the cited paper [27] and intentionally are not repeated in here.

2. In the protocol section, subheadings could be differentiated from the text a bit more clearly. For instance, subheading 1: "prepare the standard stock solutions of …." Could be replaced with "stock solutions of 3NT, Kyn, 3HKyn, 3HAA, XA" or something similar.

Appropriate changes have been introduced to the protocol section.

3. The overall text should be improved in terms of style, punctuation, and choice of words as understanding what the authors mean is sometimes difficult. JoVE seems to offer copy editing services. If so they should be used.

The manuscript has been edited by an experienced scholarly writer.

4. In the discussion section, certain considerations should be restricted to information what the data from this protocol can provide. The Kyn pathway includes many metabolites, some of them with direct and indirect effects on the levels of other, so the present analysis can only go so far in terms of providing a complete picture of the pathway.

We agree and did not aim to extensively draw conclusions from the results on determined Kyn levels. Rather we show possible directions in studying regulation of Kyn pathway by measuring changes in concentration of several metabolites (instead of Kyn only) that is possible using the presented protocol.

Minor - detailed:

1. Title: …culture medium FROM cancer cell… sounds like that the media was produced by the cancer cells.

Title has been changed and now it reads:

“Simultaneous Quantification of Selected Kynurenines analyzed by Liquid Chromatography-Mass Spectrometry in medium collected from Cancer Cell Cultures”

2. Summary, line 24: "four kynurenine pathway tryptophan metabolites" is not grammatically correct and is a confusing sentence.

The sentence now reads:

Described here is an accessible protocol for determination of four different tryptophan metabolites generated in kynurenine pathway….

3. Abstract, line 30: tryptophan derived metabolites (and in line 33, tryptophan derived products) are generally termed as tryptophan catabolites.

Corrected.

4. Abstract, line 35: "reached" is probably not the right word. Moreover, the medium is mostly analyzed because it reflects some of the metabolic events that occur at the cellular level.

The sentence was corrected accordingly.

5. Introduction, lines 59-60: description of the kynurenine pathway could be expanded and the reader would benefit from a schematic recap of it.

The description of kynurenine pathway was expanded according to the suggestion.

6. Introduction, line 62: XA is not mentioned in ref. 1

Additional references have been added.

7. Introduction, lines 68-69: sentence: "Some changes in Trp metabolism reveal in the tumor microenvironment and cancer cells" is not clear…

Corrected and it reads: “Some changes in Trp metabolism are observed in tumor microenvironment and cancer cells2, 8.”

8. Introduction, lines 69-73: there is the need to have good methods to simultaneously detect as many kyn metabolites as possible from different matrixes, with easy, flexible, and reliable protocols.

We changed this paragraph based on the suggestion and it now reads:

“Therefore, there is the need to have appropriate methods to simultaneously detect as many KP metabolites as possible in variety of biological specimens, with easy, flexible, and reliable protocol.”

9. Introduction, line 77: in my opinion, it is inappropriate to have ref. 7 there. Moreover, the whole sentence has to be rewritten. Lastly, the analysis done using commercially available kit could be very fast.

We added the original reference by Takikawa et al that can educate readers on the colorimetric assay (p-dimethylaminobenzaldehyde - Ehlirich reagent) for kynurenines quantification. It is non-specific and several compounds might interfere with Kyn therefore in modern analysis the chromatographic methods are preferred. The sentence has been also reworded for clarity.

c10. Introduction, lines 85-90: too much emphasis on the analysis of the cell culture medium. It is important, of course, but it is not an absolute novelty.

We edited this paragraph to express the intention to be more general for in vitro assays.

11. Introduction, lines 91-94: a robust protocol should work well for most currently available cell culture media. It is more relevant to have protocols that include as many kyn metabolites as possible.

Similar to comment #10.

12. Protocol, point 1.1: 0.3 mg is a very little amount of compound. Weighing this amount is risky since it could introduce a huge error in the analyte quantification.

We used in our laboratory a precise microbalance to weight reagents and analytical grade standards are usually expensive owing their instability in solutions. This is only an example how to prepare standards. Other users might use higher amounts to prepare standards keeping the given proportions. This suggestion has been added to the protocol.

13. Protocol, points 1.2s: can be very easily condensed and better phrased.

Corrected.

14. Protocol, point 1.2.4, NOTE: why three cycles?

It is recommended because of standards’ instability. The explanation has been added.

15. Protocol, point 2.5: does "cold" methanol mean ice cold or kept at -20C?

We assume the comment relates to step 3.2. Methanol was kept at -20C. This information was added.

16. Protocol, point 3.8: a cloudy sample could indicate that there is some protein left-over or that the compound is not properly solubilized.

Correct, we added this explanation.

17. Protocol, point 5.5: I would recommend to ensemble the guard column before purging and flushing the system.

Corrected.

18. Protocol, point 5.6.1: this gradient is very long considering the type of detection that you are using. Could it be improved? Alternatively, could it host other Kyn metabolites?

In this protocol we used Single Quadrupole (SQ) mass detector that is less selective compared to Triple Quadrupole, because it does not contain collision cell and is not able to fragment ions. Thus, when working with SQ, the detection and analytes identification is based on both retention time and monitoring of selected ions characteristic for the target compounds. Some sample components might generate ions with m/z selected for the analyte. In our opinion, it is better to apply lower gradient to obtain better peak separation and reduce errors associated with interfering matrix components. In such settings the protocol can be further expanded onto other compounds. It might require some adjusting in gradient program.

19. Protocol, point 6: bolt

Corrected

20. Protocol, point 6.2: redundant

The point 6.2 was removed according to the suggestion.

21. Protocol, point 6.4.2: if you do repeat charcoal purification, the medium should be cleared from your interfering analytes. Otherwise, repeat treatment one more time or do not apply any charcoal purification and consider the amount of these molecules before and after incubation with the cells.

The suggested clarification is included in point 2.5.

22. Protocol, point 6.4.3.1-2: explain it better.

Some explanation has been added.

23. Protocol, 6.4.4 NOTE: this is not correct, intraday and interday variability should be checked carefully.

The NOTE has been removed from the text.

24. Protocol, point 7.1 NOTE: if you want to compare the medium from different wells (containing the same cell type ± treatments) you should make sure that the same number of cells per well is seeded and the same number of cells is considered during the experiment. Perhaps, specify that protein content analysis could normalize the results.

We agree with Reviewer. That information was included as a NOTE in step 7. We also added a description how to estimate protein content (answering previous recommendations) in protocol step 9.

25. Protocol, point 7.5 NOTE: alternatively, you can add more points in your calibration curve (if solubility and linearity are not a problem) to adjust it to your sample concentration. Moreover, "diluted matrix" is a rather confusing definition, please, change it with a more appropriate word (this applies also in other parts of the text).

The phrase “Diluted matrix” has been changed to “100 times diluted culture medium” and the note expended with the recommended information.

26. Representative results, line 420: English, check it.

Corrected

27. Discussion: check typos (e.g. poropose, line 485, extremaly, line 525, etc), and proper terms (e.g. Kyn is the most profound Trp metabolite, what does it mean?)

Corrected.