# Submission ID #:  61128

# Scriptwriter Name: Anastasia Gomez

# Project Page Link: https://www.jove.com/account/file-uploader?src=18644088

# Title:   Preparation of Peripheral Nerve Stimulation Electrodes for Chronic Implantation in Rats

# Authors and Affiliations:

# Camilo A. Sanchez1, Jackson Brougher2, Kim C. Rahebi3, Catherine A. Thorn2

# 1Department of Bioengineering, University of Texas at Dallas, Richardson, TX, USA

# 2School of Behavioral and Brain Sciences, University of Texas at Dallas, Richardson, TX, USA

# 3Texas Biomedical Device Center, University of Texas at Dallas, Richardson, TX, USA

# Corresponding Authors:

# Catherine A. Thorn		(catherine.thorn@utdallas.edu)

# Email Addresses for All Authors:

camilo.sanchez@utdallas.edu

jjb180001@utdallas.edu

# kim.rahebi@gmail.com

catherine.thorn@utdallas.edu

## Author Questionnaire

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique?  Yes

If Yes, can you record movies/images using your own microscope camera?

No

If No, JoVE will need to record the microscope images using our scope kit (through a camera port or one of the oculars). Please list the make and model of your microscope.

https://www.amscope.com/stereo-microscopes/7x-45x-trinocular-articulating-zoom-microscope-ring-light.html

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  No

3. Filming location: Will the filming need to take place in multiple locations?   No

# Introduction

- Introductory Interview Statements

REQUIRED:

  - Catherine Thorn: Peripheral nerve stimulation is increasingly under investigation as a treatment for a variety of diseases, including inflammatory diseases and neural disorders.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Catherine Thorn: Construction of chronically implantable peripheral nerve cuff electrodes for small laboratory animals can be quite time-consuming and expensive. This protocol describes a simple, low-cost approach for constructing chronically implantable peripheral nerve stimulating electrodes for use in rats.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Camilo Sanchez: The method requires very few, easily accessible materials and supplies, and a reduced number of hazardous or technically demanding steps compared to similar techniques, which reduces both the cost of the device as well as the time required to train personnel in device assembly.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

OPTIONAL:

  - Kim Rahebi: This technique was designed to enable preclinical studies that test the effectiveness of chronic vagus nerve stimulation during stroke recovery or other types of motor dysfunction.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

Ethics Title Card

  - Procedures involving animal subjects have been approved by the Institutional Animal Care and Use Committee (IACUC) at the University of Texas at Dallas.

# Protocol

Videographer NOTE: Talent was struggling until 2.6, then filming got easier.

- Stimulating Cuff Electrode Fabrication

  - Begin by preparing the cuff tubing [1]. Use a razor blade to cut a 2.5-millimeter piece of polymer tubing [2]. Then, insert a paper clip or forceps tips through the tubing and make a lengthwise slit through the wall of the tubing on one side the cuff [3].

    - WIDE: Establishing shot of talent at the microscope.

    - SCOPE: Talent cutting the tubing.

    - SCOPE: Talent inserting paper clip into the tubing and making a slit.

  - Remove the forceps from the tubing and insert a large sewing needle through the midline of the cuff [1]. Place the needle into the foam board to pin the cuff in place during the remaining assembly steps [2].

    - SCOPE: Talent removing the forceps and inserting the large sewing needle. NOTE: 2.2.1 and 2.2.2. filmed together.

    - SCOPE: Talent pinning the needle to the foam board.

  - To place a suture for securing cuff closure during implantation, insert a small sewing needle through the wall of the cuff, on the midline, approximately 0.5 millimeters from the top slit on one side [1].

    - SCOPE: Talent inserting the small sewing needle.

  - Then, insert 2 centimeters of a 6-0 suture through the eye of the needle [1] and pull the needle through the wall of the tubing to thread the suture into the cuff [2].

    - SCOPE: Talent inserting suture through the eye of the needle.

    - SCOPE: Talent pulling the needle through the tubing.

  - Puncture a second hole through the tubing wall approximately 0.5 millimeters below the first hole, along the midline of the cuff [1]. Insert the suture through the eye of the needle and pull the needle through the tubing wall [2].

    - SCOPE: Talent puncturing a second hole through the tubing.

    - SCOPE: Talent inserting the suture through the eye of the needle and pulling it through the hole.

  - Apply a small amount of UV cure adhesive to the short end of the suture extending from the lower hole [1] and pull the longer suture end until the lower tail is nearly flush with the exterior wall of the tubing [2], then cure the adhesive with a UV wand [3]. Repeat this process to place another suture on the opposite side of the cuff [4].

    - SCOPE: Talent applying the adhesive.

    - SCOPE: Talent pulling on the longer suture.

    - Talent curing the adhesive.

    - SCOPE: Talent beginning to place another suture on the opposite side of the cuff.

  - Next, place the platinum-iridium wire leads. Use the small sewing needle to make 4 holes in the cuff wall [1], then insert the sewing needle again, this time working from exterior to interior, through the first lead hole [2].

    - SCOPE: Talent making holes in the cuff wall.

    - SCOPE: Talent inserting the sewing needle through hole 1.

  - Insert approximately 0.5 centimeters of a 7.5-centimeter platinum-iridium wire through the eye of the needle [1] and pull the needle through the tubing to thread the wire lead through the cuff wall [2]. Adjust the wire so that approximately 4.5 centimeters extends on the exterior side of the cuff [3].

    - SCOPE: Talent inserting wire through the eye of the needle. NOTE: 2.8.1 – 2.8.3 filmed together

    - SCOPE: Talent pulling the needle through the tubing.

    - SCOPE: Talent adjusting the wire.

  - Insert the needle through the first lead hole again, then through second lead hole [1]. Insert 0.5 centimeters of the shorter end of the wire through the eye of the needle [2] and pull the needle through the tubing to thread the wire lead through the cuff walls [3]. Repeat the process to thread the wire through the third and fourth lead holes [4].

    - SCOPE: Talent inserting the needle through hole 1 and then hole 2.

    - SCOPE: Talent inserting the wire through the eye of the needle.

    - SCOPE: Talent pulling the needle through.

  - Using a butane lighter, carefully remove the insulation from 5 to 6-millimeters of the wires extending from the second and fourth lead holes [1].

    - Talent removing the insulation from the wires.

  - Then, gently pull on the end of the wire extending from the first hole until the uninsulated portion of wire is flush with the hole [1]. Repeat this with the other lead to align the uninsulated end of the wire threaded through the third and fourth lead holes [2].

    - SCOPE: Talent pulling on the wire from hole 1.

    - SCOPE: Talent pulling on the other leads.

  - Apply a small amount of UV cure adhesive to the wire loops on the exterior side of the cuff at the first and third lead holes [1], then use the UV wand to cure the adhesive and secure the leads in place [2].

    - SCOPE: Talent applying the UV adhesive to the wire loops.

    - Talent curing the adhesive.

  - Use a small pipette tip to push the uninsulated wire leads against the interior wall of the cuff [1]. Fold the 1-millimeter tails of the wire flat against the exterior surface of the cuff, taking care not to short them together [2], then apply a small amount of UV adhesive to just cover the two tails and cure it [3].

    - SCOPE: Talent pushing the uninsulated wire leads against the interior of the cuff.

    - SCOPE: Talent folding the tails of the wire.

    - SCOPE: Talent applying the adhesive to the two tails and curing it.

  - Remove the large needle with the cuff assembly from the foam board [1]. Insert a 3-centimeter length of a 6-0 suture through the eye of the needle [2] and pull the needle through the tubing to thread the suture through the bottom of the cuff at the midpoint [3].

    - Talent removing the large needle. NOTE: 2.14.1 – 2.14.3 filmed together, author suggested that this portion should maybe be scaled up.

    - Talent inserting a suture through the eye of the needle.

    - Talent pulling the needle through the tubing.

  - Insert the small sewing needle through the same midline hole, working from interior to exterior to avoid deformation of the tubing and the wire leads [1]. Then, insert the exterior tail of the suture through the eye of the needle [2] and pull the needle through the cuff wall to create a loop of suture around the edge of the cuff [3].

    - Talent inserting the small needle.

    - Talent inserting the exterior tail of the suture through the eye of the needle. NOTE: 2.15.2 – 2.15.3 filmed together

    - Talent pulling the needle through the cuff wall.

  - Create a second loop around the opposite end of the cuff by tying the ends of the suture in a half knot on the exterior side of the cuff [1]. While holding the knot flat against the tubing, apply a small amount of UV adhesive to the half-knot and cure it [2].

    - SCOPE: Talent tying the ends of the suture in a knot. NOTE: Do not use take 1, use take 2

    - SCOPE: Talent applying adhesive to the knot and curing it.

  - Carefully cut the ends of the suture thread as close to the knot as possible [1]. Then, use a butane lighter to remove the insulation from approximately 3 millimeters at the end of each of the platinum-iridium wire leads [2]. Solder the cup side of a gold pin to the uninsulated end of each lead [3].

    - SCOPE: Talent cutting the ends of the suture thread.

    - Talent removing the insulation from the wire leads with the lighter.

    - Talent soldering the cup side of the gold pin to the uninsulated end of each lead.

  - Before proceeding with implantation, test the impedance of the assembled device to verify that the assembled cuff has an impedance of less than 2 kilowatts at 1 kilohertz [1].

    - Talent testing the impedance of the device.

- Device Usage

  - After assembling the headcap and implanting the cuff electrodes according to manuscript directions, stimulate the vagus nerve while the rat is awake [1].

    - Rat with the implanted device.

  - Connect the rat to a stimulus generator via the implanted head-cap and adjust the stimulation settings [1]. For VNS-induced reorganization of the motor cortical map, deliver a single train of 15 biphasic pulses, each with a width of 100 microseconds and an amplitude of 800 microampere, at a frequency of 30 Hertz [2].

    - Talent connecting the rat to the stimulus generator.

    - Talent adjusting stimulation settings.

  - A stimulation train is delivered immediately after the detection of each successful lever press throughout ten 30-minute training sessions [1]. Use an oscilloscope to monitor successful delivery of current stimulation [2].

    - Rat during stimulation.

    - Oscilloscope monitoring the current delivery.

# Results

- Results: Chronic Vagus Nerve Stimulation in Awake Behaving Rats

  - This protocol was used to chronically implant vagus nerve cuff electrodes and headcaps in rats. During initial implantation, there was a significant correlation between initial blood oxygen saturation readings and the percent change in blood oxygen saturation evoked by vagus nerve stimulation, or VNS [1].

    - LAB MEDIA: Figure 2 B.

  - To test the long-term functionality of the devices, VNS was applied 6 weeks after implantation [1]. For 7 out of 9 devices, the magnitude of stimulation-evoked change in blood oxygen saturation did not differ from that observed at initial implantation, suggesting continued excellent performance [2].

    - LAB MEDIA: Figure 2 C.

    - LAB MEDIA: Figure 2 C. Video Editor: Emphasize the 6-week data points outside the grey area.

  - In the remaining 2 devices, increasing the stimulation amplitude to 1.6 milli ampere was sufficient to evoke a reliable reduction in blood oxygen saturation of at least 5%, suggesting that these devices continued to function but that changes in impedance, nerve damage, or cuff orientation may have resulted in reduced performance [1].

    - LAB MEDIA: Figure 2 C. Video Editor: Emphasize the 6-week data points in the grey area.

  - To further test the long-term functionality of the chronically implanted devices, a second group of rats was trained on a simplified version of a skilled reaching lever-press task. The rats were required to reach 2 centimeters outside the training booth to fully depress a lever, and then to release it within 2 seconds [1].

    - LAB MEDIA: Figure 2 D.

  - Consistent with prior studies demonstrating that vagus nerve stimulation drives expansion of task-relevant motor map representations, VNS treated rats exhibited significantly larger proximal forelimb representations than sham treated rats [1].

    - LAB MEDIA: Figure 2 E. Video Editor: Emphasize the red bar.

# Conclusion

- Conclusion Interview Statements

  - Jackson Brougher: This protocol can be easily adapted to accommodate chronic stimulation experiments in other laboratory species or at other peripheral nerve sites, including sciatic or sacral nerves, for example.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.