# Submission ID #: 61161

# Scriptwriter Name: Bridget Colvin

Project Page Link: http://www.jove.com/files_upload.php?src=18654153

Title: Primary Clarification of CHO Harvested Cell Culture Fluid using an Acoustic Separator

Authors and Affiliations: Jin Sung Hong1, Nicole Azer1, Cyrus Agarabi1, and Erica J. Fratz-Berilla1

1U.S. Food and Drug Administration, Center for Drug Evaluation and Research, Office of Product Quality, Office of Biotechnology Products, Division of Biotechnology Review and Research II

☐   All author names are spelled correctly, and the affiliations are correct (city/state/country information not included in video title page).

# Corresponding Author:

# Erica J. Fratz-Berilla

# Erica.berilla@fda.hhs.gov

# Co-authors:

jinsung.hong@fda.hhs.gov

nicole.azer@fda.hhs.gov

# cyrus.agarabi@fda.hhs.gov

## Author Questionnaire

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique?  N

2. Software: Does the part of your protocol being filmed demonstrate software usage?  Y

If Yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps. Please upload all screen captured video files to your project page as soon as reasonably possible.

3. Filming location: Will the filming need to take place in multiple locations (greater than walking distance)?   N

- To ensure that your script can be filmed in one day, the Introduction, Protocol, and Conclusion sections are restricted to 55 shots (i.e., designated as 1.1.1., 1.2.1., etc).

Protocol Length

Number of Shots: 40

# Introduction

- Introductory Interview Statements

Your answers to these questions will become author interview statements, which will be incorporated into the script. Authors will memorize the statements and then deliver them on camera.

- Enter the full name of the author who will deliver the statement on camera.

- Each author should deliver no more than two statements.

- Fill out both required statements. One optional statement may also be selected.

- Please answer in full sentences in a style suitable for being spoken aloud.

- Limit the length of each statement to 30 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: Why is your protocol significant? OR What key questions can this method help answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

REQUIRED: What is the main advantage of this technique?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Do the implications of this technique extend toward the therapy (or diagnosis) of a particular disease, disability, or challenge? How so?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Are there any specific areas of research that this method could provide insight into? OR Can this method be applied to any other systems?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: How would you expect an individual who has never performed this technique to struggle? Do you have any advice to offer to somebody who is trying this technique for the first time?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Why is visual demonstration of this method critical?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Has any author been assigned more than two statements? ☐ No

# Introduction of Demonstrator on Camera

  - Enter name of author who will introduce demonstrator: Demonstrating the procedure will be Click here to enter name of demonstrator(s)., a Click here to enter demonstrator job title. from my laboratory. Include additional demonstrators as needed. [1][2].

    - INTERVIEW: Author saying the above

    - Named demonstrator(s) looks up from workbench or desk or microscope and acknowledges camera

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The one-digit numbers represent sections of the video. The text will appear onscreen.

- The two-digit numbers (e.g. 2.1., 2.2.) represent steps of your protocol. The text will be recorded by a professional voiceover talent.

- The three-digit numbers (e.g. 2.1.1., 2.2.2.) represent the shots that our videographer will capture at your lab.

Please use this draft script to help you prepare for filming day.

- Filming should take no more than 10 minutes per step. If a step will take more than 10 minutes, prepare the product from that step in advance.

- Acoustic Wave Separator (AWS) Preparation

  - To prepare the acoustic wave separator, connect the turbidity cables to their respective ports [1] and connect the chamber power BNC (B-N-C) cables to the back of the acoustic wave separator system [2-TXT].

    - WIDE: Talent connecting cable(s)

    - BNC cables being connected TEXT: Do not connect cable to acoustophoretic chamber until after fluid filling

  - Insert the turbidity probes into the turbidity meter [1] and the thermometer housing [2] and tighten the screws [3].

    - Probes being inserted into meter

    - Probes being inserted into housing

    - Screw(s) being tightened

  - Connect the feed tubing to the input of the feed turbidity port via the feed pump [1] and connect the y-tubing from the output of the feed turbidity port to the inlet ports of the acoustophoretic chamber [2].

    - Talent connecting tubing to turbidity port via feed pump

    - Talent connecting y-tubing to inlet port(s)

  - Connect the stage1 tubing from the waste port of the acoustophoretic chamber to a cell collection vessel via the stage1 pump [1] and connect the tubing from the permeate port of the acoustophoretic chamber to the input of probe1 turbidity port [2].

    - Talent connecting stage1 tubing to stage1 pump

    - Talent connecting permeate port tubing to probe1 turbidity port

  - Then connect the harvest tubing from out of the probe1 turbidity port to a product collection vessel [1].

    - Talent connecting harvest tubing to collection vessel

- System Priming

  - To prime the system with harvested cell culture fluid, turn on the acoustic wave separator [1] and open the associated software program [2].

    - WIDE: Talent turning on AWS

    - Talent at computer, opening program, with monitor visible in frame

  - In the Readings panel, click Start Test to initiate the data recording [1] and connect the feed tubing end into the harvested cell culture fluid vessel being stirred [2].

    - SCREEN: To be provided by Authors: Start Test being clicked

    - Talent connecting tubing to vessel

  - To start the feed pump, enter the pump rate and press Enter [1].

    - SCREEN: To be provided by Authors: Pump rate being entered, then Enter being clicked

  - Confirm that the Pump Direction Arrow is correctly oriented to ensure that the cell fluid is pumped from the vessel into the acoustophoretic chamber and click the Triangle icon to start [1].

    - SCREEN: To be provided by Authors: Arrow direction being checked, then triangle being clicked TEXT: 60 mL/min recommended pump flow rate

  - Monitor the feed turbidity measurements in the Percent Reduction Panel during the filling of the acoustophoretic chamber. If the harvested cell culture fluid is being mixed sufficiently within the cell fluid vessel, the turbidity values will be consistent during the loading of the chamber [1].

    - SCREEN: To be provided by Authors: Shot of Measurements Video Editor: please emphasize turbidity values when mentioned

  - Once the liquid is above the piezo transducer at the back of the acoustophoretic chamber [1], click Turn Off to stop the pump [2] and connect the other end of the BNC power cable to the acoustophoretic chamber [3].

    - Shot of liquid above transducer

    - SCREEN: To be provided by Authors: Turn Off being clicked

    - Talent connecting cable to chamber

- AWS Operation

  - Once the acoustophoretic chamber is filled and powered on, change the rate of feed pump to the desired operating rate [1].

    - WIDE: Talent change feed pump flow rate

  - To turn on the stage1 piezo power, slide the bar in the power module to 10 Watts and click Turn ON [1].

    - SCREEN: To be provided by Authors: Power being turned on, then bar being slid

  - Once the cells begin to settle to the bottom of the acoustophoretic chamber [1], start the stage1 pump with an appropriate rate based on the cell density and feed pump rate [2-TXT].

    - Shot of cells settled at bottom of chamber Videographer/Video Editor: can skip this step if no visual

    - Talent starting pump TEXT: Stage1 pump rate =Feed pump rate ×[2 ×Feed packed cell mass (%)]

  - To calculate packed cell mass, tare a scale with an empty 15-milliliter [1] and fill the tube with feed material [2].

    - Talent placing empty tube onto balance

    - Talent adding feed material to tube, with feed material visible in frame

  - After record the total weight of the tube with feed, centrifuge the tube [1-TXT] and decant the supernatant into a new container [2].

    - Talent placing tube into centrifuge TEXT: 10 min, 3700 x g, RT

    - Supernatant being decanted

  - Measure the weight of the tube with the cell pellet [1-TXT].

    - Talent placing tube onto balance TEXT: Packed cell mass feed material percentage = (decanted tube weight/filled tube weight) x 100%

  - Then monitor the turbidity profile of the stage1 turbidity as the overflow from the acoustophoretic chamber enters the turbidity probe1 [1].

    - LAB MEDIA: Figure 8

- Ending AWS

  - When the run is over, stop the feed and stage1 pumps [1] and turn off the power to the chamber [2].

    - WIDE: Talent turning off pumps, with monitor visible in frame

    - Talent turning off power

  - Disconnect the BNC power cable [1] and collect the product harvest material for downstream analysis [2-TXT].

    - Talent disconnecting cable(s)

    - Talent collecting material TEXT: Discard cell harvest material

  - To drain the remaining fluid from the acoustophoretic chamber, place the waste tubing into an empty vessel [1], disconnect the tubing from the acoustophoretic chamber inlet and permeate ports [2], and release the waste tubing from the pump head [3].

    - Talent placing tubing into vessel

    - Talent disconnecting tubing

    - Talent releasing waste tubing

  - When the chamber is empty, reconnect the tubing [1], place the waste tubing back into the pump head [2], and use a feed and stage1 pump rate of 60 milliliters/minute to flush deionized water from the end of feed tubing for 15-20 minutes [3-TXT].

    - Talent reconnecting tubing

    - Talent placing waste tubing into pump head

    - Water being flushed TEXT: Discard flow through

  - At the end of the water rinse, pump 70% isopropyl alcohol through the tubing and chamber for 15-20 minutes [1] followed by an additional 15-20-minute deionized water rinse [2].

    - Talent adding IPA to tube(s), with IPA container visible in frame

    - Use 5.4.3. water being flushed

## Protocol Script Questions

# Authors: Please use the step numbers from the script above (not step numbers from the manuscript) when answering the questions below. Please do not include steps that will be screen-captured and do not list entire sections.

A. Which steps from the protocol are the most important for viewers to see? Please list 4 to 6 individual steps.

Click here to list 4 to 6 individual steps, using the step numbers from the protocol section of the video script.

B. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1 or 2 individual steps from the script above.

Click here to list 1 or 2 individual steps, using the step numbers from the protocol section of the video script.

# Results

Please review this section to make sure that it accurately reflects your findings.

- Use Track Changes when making edits or revisions.

- If you would like the video to include different results, please revise this section.

- When revising, please keep the length of the voiceover below 200 words. Current word count: 159. (Voiceover is the text that follows the two-digit numbers.)

- Please note that the video cannot include voiceover without an accompanying visual.

- Results: Representative Turbidity Measurements and Cell Removal Efficiency Comparison

  - In this representative analysis, as the harvested cell fluid entered the acoustic wave separator chamber [1], the turbidity measurements from the feed turbidity probe remained consistent, around 1000-1100 NTU (N-T-U) [2-TXT], and measurements from the probe1 turbidity probe remained around 40-50 NTU [3].

    - LAB MEDIA: Figure 8

    - LAB MEDIA: Figure 8 Video Editor: please emphasize blue data line TEXT: NTU: nephelometric turbidity unit

    - LAB MEDIA: Figure 8 Video Editor: please emphasize orange data line

  - Keeping the harvested cell fluid within the chamber for longer time periods causes temperature increases [1], as in this example of a greater than 6-degree Celsius temperature difference of the harvested cell fluid before entering the acoustophoretic chamber [2] and after acoustic separation [3].

    - LAB MEDIA: Figure 9

    - LAB MEDIA: Figure 9 Video Editor: please emphasize blue data line

    - LAB MEDIA: Figure 9 Video Editor: please emphasize pink data line

  - Another important consideration when running high cell density harvests is the saturation of the turbidity probes [1], as the turbidity measurements for the feed turbidity probe become saturated over 4400 NTU [2], likely resulting in an underestimation in the calculation of the cell removal efficiency [3].

    - LAB MEDIA: Figure 10

    - LAB MEDIA: Figure 10 Video Editor: please emphasize data line about 4400 NTU

    - LAB MEDIA: Figure 10

  - In addition, the cell removal efficiency decreases significantly from approximately 100% to 57% as the feed pump rate is increased [1] with a slower the feed pump rate typically directly correlating with a better the cell clarification [2].

    - LAB MEDIA: Figure 11 Video Editor: please emphasize 1.5 and 3.5 ml/min data bars

    - LAB MEDIA: Figure 11 Video Editor: please emphasize 35 ml/min data bar

    - LAB MEDIA: Figure 1 Video Editor: please emphasize 1.5 and 3.5. ml data bars

# Conclusion

- Conclusion Interview Statements

Below are prompts for interview statements that can be used to further emphasize the significance of your protocol.

- Please answer one or two of the statements.

- Each statement is limited to 30 words.

- Answer the questions in full sentences; you will need to memorize and deliver the sentences as spoken interview statements during filming.

- Indicate the full name of the author who will deliver each statement.

# What is the most important thing to remember when attempting this procedure? Please indicate the steps (e.g., 2.4., 2.5.) in the Protocol section of the script that this advice applies to.

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera (Enter step numbers referred to.)

Following this procedure, what other methods can be performed? What questions would these additional methods answer?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

# After its development, did this technique pave the way for researchers to explore new questions within a specific scientific field? If so, how?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Thank you for addressing our questions. We will incorporate your answers and suggestions and send you the final script before your filming day. You will also receive detailed preparation instructions in the email accompanying the final script.