## Sheet1
Protocol step | Problem | Possible reason | Solution
1.1 | Software cannot be installed or updated | Running out of disk space on SD card | Check disk space on SD card. If required, remove unnecessary items, empty bin, or use an appropriately sized SD card
1.2 | API cannot be installed | Users ability for installation is restricted (no root user permission) | Use ‘sudo’ command in front of specified commands to gain admin rights. In Linux, this kind of access is known as the superuser.
3.1 | Troubles with GelMA | Functionalization, dialysis or lyophilization | Detailed step-by-step protocol including troubleshooting list available in Loessner et al.33.
5.1 and 6.2 | Workstation is not reacting to commands | Connection issues | Turn of everything and shut down the computer. Switch off the power supply for 10 s. Power computer and workstation back on.
5.1 and 6.2 | Workstation is not reacting to commands | Connection issues | Check if the computer is recognizing the USB connection and the USB port is defined correctly. Make sure that firewall is not preventing the connection process (see the link below Table 2).
5.1 and 6.2 | Cannot open file | Wrong directory | Check director (folder path) to ensure that the right path is being used. If a file (e.g., interface.py) cannot be found, it is likely that the wrong path is being used.
6.6.2 | Tip is not properly attached or falls during movement | Calibration issue | Repeat calibration steps for the pipette and make sure that the capillary piston is connected properly with the pipette.
6.6.2 | Tip is not properly attached or falls during movement | Attachment issue | Pipette is not properly connected to the pipette axis and moves during the movement steps. Tighten screws firmly to prevent this.
6.6.2 | Tip is aspirating above material | Calibration issue | Repeat calibration of this tray type to define height properly.
6.6.2 | Tip is aspirating above material | Calibration issue | Check volume in tube and make sure that the volume equals the volume defined in the protocol designer application.
6.6.2 | Material is dipping during movement | Too much excess material on tip | Add tip touch dock option; optionally, also the time for tip touch can be increased.
6.6.2 | Material is solid or too viscous for pipetting | Thermoresponsive behaviour of material | Check thermoresponsive material characterization and adjust heating/cooling temperature of temperature dock accordingly.
https://support.opentrons.com/en/articles/2687601-c-having-trouble-connecting-try-this-basic-troubleshooting