## Hoja1
Supplemental File 2. Main differences between the optimized and the non-optimized protocol for TAA, AA and DHAA extraction and quantification. The samples used were the same in both cases
Section | Step | Optimized protocol | Non-optimized protocol
Reagent and solution preparation | Extraction solution | 8% acetic acid (v/v), 1% MPA (w/v), 1 mM EDTA | 5% acetic acid (v/v), 3% MPA (w/v)
| Reducing solution | 40 mM DTT in 0.5 M Tris pH 9.0 | 30 mM DTT in dH2O
| Sulphuric acid (H2SO4) | 0.4 M | -
| Hydrochloric acid (HCl) | 2M | 5M
| Stock AA standard preparation | Solvent: Ultrapure water pH 2.0 acidified by 98-100% formic acid | Solvent: dH2O
| AA calibration curve | 5 points (0.5, 2.5, 5, 10, 25 µg mL-1) | 6 points (10, 20, 40, 60, 80, 100 µg mL-1)
|  | Solvent: Ultrapure water pH 2.0 acidified by 98-100% formic acid | Solvent: 1.5% MPA
Extraction of AA and DHA | Extraction solution volume (mL) | 5 | 3
| Mixing | Vortex 5 sec; orbital shaker (2000 rpm) 10 min at room temperature | Vortex 1 min
| Sonication (ultrasound bath) | 10 min at room temperature | No
| Centrifugation (4000 g) | 10 min at 4 ºC | 20 min at room temperature
| Filtration | 0.22-µm regenerated cellulose filter | 0.45-µm regenerated cellulose filter
DHAA reduction to AA | Incubation at room temperature | 30 min | 5 min
| Stop reaction | 200 mL of 0.4 M H2SO4 | No
Determination | Instrument | UPLC: Acquity H-Class | HPLC: Hewlett Packard 1050
| Detector | PDA eλ Detector λabs for AA=245 nm | eλ Detector λabs for AA=265 nm
| Column | Acquity UPLC HSS T3 (150 mm x 2.1 mm x 1.8 µm) | HPLC C18 Tracer column (250 mm x 4 mm x 5 µm)
| Channels | A: CH3OH; B/Wash: H2O:CH3OH (50:50 v:v); C: Ultrapure water pH 2.0 with formic acid; D/Seal Wash | No
| Mobile phase | 0.3 ml·min-1 of 2%A + 98%C (isocratic mode) | 1 mL min-1 of KH2PO4 30 mM ajusted pH 3.0 HCl 5M (isocratic mode)
| Column temperature | 35ºC | Room temperature
| Autosampler temperature | 5ºC | No (room temperature)
| Injection volume (µl) | 5 | 20
| AA retention time (min) | 1.874 | 2.98
| Total running time (min) | 3 | 7