# Submission ID #: 61468

# Scriptwriter Name: Bridget Colvin

Project Page Link: https://www.jove.com/account/file-uploader?src=18747788

Title: Evaluation of the Curing of Adhesive Systems by Rheological and Thermal Testing

Authors and Affiliations: Ana Díaz-Díaz1*, Bárbara Sánchez-Silva1*, Javier Tarrío-Saavedra1*, Jorge López-Beceiro1*, Silvia Gómez-Barreiro1*, and Ramón Artiaga1*

*These authors contributed equally

1Escola Politécnica Superior, Universidade da Coruña

# Corresponding Author:

# Jorge López-Beceiro

# jorge.lopez.beceiro@udc.es

# Co-authors:

ana.ddiaz@udc.es

barbara.silva@udc.es

javier.tarrio@udc.es

jorge.lopez.beceiro@udc.es

silviag.gomez@gmail.com

ramon.artiaga@udc.es

## Author Questionnaire

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique?  N

2. Software: Does the part of your protocol being filmed demonstrate software usage?  Y

Videographer: All video files provided, do not film screen capture steps

3. Filming location: Will the filming need to take place in multiple locations (greater than walking distance)?   Y, 48 km apart

Protocol Length

Number of Shots: 58

# Introduction

- Introductory Interview Statements

REQUIRED:

  - Ana Díaz: Our methodology combines thermal analysis and rheology to characterize the curing process of an adhesive and to obtain useful information for industrial adhesive selection [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

REQUIRED:

  - Javier Tarrío: This technique allows the creation of a standard guide for the curing process study of an adhesive system, making it easier to compare different adhesives [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL:

  - Bárbara Sánchez: This methodology can also be used as an acceptance criterion in the quality control of adhesive systems [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

# Protocol

- Manufacturer Curing Condition Checking: Thermogravimetric Cured Sample Testing

  - To perform thermogravimetric test to determine the inorganic filler content and the temperature at which the material starts to degrade, open the Procedure tab [1] and click Editor. Drag the Ramp segment to the Editor screen and establish the ramp as 10 or 20 degrees/minute to 900 degrees Celsius [2].

    - WIDE: Talent opening Procedure tab, with monitor visible in frame

    - SCREEN: screenshot_1: 00:00-00:13

  - Click OK and open the Notes tab. Select Air as the purge gas and set the flow rate to 100 milliliters per minute. Click Apply [1].

    - SCREEN: screenshot_1: 00:14-00:22

  - Then close the furnace [1] and start the experiment [2].

    - Talent closing the Furnace, with monitor visible in frame

    - SCREEN: screenshot_1: 00:23-00:30

- Manufacturer Curing Condition Checking: Differential Scanning Calorimetry (DSC) of a Cured Sample

  - To perform a differential scanning calorimetry test of a cured sample, open the Procedure tab [1], click Test, and select Custom [2].

    - WIDE: Talent opening Procedure tab, with monitor visible in frame

    - SCREEN: screenshot_2: 00:00-00:04

  - Click Editor and drag an Equilibrate segment indicating the temperature at which to start the experiment [1].

    - SCREEN: screenshot_2: 00:05-00:17

  - Drag a Ramp segment to the Editor screen and introduce a heating rate of 10 or 20 degrees/minute and the final temperature into the command editor window [1-TXT].

    - SCREEN: screenshot_2: 00:18-00:27 TEXT: Final temperature lower than degradation temperature

  - Drag a Ramp segment to the Editor screen and introduce a 10 or 20 degrees/minute cooling rate to a temperature tentatively below the glass transition [1].

    - SCREEN: screenshot_2: 00:28-00:39

  - Drag another Ramp segment to the Editor screen and introduce a 10 or 20 degrees Celsius/minute heating rate to a temperature slightly below the degradation temperature [1].

    - SCREEN: screenshot_2: 00:40-00:47

  - Open the Notes tab and select Nitrogen as the flow gas. Set the flow rate to 50 milliliters/minute and click Apply [1].

    - SCREEN: screenshot_2: 00:48-00:57

  - Then place a reference pan [1] and a pan with a sample inside the DSC cell [2] and launch the experiment [3].

    - Talent placing reference pan into cell

    - Talent placing sample pan into cell

    - SCREEN: screenshot_2: 01:48-01:52

- DSC Fresh Sample Analysis: Ramp Curing Test

  - To analyze the fresh sample through a heating-cooling-heating test, open the Procedure tab [1] and click Test and Custom [2].

    - WIDE: Talent opening Procedure tab, with monitor visible in frame

    - SCREEN: screenshot_3: 00:00-00:04

  - Click Editor and drag an Equilibrate at minus 80 degrees Celsius segment to the Editor screen [1].

    - SCREEN: screenshot_3: 00:05-00:13

  - Drag a Ramp segment and set the heating rate to 10 or 20 degrees Celsius/minute to a temperature slightly below the degradation temperature and insert another Equilibrate at minus 80 degrees Celsius segment [1].

    - SCREEN: screenshot_3: 00:14-00:26

  - Then drag a Ramp segment and set the heating rate to 10 or 20 degrees Celsius/minute to the same temperature as before. Click Ok [1].

    - SCREEN: screenshot_3: 00:27-00:37

  - Then place a reference pan [1] and a pan with a sample inside the DSC cell [2] and click Start to launch the experiment [3].

    - Talent placing reference pan into cell

    - Talent placing sample pan into cell

    - SCREEN: screenshot_3: 01:23-01:30

- DSC Fresh Sample Analysis: Isothermal Curing Test

  - To perform an isothermal curing test, open the Procedure tab [1], click Test, and select Custom [2].

    - WIDE: Talent opening tab, with monitor visible in frame

    - SCREEN: screenshot_3: 00:00-00:04

  - Click Editor and drag a Ramp segment to the Editor screen. Introduce a 20 degrees Celsius/minute to the chosen isothermal temperature [1].

    - SCREEN: screenshot_4: 00:02-00:18 Video Editor: can speed up

  - Then introduce an Isothermal segment with enough time to complete the cure [1].

    - SCREEN: screenshot_4: 00:19-00:23

  - To check the degree of curing reached, introduce an Equilibrate at 0 degrees Celsius segment, add a Ramp segment, and set the heating rate between 2 and 20 degrees Celsius/minute to the maximum temperature [1].

    - SCREEN: screenshot_4: 00:24-00:35

  - Drag the Mark end of cycle segment to the editor window and insert another Equilibrate segment with a temperature of minus 80 degrees Celsius [1].

    - SCREEN: screenshot_4: 00:36-00:48

  - To obtain the final glass transition, add a Ramp segment with a heating rate between 2 and 20 degrees Celsius/minute to the same temperature as indicated before and click OK [1].

    - SCREEN: screenshot_4: 00:49-00:55

  - In the Tool tab, select Instrument Preferences and DSC and set a temperature lower than the isotherm temperature of the experiment [1].

    - SCREEN: screenshot_4: 00:56-01:04

  - Click Apply and open the Control tab to select Go to Standby Temperature [1].

    - SCREEN: screenshot_4: 01:05-01:10

  - Then place a reference pan [1] and a pan with a sample inside the DSC cell [2] and click Start [3].

    - Talent placing reference pan

    - Talent placing sample pan

    - SCREEN: screenshot_4: 02:38-02:41

- Logarithmic Strain Sweep Test

  - To perform a logarithmic strain sweep test, open the Procedure tab [1] and select Oscillation Amplitude [2].

    - WIDE: Talent opening Procedure tab, with monitor visible in frame

    - SCREEN: screenshot_5: 00:00-00:09

  - Set the experimental temperature to room temperature, the frequency to 1 hertz, and the logarithmic sweep from 1 x 10-3 to 100% of strain [1].

    - SCREEN: screenshot_5: 00:10-00:30 Video Editor: please speed up

  - Place the sample on the bottom plate with the upper plate separated about 40 millimeters from the lower plate [1] and lower the upper plate until a gap of about 2 millimeters is observed between both plates [2].

    - Talent placing sample onto bottom plate

    - Upper plate being lowered

  - Then trim the excess adhesive [1] and start the experiment [2].

    - Adhesive being trimmed

    - SCREEN: screenshot_5: 00:31-00:33

- Isothermal Multifrequency Curing Test

  - To monitor the curing of the adhesive, click the Procedure tab [1] and select Conditioning Options [2].

    - WIDE: Talent clicking Procedure tab, with monitor visible in frame

    - SCREEN: screenshot_6: 00:00-00:12

  - Set the Mode to Compression, the Axial Force to 0 Newtons, and the Sensitivity to 0.1 Newtons [1].

    - SCREEN: screenshot_6: 00:13-00:19

  - Click Advance and set the Gap change limit to 2000 microns in both the up and down directions [1].

    - SCREEN: screenshot_6: 00:20-00:36 Video Editor: please speed up

  - Insert a new oscillatory time sweep step and set the experimental temperature to room temperature, the test duration as a function of the estimated curing time based on the Data Sheet of the adhesive, and the strain percentage acquired from the previous logarithmic strain sweep test [1].

    - SCREEN: screenshot_6: 00:44-01:16 Video Editor: please speed up

  - Select Discrete and set the frequencies 1, 3 and 10 hertz for all of the samples [1].

    - SCREEN: screenshot_6: 01:17-01:38 Video Editor: please speed up

  - Then load a new sample [1] and start the experiment [2].

    - Sample pan being placed

    - SCREEN: screenshot_6: 01:41-01:48

- Torque sweep and Temperature Scan Testing

  - To perform a torque sweep test, open the Procedure tab [1] and select Oscillation Amplitude [2].

    - WIDE: Talent opening Procedure tab, with monitor visible in frame

    - SCREEN: screenshot_7: 00:00-00:05

  - Then set the temperature to room temperature, the frequency to 1 hertz, and the logarithmic sweep from 10 to 10000 micronewton-meters of torque and start the experiment [1].

    - SCREEN: screenshot_7: 00:06-00:36 Video Editor: please speed up

  - From the torque sweep test, choose a torque amplitude within the linear viscoelastic region to use in the Temperature Ramp test [1]. Then, select Temperature Ramp and set up the experiment at room temperature with a ramp rate of 1 degree Celsius/minute to ensure a uniform distribution of the temperature into the sample, a frequency of 1 hertz, and the torque amplitude determined from the torque sweep test [2]. NOTE to VO: Long one, feel free to break up into 1 and 2.

    - Choosing the torque amplitude within the linear viscoelastic region

    - SCREEN: screenshot_8: 00:00-00:45

  - Close the furnace of the rheometer [1] and open the air stopcock of the furnace [2].

    - Talent closing the furnace

    - Talent opening the air stopcock

  - Then start the experiment [1].

    - SCREEN: screenshot_8: 00:52-00:54

## Protocol Script Questions

A. Which steps from the protocol are the most important for viewers to see?

n/a

B. What is the single most difficult aspect of this procedure and what do you do to ensure success?

n/a

# Results

- Results: Representative Cured Adhesive System Rheological and Thermal Testing

  - These thermogravimetric results show different degradation temperatures [1] and different inorganic fillers for each studied adhesive [2]. The mass loss observed between 600-800 degrees Celsius suggests the presence of calcium carbonate as filler [3].

    - LAB MEDIA: Figure 1A Video Editor: please emphasize data lines between 185 and 240 °C (x-axis) and 110-80% (y-axis)

    - LAB MEDIA: Figure 1A Video Editor: please emphasize data lines at 900 °C

    - LAB MEDIA: Figure 1A Video Editor: please emphasize blue and dotted green lines between 600-800 °C.

  - For this two-component adhesive, in the heat flow curve, there is no evidence of residual cure [1-TXT] and the small deviation [2] cannot be assigned with certainty to the glass transition reported by the manufacturer [3].

    - LAB MEDIA: Figure 2A TEXT: tetrahydrofurfuryl methacrylate + benzoyl peroxide

    - LAB MEDIA: Figure 2A Video Editor: please emphasize green data line at 60 °C

    - LAB MEDIA: Figure 2A

  - In this Table, the degree of curing [1] of a two-component system at different temperatures [2] was calculated by comparing the curing enthalpy acquired at each temperature to that obtained in a heating ramp [3].

    - LAB MEDIA: Table 1 Video Editor: please emphasize Degree of curing column

    - LAB MEDIA: Table 1 Video Editor: please emphasize Temperature column

    - LAB MEDIA: Table 1 Video Editor: please emphasize Curing enthalpy column

  - Through a rheological multifrequency test of a fresh two-component adhesive sample [1], the gelation time of the material can be observed as the point at which the phase angle becomes frequency independent [2].

    - LAB MEDIA: Figure 5

    - LAB MEDIA: Figure 5 Video Editor: please emphasize green data lines

  - In these isothermal multifrequency tests using the one- and two-component silane polymer adhesive systems [1], no sign of gelation is observed [2] and a comparison of the slopes of the moduli of both adhesives reveals that the two-component silane polymer adhesive cures faster [3].

    - LAB MEDIA: Figures 6 and 7

    - LAB MEDIA: Figures 6 and 7 Video Editor: please emphasize green data lines in both graphs

    - LAB MEDIA: Figures 6 and 7 Video Editor: please emphasize blue data lines in Figure 7

  - In this rheological temperature scan test of a two-component adhesive sample cured for one hour [1], the glass transition can be clearly observed [2].

    - LAB MEDIA: Figure 8

    - LAB MEDIA: Figure 8 Video Editor: please emphasize gray data line at 60.2 °C

# Conclusion

- Conclusion Interview Statements

  - Silvia Gómez: Do not delay starting the test when using a fresh sample and be sure to prepare a thoroughly mixed solution when using a two-component system [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera