- 61491_screenshot_1_t2.mp4

4. Electrode Test

4.3.1 To select the DC-ERG protocol to be performed, click Protocols and select

DC-ERG [1]. 0:00:17 – 0:00:26

4.4.1 Click Run. A dialog box will pop-up with the patient information. If the information is correct, click Yes and proceed to Step 1 of 6 [1]. 0:00:26-0:00:33

4.6.1 Verify that the values for the mouth reference, tail ground, and recording electrodes are acceptable and click Step to proceed to step 4 of 6 to test the baseline stability [1]. 0:00:37-0:00:55

4.7.1 To begin viewing the traces, click Preview. The traces should be low noise with a peak-to-peak amplitude of less than 200 microvolts [1-TXT].

0:00:56-0:03:50

- 61491_screenshot_2_t2.mp4

5. Mouse and Electrode Positioning

5.2.2  . . . and, in the ERG system software, verify that the correct patient is selected [2]. 0:00:02 – 0:00:08

5.3.1 Click Protocols. Under Protocol Description, select DC-ERG and click Run. Click Yes to verify that the correct test is being performed [1]. 0:00:09 – 0:00:15

6.1.2 To record the DC-ERGs, click Step to select step 5 of 6 [1] and click Impedance to use the Impedance Checking screen to examine the resistances of the left and right eyes [2]. 0:02:46 – 0:02:51

6.2.1 The impedance values for the recording electrodes at each eye should be similar and the impedance values for both the ground and reference electrodes should be less than 10 kiloohms [1]. 0:03:09 – 0:03:23

6.3.1 Click Preview to view the traces for the left and right eyes and wait up to 10 minutes for a stable baseline to be achieved [1]. 0:08:52 – 0:09:57

6.4.1 Then click Stop to exit the trace preview and click Run to start the recording [1]. 0:11:00 – 0:11:02 (Stop) and 0:11:11 – 0:13:25 (Run)