Dear Editorial staff and reviewers.

The authors are grateful to the editorial office and the reviewers for affirmative and constructive treatment of our manuscript. Most of editing requirements and reviewers’ suggestion were incorporated in the revised version.

Please find below a point by point response to the critique.

Editorial comments:
Changes to be made by the Author(s):
1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues.

We proof-read the revised manuscript twice.

2. Please format the manuscript as: paragraph Indentation: 0 for both left and right and special: none, Line spacings: single. Please include a single line space between each step, substep and note in the protocol section. Please use Calibri 12 points

Done.

3. Please define all abbreviations during the first-time use.

All abbreviation were defined as requested.

4. Please rephrase the Short Abstract/Summary to clearly describe the protocol and its applications in complete sentences between 10-50 words: “Presented here is a protocol …”

The short abstract was re-written.

5. Please ensure that the long Abstract is within 150-300-word limit and clearly states the goal of the protocol.

Our abstract is 240 words long and includes the statement of the protocol goal.

6. For in-text formatting, corresponding reference numbers should appear as numbered superscripts after the appropriate statement(s) without brackets.

Done.

7. JoVE cannot publish manuscripts containing commercial language. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials and Reagents.
For example: TAMRA, Chandler Loop, etc.

We completely eliminated commercial language. However, both TAMRA (tetramethylrhodamine) which now is spelled out at the first occurrence in the text, and Chandler loop apparatus are not commercial names, and thus were left in the text of the protocol.

8. Please use generic term throughout.

Done.

9. Please adjust the numbering of the Protocol to follow the JoVE Instructions for Authors. For example, 1 should be followed by 1.1 and then 1.1.1 and 1.1.2 if necessary. Please refrain from using bullets, alphabets, or dashes.

Done as requested.

10. Please ensure that all text in the protocol section is written in the imperative tense as if telling someone how to do the technique (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “Note.”

The text in the protocol section is written in the imperative mood.

11. The Protocol should contain only action items that direct the reader to do something.
12. Please ensure you answer the “how” question, i.e., how is the step performed?
13. Please ensure that individual steps of the protocol should only contain 2-3 actions sentences per step.

Everything (11-13) was corrected accordingly.

14. There is a 10-page limit for the Protocol, but there is a 2.75-page limit for filmable content. Please highlight 2.75 pages or less of the Protocol (including headings and spacing) that identifies the essential steps of the protocol for the video, i.e., the steps that should be visualized to tell the most cohesive story of the Protocol.

Done.

15. Please obtain explicit copyright permission to reuse any figures from a previous publication. Explicit permission can be expressed in the form of a letter from the editor or a link to the editorial policy that allows re-prints. Please upload this information as a .doc or .docx file to your Editorial Manager account. The Figure must be cited appropriately in the Figure Legend, i.e. “This figure has been modified from [citation].”

All figures are new. No permission from previous publications is required.

16. As we are a methods journal, please ensure that the Discussion explicitly cover the following in detail in 3-6 paragraphs with citations:
a) Critical steps within the protocol
b) Any modifications and troubleshooting of the technique
c) Any limitations of the technique
d) The significance with respect to existing methods
e) Any future applications of the technique

Our revised discussion covers all of the suggested topics.

17. Please do not abbreviate the journal titles in the references section.

The official JoVE style in the Endnote use abbreviated journal titles. Nevertheless we changed them for the unabbreviated titles after Endnote formatting.

18. Please sort the materials table in alphabetical order.
Done.

Reviewer 1:

1. In this work, there are many steps of append therapeutic peptide moieties to a stainless steel surface…. could you add some markers to make them visible?
Since the quantification of the immobilized peptide is based on the use of fluorescently tagged CD47, we used fluorescence microscopy to visualize the appended peptide. A new figure (Figure 3) was added to the revised protocol to demonstrate the difference in the surface fluorescence between bare and PABT/PEI-PDT-modified metal surface treated with TAMRA-conjugated pepCD47.

2. … should contain blood compatibility evaluation, such as the proliferation and migration of blood cells on coatings.

3. There is no enough result show the new therapeutic peptide moieties in a stainless steel surface has anti-inflammatory and anti-thrombotic ability, only Figure 3 can't prove the purpose of your work.

In the revised protocol, the authors present a new data (Fig. 5) showing the specific inhibition of macrophage expansion on the pepCD47-modified compared to the bare metal surfaces. Significant reduction in the number of macrophages associated with pepCD47-functionalized surfaces 6 days after the monocyte seeding, reflects both the decreased adhesion of monocytes and diminished macrophage proliferation and supports the anti-inflammatory effects of pepCD47 coating.

Reviewer 3:

1. The first paragraph is not up to date with the current status of stent bioengineering.

The text was revised according to the reviewer’s suggestion.

2. Throughout the paper, the authors should acknowledge that the modified Chandler loops to assess materials can be used to evaluate the acute thrombogenicity and acute inflammatory cell response to materials.

The text was modified to reflect the limitation of the short-term studies performed using the Chandler loop. Additionally, a new Fig. 5 presents the data pertaining to the long-term monocyte culture on the bare metal and pepCD47-modified metal surfaces.

3) What was the blood anticoagulated with? Were the donors screened for recent medication e.g. aspirin?

Blood was anticoagulated with sodium citrate. The donors were free of anti-platelet drugs. The text was modified to convey this information.  This information is now included in the revised manuscript.

4) Chandler loops are usually just rotated, so it would be more accurate to say 'Blood was rotated for 4 hours…

The text was corrected as suggested by the reviewer.

5) Error in x axis of Figure 2. Highest dose is 200ug/ml in the paper, but the figure axis shows 300ug/ml.

This error was corrected.