Troubleshooting

This section describes how to troubleshoot several common errors that arise when performing microinjection.

- Error: The software reports an error message when loading in step 1.1.

  - Causes:

    - The manipulator is not plugged into the computer via Ethernet.

    - The manipulator is not turned on.

    - Two instances of the application are launched (two windows are open).

  - Solution: Verify the manipulator is plugged into the computer and turn on the manipulator. Close the programs and wait two minutes to verify programs are shut. Relaunch the application.

- Error: The video stream is not displayed in the software in step 1.1 and an error message says “CAMERA ERROR”.

  - Causes:

    - The camera is not turned on.

    - The camera is not properly connected.

    - Two instances of the application are launched (two windows are open).

  - Solution: Verify the camera is turned on and properly connected to the computer. Close all instances of the application.

- Error: The pipette goes to first location but freezes in place when starting the microinjection in step 4.6.

  - Causes: There is a connection problem with the microcontroller.

  - Solution: Verify the microcontroller is plugged into the computer via USB. If the problem persists, plug the USB into a different USB port on the computer.

- Error: The pipette is clogged.

  - Causes:

    - Molecules in solution aggregated and clogged the tip.

    - The pipette is not the ideal shape.

  - Solution: Shake the injection solution and reload a different pipette. Pull the pipette at two degrees lower.