# Submission ID #: 61653

# Scriptwriter Name: Bridget Colvin

Project Page Link: https://www.jove.com/account/file-uploader?src=18804213

Title: Versatile Dual-Inlet Sample Introduction System for Multi-Mode Single Particle Inductively Coupled Plasma Mass Spectrometry Analysis and Validation

Authors and Affiliations: Daniel Rosenkranz1,*, Fabian L. Kriegel1, Emmanouil Mavrakis2, Spiros A. Pergantis2, Philipp Reichardt1, Jutta Tentschert1, Norbert Jakubowski4, Peter Laux1, Ulrich Panne3, and Andreas Luch1

1German Federal Institute for Risk Assessment (BfR), Department of Chemical and Product Safety, Max-Dohrn-Strasse

2Environmental Chemical Processes Laboratory, Department of Chemistry, University of Crete, Voutes Campus

3Federal Institute for Materials Research and Testing (BAM), Richard-Willstätter-Strasse

4SPETEC GmbH, Am Kletthamer Feld

☐   All author names are spelled correctly, and the affiliations are correct (city/state/country information not included in video title page).

# Corresponding Author:

# Fabian Kriegel

# fabian.kriegel@bfr.bund.de

# Co-Authors:

daniel.rosenkranz@bfr.bund.de

mavrmanos@gmail.com

spergantis@uoc.gr

philipp.reichardt@bfr.bund.de

jutta.tentschert@bfr.bund.de

norbert.jakubowski@spetec.de

peter.laux@bfr.bund.de

ulrich.panne@bam.de

andreas.luch@bfr.bund.de

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or similar?  N

2. Software: Does the part of your protocol being filmed demonstrate software usage?  Y

If Yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps. Please upload all screen captured video files to your project page as soon as reasonably possible.

3. Interview statements: Considering the Covid-19-imposed mask-wearing and social distancing recommendations, which interview statement filming option is the most appropriate for your group? Please select one.

☐ 	Interviewees wear masks until the videographer steps away (≥6 ft/2 m) and begins filming. The interviewee then removes the mask for line delivery only. When the shot is acquired, the interviewee puts the mask back on. Statements can be filmed outside if weather permits.

☐ 	Interviewees self-record interview statements outside of the filming date. JoVE can provide support for this option.

☐ 	Interview Statements are read by JoVE’s voiceover talent.

☐ 	Author interview statement opt out (interview statements removed completely from paper).

4. Filming location: Will the filming need to take place in multiple locations (greater than walking distance)?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

- To ensure that your script can be filmed in one day, the Introduction, Protocol, and Conclusion sections are restricted to 55 shots (i.e., designated as 1.1.1., 1.2.1., etc).

Protocol Length

Number of Shots: 39

# Introduction

- Introductory Interview Statements

Your answers to these questions will become author interview statements, which will be incorporated into the script. Authors will memorize the statements and then deliver them on camera.

- Enter the full name of the author who will deliver the statement on camera.

- Each author should deliver no more than two statements.

- Fill out both required statements. One optional statement may also be selected.

- Please answer in full sentences in a style suitable for being spoken aloud.

- Limit the length of each statement to 30 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: Why is your protocol significant? OR What key questions can this method help answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

REQUIRED: What is the main advantage of this technique?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Do the implications of this technique extend toward the therapy (or diagnosis) of a particular disease, disability, or challenge? How so?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Are there any specific areas of research that this method could provide insight into? OR Can this method be applied to any other systems?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: How would you expect an individual who has never performed this technique to struggle? Do you have any advice to offer to somebody who is trying this technique for the first time?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Why is visual demonstration of this method critical?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Has any author been assigned more than two statements? ☐ No

# Introduction of Demonstrator on Camera

Complete this statement ONLY if any of the individuals who will be demonstrating the procedure on camera will not be delivering an Introductory Interview Statement.

  - Enter name of author who will introduce demonstrator: Demonstrating the procedure will be Click here to enter name of demonstrator(s)., a Click here to enter demonstrator job title. from my laboratory. Include additional demonstrators as needed.[1][2].

    - INTERVIEW: Author saying the above

    - The named demonstrator(s) looks up from workbench or desk or microscope and acknowledges the camera

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The one-digit numbers represent sections of the video. The text will appear onscreen.

- The two-digit numbers (e.g. 2.1., 2.2.) represent steps of your protocol. The text will be recorded by a professional voiceover talent.

- The three-digit numbers (e.g. 2.1.1., 2.2.2.) represent the shots that our videographer will capture at your lab.

Please use this draft script to help you prepare for filming day.

- Filming should take no more than 10 minutes per step. If a step will take more than 10 minutes, prepare the product from that step in advance.

- Dual-Inlet Sample Introduction Setup Assembly

  - To set up the conventional sample introduction system, combine an ICP-MS spray chamber with a pneumatic nebulizer [1] and use a clamp to connect the spray chamber outlet to the male ball joint of the T-piece connector [2].

    - WIDE: Talent inserting PN into chamber

    - Outlet being connected to ball joint

  - To set up the microdroplet transport unit, use clamps to attach a demountable quartz torch with the injector tube removed to a laboratory stand with the torch inlet [1].

    - Torch being clamped to stand

  - Use closed-end gas connectors to block the torch plasma and auxiliary gas inlets [1] and use an appropriate gas connector to connect a helium gas line to the torch via its cooling gas inlet [2].

    - Talent placing connector onto inlet(s)

    - Talent connecting gas line to torch

  - Then use an adapter to connect a 30-centimeter-long conductive and flexible silicone tube to the exit end of the torch [1] and stretch the flexible silicone tubing over the remaining vertical metal connection to connect the downward end of the silicone tubing to the T-piece connector unit [2].

    - Tube being connected to torch

    - Tubing being stretched over connection/tubing being connected to unit

  - To connect the microdroplet generation and microdroplet generation control units, insert the microdroplet generator head into the sample inlet end of the torch [1] and connect the power supply to the microdroplet generator control unit [2].

    - Talent inserting head into sample inlet

    - Talent connecting power supply

- Multi-Mode Nanoparticle Sample Measurement

  - To prepare the microdroplet generator control unit for multi-mode nanoparticle sample measurement, turn on the power supply of the microdroplet generator control unit [1] and click Start to start the control unit [2].

    - WIDE: Talent turning on power supply

    - Talent clicking start, with monitor visible in frame

  - Then click Global Settings and click the right graphical button to select the triple pulse mode [1].

    - SCREEN: To be provided by Authors: Global settings being clicked, then triple pulse mode being selected

  - To prepare the microdroplet generator unit, click On-Off to start the microdroplet generator [1] and fill the sample vessel with the sample solution to be measured [2].

    - Talent clicking on/off

    - Talent filling vessel

  - Connect the sample vessel to the microdroplet generator unit [1] and connect a 10-milliliter syringe to the syringe port on the sample container vessel [2].

    - Talent connecting vessel to unit

    - Talent connecting syringe

  - Depress the syringe plunger until a constant liquid stream is observed coming out the microdroplet generator head [1].

    - Plunger being depressed/stream leaving head

  - After 10 seconds, remove the syringe [1] and place the microdroplet generator unit into the focus zone of a CCD camera connected to a PC or laptop [2].

    - Talent removing syringe

    - Talent placing unit in camera focus area

  - To observe droplet formation, initiate the CCD camera software [1] and click Start to obtain a live view of the droplets [2].

    - Talent starting camera software, with camera screen visible in frame

    - Start being clicked

  - Place the microdroplet generator head onto the inverted torch on the dual-inlet sample introduction system [1] and measure repeated multi-point calibrations to validate both the microdroplet generator and pneumatic nebulizer for each element of interest [2].

    - Talent placing head into torch

    - Calibration being measured Authors: Is this measurement performed on a computer screen?

  - When all of the elements have been calibrated, select a concentration within the linear range of the calibration curves for downstream one-point calibrations [1] and select a nanoparticle and an ionic standard according to the analyte of interest [2].

    - Talent at computer/lab notebook, selecting concentration

    - Talent selecting nanoparticle and ionic standard

  - Load the selected ionic standard solution into the sample vessel as demonstrated [1] and add an appropriate diluted acid solution via the pneumatic nebulizer [2].

    - Talent adding solution to vessel, with standard container visible in frame

    - Talent adding acid to PN, with acid container visible in frame

  - Start the measurement of the ICP-MS system in time resolved mode [1].

    - Talent starting measurement, with monitor visible in frame

  - After 120 seconds, click On-Off to stop the microdroplet generator [1] and replace the acid solution with the ionic standard [2].

    - Talent clicking On/Off

    - Talent adding standard to PN, with standard container visible in frame

  - After 330 seconds, replace the standard solution with fresh dilute acid solution [1] and remove the microdroplet generator unit from the setup [2].

    - Talent adding acid to PN, with acid container visible in frame

    - Talent removing unit

  - Exchange the sample vessel of the microdroplet generator unit with a vessel containing a diluted acid solution [1] and purge the unit with 10 milliliters of air as demonstrated [2-TXT].

    - Talent placing vessel of acid into generator

    - Plunger being depress/liquid leaving unit head TEXT: Maintain air pressure 30 s

  - Then add the nanoparticle sample solution to the sample vessel [1] and analyze the nanoparticle droplets as demonstrated for 510 seconds [2].

    - Talent adding solution to vessel, with solution container visible in frame

    - Talent clicking Start

  - At the 810 seconds, click On-Off and replace the dilute acid solution at the pneumatic nebulizer with the nanoparticle sample [1] and measure the sample for 300 seconds [2].

    - Talent adding sample to PN, with sample container visible in frame

    - Talent clicking Start

  - After 1200 seconds, clean the microdroplet generator unit as demonstrated [1].

    - Talent adding acid to sample vessel, with acid container visible in frame

## Protocol Script Questions

# Authors: Please use the step numbers from the script above (not step numbers from the manuscript) when answering the questions below. Please do not include steps that will be screen-captured and do not list entire sections.

A. Which steps from the protocol are the most important for viewers to see? Please list 4 to 6 individual steps.

Click here to list 4 to 6 individual steps, using the step numbers from the protocol section of the video script.

B. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1 or 2 individual steps from the script above.

Click here to list 1 or 2 individual steps, using the step numbers from the protocol section of the video script.

# Results

Please review this section to make sure that it accurately reflects your findings.

- Use Track Changes when making edits or revisions.

- If you would like the video to include different results, please revise this section.

- When revising, please keep the length of the voiceover below 200 words. Current word count: 170. (Voiceover is the text that follows the two-digit numbers.)

- Please note that the video cannot include voiceover without an accompanying visual.

- Notes in blue italics are for JoVE’s video editor.

- Results: Representative Droplet Size Analysis and Validation

  - The microdroplet generator droplet formation, including the droplet size [1], is characterized before initiating the experiment [2].

    - LAB MEDIA: Figure 3 Video Editor: please emphasize droplet measurement in Figure 3B

    - LAB MEDIA: Table 3 Video Editor: please emphasize NP size data column

  - After the setup assembly and droplet size determination, both injection systems are validated with ionic standards [1].

    - LAB MEDIA: Figure 4

  - An accuracy of r-squared greater than 0.99 can be achieved with both injection systems for all of the investigated elements [1].

    - LAB MEDIA: Figure 4 Video Editor: please emphasize r-squared values in both graphs

  - For these representative silver samples, the first linear range can be observed between 0 and 0.5 femtograms/event [1] and the second between 0.5 and 0.8 femtograms/event [2].

    - LAB MEDIA: Figure 4 Video Editor: please emphasize blue and yellow data lines between 0 and 0.5

    - LAB MEDIA: Figure 4 Video Editor: please emphasize blue and yellow data lines from 0.5 to end of data lines

  - In contrast, the first linear range for cerium is between 0 and 0.25 femtograms/event [1] and the second between 0.25 and 3 femtograms/event [2].

    - LAB MEDIA: Figure 4 Video Editor: please emphasize pink data lines between 0 and 0.25

    - LAB MEDIA: Figure 4 Video Editor: please emphasize pink data lines from 0.25 to end of data lines

  - The linear range for the pneumatic nebulizer for the measured concentrations appears to be higher, likely related to the difference of introduced mass into the ICP-MS per detection event [1].

    - LAB MEDIA: Figure 4 Video Editor: please emphasize PN graph

  - The microdroplet generator injects a constant absolute quantity in a low volume per drop and detection event, resulting in a lower detected mass compared to the introduction of samples with the pneumatic nebulizer [1].

    - LAB MEDIA: Figure 4 Video Editor: please emphasize MDG graph

  - After a successful validation, the particle size and number concentration of cerium dioxide nanoparticles can be determined [1].

    - LAB MEDIA: Figure 5 Video Editor: please emphasize data lines or no animation

# Conclusion

- Conclusion Interview Statements

Below are prompts for interview statements that can be used to further emphasize the significance of your protocol.

- Please answer one or two of the statements.

- Each statement is limited to 30 words.

- Answer the questions in full sentences; you will need to memorize and deliver the sentences as spoken interview statements during filming.

- Indicate the full name of the author who will deliver the statement.

# What is the most important thing to remember when attempting this procedure? Please indicate the steps (e.g., 2.4., 2.5.) in the Protocol section of the script that this advice applies to.

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera (Enter step numbers referred to.)

# Following this procedure, what other methods can be performed? What questions would these additional methods answer?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

# After its development, did this technique pave the way for researchers to explore new questions within a specific scientific field? If so, how?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Thank you for addressing our questions. We will incorporate your answers and suggestions and send you the final script before your filming day. You will also receive detailed preparation instructions in the email accompanying the final script.