Reviewers' comments:
Reviewer #1: 
Manuscript Summary:
The authors developed a dual-inlet setup for characterizing nanoparticles (NPs) by single particle inductively coupled plasma mass spectrometry (spICP-MS) without a reference NP standard. It consists of a pneumatic nebulizer (PN) for NP solutions and a microdroplet generator (µDG) for ionic calibration solutions. As a key part of the setup, they developed a new and flexible interface to facilitate the coupling of PN, µDG, and the ICP-MS system. They found that by using the setup, three independent analysis modes are available for determining the particle size and number concentration of NPs. The developed dual-inlet setup and its application will be of interest toresearchers using spICP-MS techniques for analysis of NPs. However, the authors have not consistently reached the same results among the three modes and evaluated the robustness of the suggested setup. Please consider the following comments to improve the clarity and comprehensiveness of the manuscript.

Major Concerns:
Title
It is unclear where the terms "improved" and "robust" in the title come from and which data given in the text can support them. This reviewer believes that the terms should be corrected to be matched with the content of the text.

We thank the reviewer for his comment about the title of the submitted manuscript. Due to reviewers comment the title was changed into the following:

“Versatile dual-inlet sample introduction system for multi-mode single particle inductively coupled plasma mass spectrometry analysis and validation”

Page 2, line 105
In conjunction with the above concern, how did the authors evaluate the robustness of the suggested setup? If they did not evaluate it, they cannot use the term "robust" here as well as in the title.

The title was changed according to the reviewer’s comment, and we hope that our following explanation provides clarity. ‘’Robust’’ is used to describe the ICP-MS stability and uninterrupted operational state, even during the time that the µDG is removed and therefore the inlet remains open. Removing the µDG does not cause plasma extinction, and the recorded ICP-MS signal is rapidly restored when the µDG is brought back to its operational position on the system. This “robust” system allows for rapid sample exchange as well as convenient µDG rinsing.

Page 11, Eq. (6)
Why can the two factors "ηPN" and "ηµDG" be cancelled during the transformation of the equation? It is not always true that their values are the same. This point should be clarified.

We thank the reviewer for his careful observation. An explanation of equation 6 was added at the end of section “6. Data Analysis” and we hope that it enhances the readability of the submitted manuscript:

The transport efficiency calculated by equation 6 assumes that ηµDG equals 1, as 100% of the analyte reaches the plasma through the µDG (as mentioned in ref 17, but also we observe the correct number of particles according to operating frequency of the µDG). This in turn allows for determining the transport efficiency of the pneumatic nebulizer, with the only prerequisite of introducing a dissolved calibration standard through both inlets.

Page 13, Figure 4
Page 15, lines 415 and 416
In the case of the µDG, the linearity of the calibration curves for 107Ag and 140Ce was worse than that for 197Au. Regarding this trend, the authors mentioned later as follows.
"However, the measured concentrations introduced by the μDG have to be separated into two linear ranges."
If so, how can they explain that 197Au only exhibited the linear calibration curve with no separation? It seems to this reviewer that the µDG has unnoticed problems. In this context, the calculation of r2 values for three elements should be confirmed again whether correct or not, because the calibration curve for 197Au (with r2 of 0.9997) looks more linear than that for 140Ce (with r2 of 0.9999). These two points should be clarified.

For 197Au, ionic concentrations below 5 ng mL-1 were not tested, and this is probably the reason why the degree of linearity seems greater. We assume that the picture of the Au calibration curve for the low-end concentration range would be similar to the case of Ag and Ce, elements which were indeed tested for lower analyte concentrations. As mentioned in the paper (page 15 line 413 -415 of the original manuscript) the solution volume introduced by the µDG is by far lower compared to the PN (detecting fg of analyte per droplet compared to 1000s of fg in PN) and so an analyte concentration at the low ng mL-1 range (Ag range) cannot be distinguished clearly from a blank signal. Due to these points a direct comparison of the two introduction systems at the low ng mL-1 range is not possible (shown in the supplementary information of the original paper: https://doi.org/10.1016/j.aca.2019.11.043).

In order to address the remark of the reviewer and to clarify these points the following sentence was added into the discussion of the revised manuscript:

Below the overlapping region the observed signals are close to the element specific background. Above these limit the linear working range of the µDG can be identified.

Page 14, Table 6
Why were the recoveries of Au 56 nm and Ag 75 nm in the three analysis modes were so different each other? In the case of Au 56 nm, as a typical example, the difference in the recoveries obtained in Mode-I and Mode-III reached 30 %. In the both cases of Au 56 nm and Ag 75 nm, the difference in their recoveries obtained in Mode-I and Mode-II reached 5-9 %, although the transport efficiencies of the NPs were almost the same (i.e. ca. 2 %). Moreover, in the Excel file provided as an electronic supplementary material, the difference in the recoveries of Au 56 nm obtained in Mode-I and Mode-II amazingly reached 20 %. In addition to the above, why were the transport efficiencies of CeO2 JRC obtained in Mode-I and Mode-II different by 3.2 %? This difference is significantly larger than the cases of Au 56 nm and Ag 75 nm. These two points should be clarified.

We thank the reviewer for this interesting question about the recoveries with the different modes of analysis. We would like to state that small variations (c(NP) in transport efficiency (η) lead to a large variation in the particle number concentration. Attached below is an example, showing the effect of varying η% to the c(NP) determination.

The reason for which mode III exhibits a larger deviation in the determined c(NP) compared to modes I and II could be attributed to the higher number concentration of particles injected through the µDG, leading to particle aggregation and therefore accounting for the lower recoveries.

CeO2 JRC is a challenging material for spICP-MS analysis mainly due to its polydispersity and uncertainty in particle geometry. Granted that Mode I determines η by counting the number of particles detected, a possible particle aggregation could lead to a certain bias in the single particle events detected; meaning that signals coming from multiple particles could be counted as a single particle event.

For example: The following calculations are done by using equation 12 with qp = 100 NP detected, Vinjected = 0.25 mL, and the expected particle number concentration is 22,500 NP mL-1. Please keep in mind that η is used as relative value in equation 12.

| η (%) | cp,detected | Bias (%) |
| --- | --- | --- |
| 1.5 | 26667 | 18.52 |
| 1.7 | 23529 | 4.57 |
| 1.8 | 22222 | -1.24 |
| 1.9 | 21053 | -6.43 |
| 2 | 20000 | -11.11 |
| 2.2 | 18182 | -19.19 |

We hope this example answers the raised question sufficiently.

Minor Concerns:
Page 1, lines 43-46
Regarding the sentences below, which mode does "the third mode" indicate among the three modes given in the Page 11?
"While two analysis modes are known from the literature the third mode, introduced by this study, determines the transport efficiency using the data of inorganic ionic standard solutions. It is independent of reference materials."
It is unclear to this reviewer.

We thank the reviewer for spotting this misleading sentence. In order to clarify the named modes, the following sentences of Page 1 line 43 -46 was changed as follows:

While mode I (counting) and mode III (µDG) are known in the literature, mode II (sensitivity), introduced by this study, is used to determine the transport efficiency by simply introducing ionic calibration standard via each mode and taking the sensitivity ratio obtained for each mode.

Page 1, lines 61 and 62
The sentence below is understandable to this review.
"The transport efficiency describes the ratio of the mass or particle number injected to the intensity values measured by the ICP-MS."
It should be corrected; otherwise additional explanation should be provided for that.

The mentioned sentence was corrected in order to enhance the readability of the revised manuscript:

The transport efficiency describes the ratio of the mass or particle number injected to the mass (waste collecting method) or particle number (counting method) detected by the ICP-MS.

Pace, H. E.;  Rogers, N. J.;  Jarolimek, C.;  Coleman, V. A.;  Higgins, C. P.; Ranville, J. F., Determining Transport Efficiency for the Purpose of Counting and Sizing Nanoparticles via Single Particle Inductively Coupled Plasma Mass Spectrometry. Analytical Chemistry 2011, 83 (24), 9361-9369.

Page 1, line 67
There are some sources of the differences other than the "elemental composition" such as the structure and the dispersant as mentioned. Thus, this part should be corrected.

The reviewer correctly states other sources for the observed differences, we changed the mentioned section and thank the reviewer for his careful observation:

However, the transport properties depend on the structure of the NP, like the composition and the sample dispersant. Other influencing factors are instrumental parameters, like sample uptake rate, nebulizer gas flow rate, dwell time and total measurement time.

Page 1, lines 70-73
Appropriate references should be cited for the sentence below to clearly show the superiority of a µDG over a PN in terms of the probability of multi-particle events.
"It has been demonstrated that a continuous sample flow of PN increases the probability of multi-particle events, whereas injection of discrete sample volumes by a μDG (pL in a given time interval) lowers the probability of such events."
This reviewer believes that the probability of multi-particle events totally depends on the particle number concentration in the measured sample and therefore, a PN can work at the same level as µDG depending on the particle number concentration in the measured sample.

Unfortunately, the others of the manuscript were not able to find an appropriate reference that shows the superiority of a µDG over a PN in terms of the probability of multi-particle events.

The probability which was addressed within the sentence was related to the low volume injected by the µDG and is not based on Poisson statistics. In order to prevent confusion of the reader we decided to remove the mentioned sentence from the manuscript.

Page 1, line 75
What is the reason for "low injection volumes" being preferable? If there is a large amount of sample, small injection volume can lead to biased result.

We thank the reviewer for his question. The low injection volume is not only utilized to analyze higher particle number concentrations as for conventional injection systems. These low volume systems are also able, to detect particles which have a close proximity to each other, and to differentiate whether these particles are individual particles or have formed an agglomerate.

The sentence addressed by the reviewer was changed into following:

To be independent of reference materials, ideally, a sample introduction system with a transport efficiency of almost 100% is preferable. At the same time when a low volume is used compared to conventional introduction systems, there can be better distinguished between particles close to each other and higher particle number concentrations used.

Page 2, lines 85-89
Appropriate references should be cited for the sentence below to clearly show the superiority of a µDG over a PN in terms of the analytesensitivity, etc.
"Due to the uniform sample introduction at high transport efficiency of 100% of the μDG, high instrument specific analyte sensitivity can be achieved. Depending on the matrix used, this leads to lower limits of detection (LOD) of particle mass and size when compared to the results of conventional introduction systems based on PN."

The missing reference was added:

Mehrabi, K., Günther, D., Gundlach-Graham, A., Single-particle ICP-TOFMS with online microdroplet calibration for the simultaneous quantification of diverse nanoparticles in complex matrices. Environ. Sci.: Nano 2019, 6.

https://doi.org/10.1039/C9EN00620F

Page 3, Table 1
Regarding Part 2, why do the authors suggest these types of ICP-MS spray chambers and PNs? The reason for that should be mentioned somewhere because it could be informative.

We thank the reviewer for raising this issue. In fact, we have only experience for the available spray chambers, which are used in our lab. We also found, no other combinations with µDG in the literature. Irrespective of this other spray chambers can be attached to the shown setup.

In order to address the question a note was added to Table 1 Part 2:

Instead of the impact bead spray chamber other spray chambers with transport efficiencies typically in range of 2 to 10 % or higher can be used.

Page 7, line 236
What is the procedure below based on?
"Vortex or shake each solution for no more than 10 s."
In the nanoComposix protocol, for example, a longer time than 10 s is acceptable as follows.
"During storage, the nanoparticles may settle to the bottom of the vial (especially nanoparticles > 30 nm in diameter). Prior to aliquoting or use, resuspend thesettled nanoparticles by vigorously shaking the bottle until the solution is homogenous. This will typically require ~30 seconds of mixing. ..."
Thus, appropriate references should be cited here to show the validity of the procedure.

We thank the reviewer for raising this point. The sample preparation was done in according to the manufacturer specifications. For AuNP NIST 8013 30 seconds of vortexing and for AgNP NIST 8017 30 seconds of shaking was used. The mentions 10 s was a typing error and corrected:

Vortex (NIST 8013) or shake (NIST 8017) each solution for 20 - 60s.

The investigation report for NIST 8017 was cited according to the reviewers suggestion: Small, A. J., Watters, R. L., National Institute of Standards and Technology. Report of Investigation Reference Material 8017; 2015.

Page 7, line 238
What the value "approx. 2.56 mg mL−1" indicate? It is unclear to this reviewer.

The concentration 2.56 mg mL-1 of nano particles in BSA based solution was derived from the NanoGenoTox protocol. Slightly deviation from this value are acceptable to achieve a stable particle dispersion.

Page 8, Table 4
Why did the authors set the sampling depth (i.e. sampling position from the load coil) to 4 mm? This is far away from the normal sampling depth in the 8-10 mm range. The reason for that should be mentioned somewhere.

The sampling depth was optimized to reach the highest possible sensitivity for gold with the developed dual-inlet setup. The sampling depth was listed as the parameter used with our instrument iCAP Q from Thermo Scientific. For other instruments or analytes, the sampling depth can be different. To clarify, that the sampling depth can be optimized point 4.5 (in corrected manuscript) on page 7 was changed to:

Optimize instrumental parameters to improve analyte sensitivity if necessary, e.g. nebulizer gas flow rate, sampling depth, plasma power.

Page 9, lines 279 and 287
It seems strange to this reviewer that the sentence "Start the µDG" appears twice in one sequence of the µDG preparation without stopping it once. For clarifying that, the sentence "Stop the µDG" should be inserted between them.

To overcome the addressed problem, point 5.4.4 was changes as follows: Start the measurement of the µDG.

When the µDG is prepared (5.1), it is started in the 1st step of µDG preparation (5.1.1).

Page 10, line 341
Is the term"larger" correct? It would be "smaller."

The reviewer’s comment is correct, smaller signals indicated as background are removed from entire data set. As the reviewer suggested “larger” was exchanged by “smaller”.

Page 11, line 353 and Eq. (11)
The authors only defined (or assumed) that the factor "ηµDG" is 1; in other words, they did not calculate it. Thus, the previous phrase (given below) looks strange to this reviewer.
"Calculate the specific transport efficiency of the analysis modes:"
It should be changed to be correct.

In order to address this question we would like to highlight, that the transport efficiency of the µDG is 1 because a particle enters the plasma and is detected by the ICP-MS according to the operation frequency of the µDG, which means there are no missing particles. If one assumes that the µDG device is set to 10 Hz and 10 signals per seconds are detected via ICP-MS than naturally the ηµDG is 1.

The transport efficiency for the µDG is not calculated, it is only assumed to be 1 (relative) based on literature. Therefore, the calculation part for the µDG in the revised manuscript was changed as follows:

Assume that the transport efficiency of Mode III is equal to 1:17

Page 11, Table 5
The two factors "ni,p" and "td" should also be defined here in this table.

We thank the reviewer for spotting the missing factors. We added the missing factors into table 5.

Page 16, lines 467 and 468
The sentence below is not true based on the Table 6.
"In addition, the NPs introduced into the ICP-MS with the μDG have a narrower particle size distribution."
In the case of CeO2 JRC, for example, the standard deviation of the mean for the NP size obtained by Mode-III using the PN is larger than that obtained by Mode-I using the µDG. Thus, the sentence should be corrected.

To clarify the difference of the observed size distributions a more detailed explanation was added:

In addition, the NPs introduced into the ICP-MS with the µDG have a narrower (AuNP) or similar (AgNP) particle size distribution. In contrast for CeO2 a broader size distribution for the µDG was observed and can be attributed to the higher polydispersity of the analyzed sample.

Page 16, legend for Table 6
Since the reference #15 cited in the legend is the same as the reference #13, it should be cited correctly as the reference #13.

The doubling of reference 13 (now 14) was changed. 

Reviewer #2: 
In "Improved dual-inlet system for robust multi-mode single particle ICP-MS validation," the authors present a strategy for calibration of nanoparticle number concentration and nanoparticle mass distribution using only microdroplet based element standard solutions and nebulized element standard solutions. In my opinion, the biggest contribution of the authors with this manuscript is the modification of the dual-inlet setup by the use of conductive silicon tubing and modification of the ICP torch to accommodate a microdroplet dispenser head. In addition, the current manuscript provides a means for other practitioners of ICP-MS to build their own dual-inlet system; this is also a valuable contribution. However, before publication of this video manuscript, I think that several points should be addressed, which I've listed below:

* In the abstract, and at several points in the manuscript, the authors claim that a "third analysis mode" is used to determine transport efficiency based on sensitivities obtained from the same element introduced in both droplets and with the pneumatic nebulizer. However, this analysis mode is actually "Mode II" in the manuscript. Calling this analysis mode the "third mode" and "Mode II" is confusing and should be fixed. Also, the claim that this Mode II transport efficiency calibration strategy is introduced in this manuscript is false. While, it is true that the authors published this work in Rosenkranz, D. et. al. Anal. Chim. Acta 2020, 1099, 16-25, a variation of this calibration strategy was actually already published before the manuscript by Rosenkranz et. al. under the name "online microdroplet calibration":Mehrabi, K.; Günther, D.; Gundlach-Graham, A., Environ. Sci.: Nano 2019, 6 (11), 3349-3358. K. Mehrabi also presented this calibration strategy in a poster at the European Winter Plasma Conference in 2019 in Pau, France. Appropriate citations should be added.

The mentioned part into the abstract was fixed:

While mode I (counting) and mode III (µDG) are known from the literature mode II (sensitivity), is used to determine the transport efficiency by inorganic ionic standard solutions only

The second citation was used in line 105 page 1.

* The authors claim that "injection of discrete sample volumes by a µDG … lowers the probability of [muliti-particle events]." This is, in principle, not true. It could be true that higher NP concentrations can be run with a µDG and still just get one particle per droplet; however, if a correct NP concentration is run via the pneumatic nebulizer, one also has low chances of coincident particle events. Moreover, with a pneumatic nebulizer, the sample is introduced continuously, so, if microsecond time resolved ICPMS analysis is made, one could even separate multi-NP events based on the shape of the measured time-resolved signals.

We thank the reviewer for his explanation and we are agreeing. To be correct, the addressed sentence was deleted.

* In many microdroplet sample instruction setups, the microdroplet is viewed (and sized) online instead of prior to the placement of the µDG in its holder. In my experience, moving the µDG can lead to droplet instability, either in the form of a change in dispensed droplet size or in loss of the droplet. The authors should comment on challenges associated with removing and replacing the µDG so often in their workflow.

We thank the reviewer for raising this point. The µDG head is removed from the setup and after the sample solutions of the µDG are changed the µDG is placed carefully back to the setup while already running. Every time, when the µDG head was installed again into the setup signals were observed resulting from droplets. The variation into droplet size are not taking into account until now, because of the strong signal fluctuation depending on the wet-plasma conditions used.

* In the section on LODs with microdroplets there is a mistake. The LOD for a microdroplet analysis of NPs must take into account the intensity and standard deviation of elemental signals from droplets that have no NPs in them, not the background in which droplet signals are not present. This should be corrected.

In order to correct this issue a more detailed description was added as note in line 364 page 11

To determine the background intensity in point 6.3 of the µDG the droplet signals which contains no particles should be used.

* In figure 4, for the µDG calibration curve, the curve for silver is shallower than that for gold, but a higher sensitivity is reported for Ag107 than Au197. This is sloppy and inappropriate for publication. Also, in these plots, counts are plotted versus fg, but then the sensitivity is reported as cps/fg. Again, this makes no sense. Care should be taken in re-constructing these figures. Also, I do not know why droplet signals wouldn't produce entirely linear calibration curves. In the past, other authors have shown good linearity for microdroplet signals. Perhaps contamination is the reason for non-linear calibration? This should be better explained.

Figure 4 was checked ,redrawn and reanalyzed for mass per event injected according to the reviewers suggestion.

5

14201-
* Citation 13 and 15 are identical. Please correct.

The citations were corrected.

Reviewer #3: 
Manuscript Summary:
This manuscript provides a detail procedure for nanoparticle measurement using a dual-inlet system. The authors presents most of the key steps to operate this correctly. It can be accepted after some minor revision.

Major Concerns:
1. uDG，Is it neccessary to be turned on the He gas of the head ？ please emphasize this.

There was no possibility to use He gas on the µDG head directly. We know such a head is commercially available. The He was applied via the ICP-MS torch. By removing the µDG head from the setup the He gas reduce the entrance of oxygen into the setup. By removing the µDG head the setup is now opened and He gas can also leave through this opening against the flow direction of air.

An explanation why the He gas has to be used can be found in the revised manuscript at the protocol section (1. Assembly of the Dual-inlet sample introduction setup, 3rd point of Notes)

2. Figure 4, please give some explanation of why the signal response of Ag and Ce is not linear at the low concentration?

Because the lower concentrations are too close to the background, i.e. they cannot be fully separated from the background. As a result, ions getting lost and the observed signal intensity is reduced. The linear working range has to be defined when the used calibration point can be separated from the background.

In the manuscript the following sentence was improved:

If different concentrations for calibration along this limit are used two linear regions can be observed with an overlap at approximately 0.5 ng mL-1 for Ce and 2 ng mL-1 for Ag. Below the overlapping region the observed signals are close to the element specific background. Above these limit the linear working range of the µDG can be identified.

3. Figure 6, I suggest to provide some icons that mark the steps according to Figure 2

Figure 6 was changed to select the steps in according to figure 2

Figure 5: Representing measurement for the quantification of CeO2 NP with the dual-inlet system with colored bars in according to figure 2 for the different injection steps.