Dear editor,

Thank you for allowing us to submit the manuscript by December 11, 2020. Please find below our response to the reviewers. We have revised the manuscript according to their suggestions. We kept track of our revisions.

Looking forward to hearing from you.

Best wishes,

Hassina Bilheux on behalf of the co-authors

____________________________________


Please note that novelty is not a requirement for publication and reviewer comments questioning the novelty of the article can be disregarded.


Editorial comments:

Dear editor,

Thank you for your comments and guidance. See our response below.

Changes to be made by the Author(s):
1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues.

Done.

2. Please adjust the numbering of the Protocol to follow the JoVE Instructions for Authors. For example, 1 should be followed by 1.1 and then 1.1.1 and 1.1.2 if necessary. Please refrain from using bullets or dashes.

Done.

3. Please move the statement about ethics committee approval from 1.2 to before the numbered protocol steps, indicating that the protocol follows the guidelines of your institution’s animal research ethics committee. Did you get the approval for implantation into the rat femurs?

The statement about the ethics committee approval has been moved to Section 2. Data acquisition.

4. Please ensure that all text in the protocol section is written in the imperative tense as if telling someone how to do the technique (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “Note.” However, notes should be concise and used sparingly. Please include all safety procedures and use of hoods, etc.

Done.

5. The Protocol should contain only action items that direct the reader to do something. Please move the discussion about the protocol to the Discussion.

Done.

6. Please note that your protocol will be used to generate the script for the video and must contain everything that you would like shown in the video. Please add more details to your protocol steps. Please ensure you answer the “how” question, i.e., how is the step performed? Alternatively, add references to published material specifying how to perform the protocol action. Please add more specific details (e.g., button clicks for software actions, numerical values for settings, etc) to your protocol steps. There should be enough detail in each step to supplement the actions seen in the video so that viewers can easily replicate the protocol.

Done.

7. Please revise the protocol to break it into the main sections (using only the imperative tense for the steps): neutron source and beamline setup, specimen preparation, data acquisition for each specimen, volume reconstruction and data processing. Please include enough details for each procedure so that the steps can be filmed for the entire process for each specimen. After revising, please highlight up to three pages of text for filming.

Done.

8. Step 1 of the instrument setup: what are the pinhole aperture size and opening of the slit system?

The pinhole aperture is what defined the beam geometry and it is mentioned in the manuscript. Since this is the protocol, we don’t explain what each component is again, per your recommendation.

9. Step 2: How are the XY stages installed?

Revised to explain how to do this.

10. Step 3a: How and what lens is selected for sCMOS and how do you focus the camera?

Thank you for your comment. We revised the protocol accordingly.

11. Step 3b: What is involved in using a resolution mask to tune the focus?

Explained.

12. Step 4: How do you secure the sample and use neutron shielding?

The samples are positioned at the beamline in containers, often Al cylinders. Neutron shielding is placed against the detector. We have now clarified this in the manuscript.

13. Step 5: How is pixel size evaluated with the resolution mask?

Pixel size is evaluated by measuring a known feature dimension on the radiograph. Then, by knowing the number of pixels across a dimension, one can determine the size of the pixel. We revised the protocol accordingly.

14. Step 7: How do you align the sample with the neutron beam using the EPICS interface?

Thanks for making this comment. We revised to explain that this is done by taking a series of fast radiographs until the sample is in full view of the detector.

15. 1.1: How much saline-soaked gauze is used to freeze the implants? How is this done?

We use Curity Gauze Sponges, 2 in x 2 in (CardinalHealth, part number 2146), fully submerge them in phosphate buffered saline (PBS) and wrap each sample fully. The manuscript has been revised accordingly.

16. What do you mean by “it was position spending up on the rotation stage…”?

Our apologies, we meant that the sample was positioned sitting up or vertically. We corrected the manuscript, thank you for catching this error.

17. What do you mean by “water was periodically added”—please specify the frequency.

We have expanded the description of plant propagation and watering to indicate that during propagation water is applied daily based on mass loss of the plant system, which allows for precise control of the hydration within the soil.

18. Please rewrite the EPICS data acquisition protocol to specify the various parameters to be entered and selected, which will facilitate filming.

Done.

19. Please remove the embedded figure(s) from the manuscript. All figures should be uploaded separately to your Editorial Manager account. Each figure must be accompanied by a title and a description after the Representative Results of the manuscript text. Please include only Figure Legends after the Representative Results.

Done.

20. Please reorder the figure legends so that fig 1’s legend appears first.

Done.

21. Please sort the Materials Table alphabetically by the name of the material.
Done.

____________________________________
Reviewers' comments:
Reviewer #1:

Dear reviewer,

Thank you for your comments. We value your input and have done our best to respond to your review comments and suggestions. Please see below our response below in green.

Manuscript Summary:
The manuscript details how neutron CT on biological samples is performed on the CG-1D beamline at the High Flux Isotope Reactor, Oak Ridge National Laboratory. Measurements can be tailored to different sample types. In particular, dynamics (on the time scale of hours) can be studied, at the cost of signal to noise ratio. Three representative samples are showcased: 1) a rat femur, 2) a mouse lung, and 3) dynamic water uptake in a root system. Overall, the methods are detailed, and the manuscript is well-written.

Minor suggestions:
Line 96 - The microscopic cross section appearing in Eqn. 2 should be called the "total collision cross section" - (diffuse plus absorption cross sections), as opposed to the "total cross section", which could be confused with the "total scattering cross section" (coherent plus incoherent cross sections). For a thorough treatment, see Chapter 4 of Sears, Neutron Optics.

Thank you for this comment. We believe that calling the total neutron cross section can lead to confusion. Thus, we decided to use neutron attenuation cross section, as enunciated in the ASTM E1316-17a standard terminology for nondestructive examinations.

Line 118 - many authors have discussed the complementarity of neutron and x-ray CT. I would suggest referencing a review such as Strobl, et al. 2009. Journal of Physics D: Applied Physics, 42(24), 243001. Additionally, the complementarity of neutrons and x-rays is very-well showcased by the simultaneous Neutron and X-ray Tomography (NeXT) system, at the National Institute of Standards and Technology, as described in LaManna et al. 2017. Review of Scientific Instruments 88(11), 113702.

Thank you for the suggestion. These references have been also added. We added Tengattini’s NeXT (same facility name as NIST but at ILL) Nuclear Instr. And Methods A 2020 manuscript.

Line 366 - The claim of 100 um resolution is not very-well motivated. Is there a specific structure in Fig 4E that demonstrates this spatial resolution? A scale bar in Fig 4E and/or a more thorough description in the text would solve the problem.

Our resolution is based on Nyquist’s theorem, i.e., the spatial resolution is approximately 3 times the pixel. Here, the pixel size is approximately 36 m. These measurements are done with a SIEMENS star during calibration of the instrument (not during the CT scan). We have added a reference to the resolution mask we use at the beamline (in the protocol section).

A dimension scale is available on the x and y axes (10 mm x 50 mm). Thank you for noticing that the x-scale has been cropped on Figure 5C. It has been corrected. The scale bars are also indicated on the x and y axes. We added a comment in the figure caption for clarification purposes.

Line 493 - The authors should state how long it takes to complete a complementary scan using either a synchrotron or laboratory x-ray source. It is implied that x-ray CT would take much less time, but many readers might not recognize that.

It really depends on the XCT scanner and what resolution is required. We revised the sentence to provide an idea of the time and field-of-view for synchrotron radiation: “This is mostly due to the difference in overall neutron flux available at a reactor compared to a synchrotron source, where X-ray CT scans can take seconds to minutes for a few mm2 field-of-view.”

Reviewer #2:

Dear reviewer,

Thank you for your comments. We value your input and have done our best to respond to your review comments and suggestions. Please see below our response below in green.

Manuscript Summary:
The manuscript describes the workflow for image acquisition and data analysis at the CG-1D neutron imaging beamline at ORNL and demonstrates some imaging applications for biological systems

Major Concerns:
The introduction is very ONRL focused it should be more general even though the purpose of the publication is to describe the workflow at ONRL. It should be extended with some more references to investigation of anatomical structures. E.g., Bone-metal studies have been done by Isaksson, Tudisco, and Le Cann which should be mentioned in the introduction. Lung studies were done at FRM-II before the cited works.
It is also a pity that there are no reports about end applications, the authors mainly show that you can make nice images with neutrons, but not how they can be used in a quantitative analysis.

We thank you for your comments. It is true, indeed, that our manuscript focuses on neutron radiography and computed tomography at ORNL, so we revised our title to reflect this. Our goal was to illustrate how we perform experiment at our facility, rather than a general overview. It now reads “Neutron Radiography and Computed Tomography of Biological Systems at The Oak Ridge National Laboratory’s High Flux Isotope Reactor”. Thank you for pointing the references out. We have added them to the introduction.

Minor Concerns:
Page 4: Lines 197-201: Don't you do (a) and (b) with the CCD?

The reviewer is correct. We can do this for both detectors (CCD and sCMOS), however in this case, the CCD was already setup, so we did not need to tune it. For completeness and comprehension, we will put both detectors in sections (a) and (b).

Page 5: Line 236: What kind of scintillator did you use and what was the neutron flux? Why do you over scan to 182deg? I hope you don't have a fan angle of 2 degrees when you need 150s exposure time. Over scan is anyway only relevant when you do cone/fan beam reconstruction. Also, didn't the camera saturate with 150s exposures? Did you sum up multiple acquisitions?
We use a cone beam geometry. We’ve tried both a parallel beam and a cone beam reconstruction using commercial software at the early stages of the facility. The camera did not saturate since we used a 25 m thick LiF/ZnS scintillator, the only thickness available at the time. Now, we have thicker scintillators, and the camera does saturate in less than 60 s. We have added the scintillator description in the Neutron source and beamline setup section.

Page 7 line 322 don't you use something more sophisticated than median filtering to remove spots? What algorithm are you using to remove ring artefacts?

Median filter is adequate since the affected pixels are limited. In the iMARS3D CT reconstruction workflow, two different algorithms can be used for ring artifact removal

- An algorithm implemented in tomopy that works on the reconstructed images. I did not find a reference for that algorithm

- The Ketcham algorithm we implemented that works on the sinograms. The original paper of Ketcham is here https://doi.org/10.1117/12.680939

We can also use AMIRA for ring filtering after the reconstruction has been performed and slices have been generated (https://www.thermofisher.com/us/en/home/industrial/electron-microscopy/electron-microscopy-instruments-workflow-solutions/3d-visualization-analysis-software/amira-life-sciences-biomedical.html).


Reviewer #3:

Dear reviewer,

Thank you for your comments. We value your input and have done our best to respond to your review comments and suggestions. Please see below our response below in green.

Manuscript Summary:
This manuscript describes the protocol to perform neutron radiography and computed tomography of three biological samples at the High Flux Isotope Reactor (HFIR) CG-1D beamline.

Major Concerns:
* The title is too general, the work presented here is from a specific facility HFIR from Oak Ridge, so I'd recommend changing the title to clearly emphasize where this work was done, e.g. Neutron Radiography and Computed Tomography of Biological Systems at HFIR, Oak Ridge. If the authors want to keep their title, then they have to add references to what has been done on this topic (biological systems) at other neutron imaging facilities (from UK, Switzerland, Germany) on biological systems.

This is a very good point. We did change the title to: “” and we added references from colleagues in Europe in the introduction section (Schillinger in Germany, Burca et al. in the UK, and Kaestner in Switzerland). Thank you for the comment.

* Line 118: The authors say 'This manuscript aims to demonstrate that neutron imaging (nR and nCT) is a complementary technique to X-ray imaging' but unfortunately nothing from X-ray imaging measurements/ protocol is presented here to prove the claim. Entire manuscript is based only on neutron imaging results.

Excellent point. This is a remnant to our original manuscript. Once we had a better understanding of the manuscript focus (in order to publish videos), we changed our initial focus. We have rephrased the sentence as follows:
“ This manuscript aims to demonstrate the procedure of using neutron imaging (nR and nCT) at HFIR CG-1D beamline.” We have moved the comment on the complementarity between neutron and X-ray imaging in the introduction.

Minor Concerns:
1. Line 96: σ_i doesn't appear in equation (2).

Thank you for pointing this out. It should be σ. We have corrected our mistake.

2. Line 202: What kind of material was used for container? What kind of neutron shielding was used for detector and equipment?

Good points. We usually use Al as a container as it mostly does not disrupt the neutron beam. The shielding for the detector is two-fold: boron rubber and Pb bricks, in this order from the beam/sample to the detector. Boron carbide absorbs neutrons but produces gammas. Pb bricks are needed because most samples or sample holders produce gammas. They also scatter neutrons thus the need to absorb them and stop the subsequent gammas too. We have revised the manuscript accordingly.

3. Line 190: 'A photograph of the CG-1D beamline is given in Figure 2' - not true. Please double check the Figure's number

Thank you. Figure 2 now shows a photograph of the facility or CG-1D beamline.

4. Section 1.3.: It is unclear how many samples were measured. Two plants (mulberry weed and mint)? Why mulberry plant was grown in chamber for 1 week? The plant had already 6.3 cm height when it was excavated; How was mint plant prepared before being transplanted in Al container?
Thank you for your comment. The mint plant is not shown in this article, it was a typo that should have been removed before submitting the manuscript. We transplant our plant before the experiment to acclimate them to their new environment (sand) so that water uptake is not measured when the plant is stressed due to the transplant.

We revised the manuscript to indicate the plant preparation, including water cycles and its transplant to the sand soil.

5. Line 242: Is any type of quartz suitable?

Pure polycrystalline quartz is fine. We try to avoid quartz because it is commonly used as glass for chemistry applications and contains a small amount of boron which absorbs neutrons. Ideally, we would measure the container alone to make sure it is not too attenuating. The advantage is the ability to see through which is of interest to plant physiologists.

6. Line 247: How was the water periodically added to maintain the plant during the experiment?

We used a needle that we inserted at the bottom of the Al chamber through a punctured hole. We added water once a day to maintain the plant’s health. We revised the manuscript accordingly.

7. Line 516: Why each sample was reconstructed with a different type of reconstruction software?

We used three software packages, one commercial (Octopus) and another based on the open-source code TomoPY. Octopus has been discontinued and the company that created it does not offer any technical support, so we transitioned to iMARS3D (based on TomoPY). For fast CT (less than 2 hr), we cannot use Octopus nor IMARS3D because they are based on the filtered-back projection method, which requires high signal-to-noise data. Thus, we decided to utilize an iterative method which was more successful than our attempt to use the previously mentioned tools.

8. Line 338: Why is it mentioned only iMARS3D in the protocol steps and not also Octopus or pyMBIR? There is nothing mentioned about the filters (pre- or post-processing filters) in the protocol steps.

Thank you for your comment. We do mention that ring artifacts, median filtering and smoothing are available with our software. We will explicitly demonstrate them during the video. We have added these steps in our protocol for image processing.

9. Line 438: It's Figure 5 instead of Figure 3

Thank you. The figure numbering has been corrected.

10. Line 456: It's Figure 6 instead of Figure 4

Thank you. The figure numbering has been corrected.

11. Line 479: It's Figure 7 instead of Figure 5
Thank you. The figure numbering has been corrected.

Reviewer #4:

Dear reviewer,

Thank you for your comments. We value your input and have done our best to respond to your review comments and suggestions. Please see below our response below in green.

Manuscript Summary:
The manuscript is to demonstrate the neutron radiography and tomography as suitable imaging technique for biological samples. As samples a plant root system, a rat femur implant and a mouse lung. The aim is to demonstrate the applicability of neutron imaging as complementary technique to x-ray imaging. The paper focuses on the advantages of neutron bioimaging as neutrons interact strongly with hydrogen and hydrogen is an abundant material in biological systems.

Major Concerns:
My major concern is the lacking novelty. Neutron imaging to biological systems is applied since longer time. However, when the paper aims to present the methodology without claiming novelty that is okay. Another major concern is that the authors claim to demonstrate neutron imaging as complementary method to x-ray imaging. However, no results on x-ray imaging are shown and the complementary information that are gained are not discussed. Another concern is that the examples that are chosen are not convincing. At first neutron imaging is not non-invasive as gamma radiation could also influence living plants. Also, the plant had to be potted in sand. Furthermore, biological tissues cannot be investigated in vivo but the animals have to be sacrificed for imaging. I don't see the large advantage of neutron imaging in that cases.
Thank you for your comments. We have removed the claim that we are comparing neutrons with X-rays and solely focus the manuscript on neutron imaging. The plants are not damaged by neutron imaging or low energy gammas produced when neutrons are in contact with the Al container. As a matter of fact, plants have survived weeks and months after neutron exposure. The advantage of neutron imaging in animal tissue is the capability to look at thick layers of tissue rather than histology where slices are ~ 7 m. In the case of neutron CT, samples can be a couple of mm thick, thus being equivalent to measuring hundreds of histology slices. In that sense, nCT is not as invasive and destructive.

In the introduction, we have added references from other teams who have measured similar samples as we did.

Minor Concerns:
Please refer to the complementarity between x-ray und neutron imaging.

Thank you for your recommendation. We have added references that directly compare neutrons and X-ray imaging.

line 233: what was the spatial resolution

The spatial resolution of all our measurements is on the order of 75 m (see lines 164-167).

Fig. 4 Please provide scale bars with units

Thank you for your recommendation. The scale bars are indicated on the x and y axes. We added a comment in the figure caption for clarification purposes.

Fig. 5 Please use consistent terminology for the presented methods

Thank you. We have changed the terminology of both Figures 6 and 7 to match.

line 308: Is it possible to visualize the mouse lung in-vivo?

Since we used cold neutrons, they are not energetic enough to pass through tissues that are thicker than 2-3 mm. There has been an attempt in the 1960’s to visualize human lungs on a live person. See reference: J. Anderson, Neutron Radiography of Man, British Journal of Radiology, Vol. 37, 1964.

line 442: complementarity is not shown

Understood. Removed this part of the discussion.

line 137ff: What is the beam intensity? Please provide

Thank you for your comment. The flux on sample is on the order of 107 n/cm2/s. We have added this value in the Neutron source and beamline setup section.

line 240ff: The data acquisition system is very special to the oak ridge laboratory and not of such high use for other neutron sources

Good point. Our manuscript is indeed specific to Oak Ridge, on purpose, as we are trying to illustrate the steps of doing an experiment at our facility. Hence, we have changed the title to reflect this. It now reads: Neutron Radiography and Computed Tomography of Biological Systems at The Oak Ridge National Laboratory’s High Flux Isotope Reactor

line 452: the chosen examples are not appropriate

Thank you for your comment. The chosen examples were selected because they represent 90% of our scientific community in the biological field. In this sense, these examples are representative of the community that utilizes our facility.