Title:

Measurement of tissue oxygenation using near-infrared spectroscopy in patients undergoing hemodialysis

# Authors and Affiliations:

# Kiyonori Ito1,*, Susumu Ookawara1,*, Takayuki Uchida2, Hideyuki Hayasaka2, Masaya Kofuji2, Haruhisa Miyazawa1, Akinori Aomatsu1, Yuichiro Ueda1, Keiji Hirai1, Yoshiyuki Morishita1

1Division of Nephrology, First Department of Integrated Medicine, Saitama Medical Center, Jichi Medical University

2 Department of clinical engineering, Saitama Medical Center, Jichi Medical University

*These authors contributed equally.

# Corresponding Authors:

Kiyonori Ito

Susumu Ookawara

Co-authors:

# Kiyonori Ito kiyonori.ito@gmail.com

# Susumu Ookawara su-ooka@hb.tp1.jp

# Takayuki Uchida t.uchida@jichi.ac.jp

# Hideyuki Hayasaka hide7hao1026@yahoo.co.jp

# Masaya Kofuji la-mer-bleue67@live.jp

# Haruhisa Miyazawa haruhisa_m@hotmail.com

# Akinori Aomatsu ayksera@gmail.com

# Yuichiro Ueda mini_bouz_butterfly@yahoo.co.jp

# Keiji Hirai keijihirai@kfy.biglobe.ne.jp

Yoshiyuki Morishita ymori@jichi.ac.jp

# KEYWORDS:

# tissue oxygenation, hemodialysis, near-infrared spectroscopy, regional oxygen saturation, intradialytic hypotension, hypoxia, diabetes mellitus

# SUMMARY:

We present a protocol to measure regional oxygen saturation (rSO2) in hemodialysis (HD) patients by using a near-infrared spectroscopy monitor. The rSO2 value is an index of tissue oxygenation. This noninvasive and real-time monitoring could be useful for confirming changes in organ oxygenation during HD.

ABSTRACT:

Near-infrared spectroscopy (NIRS) has recently been applied as a tool to measure regional oxygen saturation (rSO2), a marker of tissue oxygenation, in clinical settings including cardiovascular and brain surgery, neonatal monitoring and prehospital medicine. The NIRS monitoring devices are real-time and noninvasive, and have mainly been used for evaluating cerebral oxygenation in critically ill patients during an operation or intensive care. Thus far, the use of NIRS monitoring in patients with chronic kidney disease (CKD) including hemodialysis (HD) has been limited; therefore, we investigated rSO2 values in some organs during HD. We monitored rSO2 values using a NIRS device transmitting near-infrared light at 2 wavelengths of attachment. The HD patients were placed in a supine position, with rSO2 measurement sensors attached to the foreheads, the right hypochondrium and the lower legs to evaluate rSO2 in the brain, liver and lower leg muscles, respectively. NIRS monitoring could be a new approach to clarify changes in organ oxygenation during HD or factors affecting tissue oxygenation in CKD patients. This article describes a protocol to measure tissue oxygenation represented by rSO2 as applied in HD patients.

INTRODUCTION:

Near-infrared spectroscopy (NIRS) has been used to evaluate regional oxygen saturation (rSO2), a marker of tissue oxygenation, especially cerebral oxygenation in various clinical settings1-3 and has recently been applied to patients undergoing hemodialysis (HD)4-11. Cerebral rSO2 is reportedly associated with cognitive function in patients undergoing HD or those with non-dialyzed chronic kidney disease (CKD)11,12. However, thus far, the use of NIRS monitoring has been limited in patients with CKD.

As NIRS monitoring is real-time and noninvasive, we assessed its usefulness as a monitoring device in patients undergoing HD. Although NIRS is mainly used to measure cerebral rSO2, we also investigated rSO2 values in other organs during HD. Specifically, the rSO2 measurement sensors were attached to the forehead, the right hypochondrium and the lower legs to evaluate rSO2 in the brain, liver and lower muscles, respectively. The results showed that NIRS monitoring could be a new approach to clarify changes in organ oxygenation during HD or factors affecting tissue oxygenation in CKD patients.

To date, continuous monitoring was performed during HD, blood volume monitoring, central venous oxygen saturation, thoracic admittance and electronic stethoscope-guided estimated blood pressure (BP) in clinical settings13-15; however, there are limitations for the prediction of hypotension or the wide use of devices. In contrast, the new noninvasive approach here could provide real-time information on intradialytic oxygen dynamics in individual organs. Therefore, this monitoring method may allow the detection of transient organ ischemia in the early phases of intradialytic hypotension and may also permit the safe performance of HD. This article describes a protocol to measure tissue oxygenation represented by rSO2, as applied in patients undergoing HD.

PROTOCOL:

All participants provided written informed consent. The study was approved by the Institutional Review Board of the Saitama Medical Center, Jichi Medical University, Japan (RIN 15–104).

- Device for the monitoring of rSO2

  - Prepare a NIRS device for measuring tissue oxygenation. This device has four channels and can perform measurement in up to four organs at the same time.

  - Prepare a measurement sensor for NIRS monitoring, to evaluate rSO2 values in each organ via transmitting near-infrared light at two wavelengths of attachment.

- Attaching the measurement sensor

  - Allow each patient to rest in a supine position for at least 5 minutes before HD.

  - Attach measurement sensors to the forehead, the right hypochondrium and lower legs to evaluate rSO2 in the brain, liver and lower leg muscles, respectively.

  - Monitoring of cerebral oxygenation

    - Attach measurement sensors to the forehead of the dominant hemisphere.

  - Monitoring of hepatic oxygenation

    - Prepare echocardiography to measure the depth to the patients’ liver from the body surface. Confirm that this measurement is within 20–30 mm from the body surface. Next, attach the measurement sensors to the right hypochondrium.

NOTE: In this device, rSO2 values should be obtained in deep tissue 20–30 mm from the body surface. In some instances, the liver may be located in more than 30 mm from the body surface due to the presence of thick subcutaneous fat.

  - Monitoring of muscle oxygenation

    - Attach measurement sensors to the right or bilateral lower legs.

  - Sensor connection and powering the device

    - Connect each sensor to the leads from the device. Next, turn on the device, and start measuring oxygenation.

- Puncturing the dialysis shunt and starting monitoring

  - Puncturing the dialysis shunt

    - Puncture the patients’ dialysis shunt to start HD therapy. At this time, measure BP using a digital blood pressure monitor equipped with the dialysis machine and collect blood samples using syringes.

  - Start monitoring

    - After starting HD therapy, start monitoring the tissue oxygenation of the three organs: the brain, liver and lower leg muscle.

  - Monitoring of rSO2 during HD

    - Observe changes in rSO2 values of each organ and measure BP regularly in addition to the usual monitoring performed during HD therapy including heart rate, venous pressure and blood volume. Confirm the attachment area and connection between the sensors and leads.

REPRESENTATIVE RESULTS:

Cerebral rSO2 values before HD were lower than those in healthy subjects and cerebral rSO2 in HD patients with diabetes mellitus (DM) were lower than those in HD patients without DM (Figure 1)16. Furthermore, although tissue oxygenation continues without a decrease of BP during HD, we incidentally observed changes in cerebral and hepatic rSO2 due to intradialytic hypotension (Figure 2). Due to the continuous monitoring, the changes in tissue oxygenation were observed more quickly than by intermittently monitored BP. Data were expressed as means ± standard error. The analysis of variance for non-paired values was used to compare three groups.

FIGURE AND TABLE LEGENDS:

Figure 1. Comparison of cerebral rSO2 before HD among HD patients with diabetes mellitus (n = 27), HD patients without diabetes mellitus (n = 27) and healthy subjects (n = 28). The patients included 38 men and 16 women with mean age of 67.7 ± 1.2 years and HD duration of 6.5 ± 1.9 years. The causes of chronic kidney disease were DM (27 patients), chronic glomerulonephritis (14 patients), nephrosclerosis (4 patients), polycystic kidney disease (4 patients), and other (5 patients). The error bars indicate the standard error. The data were based on and the figure has been modified from a previous report16. DM; diabetes mellitus, HD; hemodialysis, rSO2; regional oxygen saturation

Figure 2. Changes in cerebral and hepatic rSO2 in a patient with acute intradialytic hypotension. BP; blood pressure, h; hour, rSO2; regional oxygen saturation, UFR; ultrafiltration rate

DISCUSSION:

NIRS monitoring has been mainly used to evaluate cerebral rSO2, especially in cardiovascular or cerebrovascular surgeries, which require extracorporeal circulation. During extracorporeal circulation including HD therapy, some organs could show relative ischemia7,17,18; however, it remains unclear whether tissue oxygenation becomes low or not. Muscle cramps or abdominal pain during HD could be one of the prodromal symptoms of intradialytic hypotension via organ hypoperfusion. However, in HD therapy, there is currently no method for the real-time evaluation of tissue oxygenation. Therefore, we focused on using this monitoring device to evaluate organ oxygenation using the protocol described above. This protocol is noninvasive for HD patients and is useful for confirming changes in tissue oxygenation in real-time.

As shown in Figure 1, cerebral rSO2 in HD patients with DM was lower than that in patients without DM. Furthermore, higher vascular calcification was associated with lower cerebral oxygenation19. Thus, micro- and macro-vascular disorders could be associated with impairment of cerebral oxygenation. Furthermore, cerebral rSO2 was relatively maintained constant within 60-150 mmHg in HD patients4. However, in intradialytic hypotension, an acute decrease in BP could lead to changes in organ oxygenation (Figure 2). Before observing changes in rSO2 values by this protocol, we could not confirm the influence of tissue oxygenation during HD. Besides continuous arterial pressure monitoring, BP is generally evaluated intermittently. In contrast, the continuous monitoring by NIRS might be able to detect changes in organ oxygenation before being detected by changes in BP during HD. Thus, we could observe the state of hypoxia ahead of confirming the lowering of BP. In addition to changes in BP, blood transfusion, low-density lipoprotein apheresis and ultrafiltration might cause changes in organ oxygenation such as the rSO2 of the lower-legs20-22. Therefore, we should pay attention to acute changes in organ oxygenation during HD.

This protocol has several limitations. First, cerebral rSO2 could be only measured from the forehead; however, it is difficult to perform this evaluation in the posterior brain circulation. As the measurement sensors are a seal type, their sensors could be fixed on the hair. Next, measurement of hepatic rSO2 requires confirmation of the subcutaneous fat thickness. In patients with obesity, the measured rSO2 might be not accurate, because the near-infrared lights could not reach target organs. Third, rSO2 values might be affected by body motion or position (i.e., supine and seated positions). Therefore, during HD, patients should be measured while in their beds and in the same position, as possible.

Furthermore, the rSO2 values measured in this protocol represents mixed venous saturation, which reflects tissue oxygenation in venous (70–80%), capillary (5%), and arterial (20–25%) blood23. Therefore, changes in rSO2 values do not necessarily parallel changes in percutaneous oxygen saturation24,25. Thus, the measured rSO2 values should be carefully interpreted. Furthermore, this protocol is easy to perform and noninvasive for patients if NIRS monitoring device is available. Therefore, this method would provide widely general versatility. We hope that this NIRS monitoring would be equipped with dialysis machines as a dialysis monitor in the future.

In conclusion, we have described a protocol for the measurement of tissue oxygenation by NIRS in patients undergoing HD. This monitoring during HD might provide new findings regarding changes in tissue oxygenation affected by HD therapy.

ACKNOWLEDGMENTS:

We thank the dialysis staffs and members of the department of nephrology in Saitama medical center of Jichi Medical University. We would like to thank Editage (www.editage.com) for English language editing.

DISCLOSURES:

No conflicts of interest.

REFERENCES:

# Nishiyama, K. et al. Regional cerebral oxygen saturation monitoring for predicting interventional outcomes in patients following out-of-hospital cardiac arrest of presumed cardiac cause: A prospective, observational, multicentre study. Resuscitation. 96, 135-141 (2015).

# Kobayashi, K. et al. Factors associated with a low initial cerebral oxygen saturation value in patients undergoing cardiac surgery. Journal of Artificial Organs. 20 (2), 110-116 (2017).

# Cruz, S.M. et al. A novel multimodal computational system using near-infrared spectroscopy predicts the need for ECMO initiation in neonates with congenital diaphragmatic hernia. Journal of Pediatric Surgery. 53 (1), 152-158 (2018).

# MacEwen, C., Sutherland, S., Daly, J., Pugh, C., Tarassenko, L. Relationship between Hypotension and Cerebral Ischemia during Hemodialysis. Journal of the American Socociety of Nephrology. 28 (8), 2511-2520 (2017).

# Polinder-Bos, H.A. et al. Changes in cerebral oxygenation and cerebral blood flow during hemodialysis - A simultaneous near-infrared spectroscopy and positron emission tomography study. Journal of Cerebral Blood Flow & Metablism. 40 (2), 328-340 (2020).

# Ookawara, S. et al. Differences in tissue oxygenation and changes in total hemoglobin signal strength in the brain, liver, and lower-limb muscle during hemodialysis. Journal of Artificial Organs. 21 (1), 86-93 (2018).

# Malik, J. et al. Tissue ischemia worsens during hemodialysis in end-stage renal disease patients. The Journal of Vascular Access. 18 (1), 47-51 (2017).

# Ito, K. et al. Cerebral oxygenation improvement is associated with hemoglobin increase after hemodialysis initiation. The International Journal of Artificial Organs. (2020).

# Valerianova, A. et al. Factors responsible for cerebral hypoxia in hemodialysis population. Physiological Research. 68 (4), 651-658 (2019).

# Ookawara, S. et al. Associations of cerebral oxygenation with hemoglobin levels evaluated by near-infrared spectroscopy in hemodialysis patients. PLoS One. 15 (8), e0236720 (2020).

# Kovarova, L. et al. Low Cerebral Oxygenation Is Associated with Cognitive Impairment in Chronic Hemodialysis Patients. Nephron. 139 (2), 113-119 (2018).

# Miyazawa, H. et al. Association of cerebral oxygenation with estimated glomerular filtration rate and cognitive function in chronic kidney disease patients without dialysis therapy. PLoS One. 13 (6), e0199366 (2018).

# Locatelli, F. et al. Haemodialysis with on-line monitoring equipment: tools or toys? Nephrology Dialysis Transplantation. 20 (1), 22-33 (2005).

# Cordtz, J., Olde, B., Solem, K., Ladefoged, S.D. Central venous oxygen saturation and thoracic admittance during dialysis: new approaches to hemodynamic monitoring. Hemodialysis International. 12 (3), 369-77 (2008).

# Kamijo, Y. et al. Continuous monitoring of blood pressure by analyzing the blood flow sound of arteriovenous fistula in hemodialysis patients. Clinical and Experimental Nephrology. 22 (3), 677-683 (2018).

- Ito, K. et al. Factors affecting cerebral oxygenation in hemodialysis patients: cerebral oxygenation associates with pH, hemodialysis duration, serum albumin concentration, and diabetes mellitus. PLoS One. 10 (2), e0117474 (2015).

- Imai, S. et al. Deterioration of Hepatic Oxygenation Precedes an Onset of Intradialytic Hypotension with Little Change in Blood Volume during Hemodialysis. Blood Purification. 45 (4), 345-346 (2018).

- Cho, A.R., Kwon, J.Y., Kim, C., Hong, J.M., Kang, C. Effect of sensor location on regional cerebral oxygen saturation measured by INVOS 5100 in on-pump cardiac surgery. Journal of Anesthesia. 31 (2), 178-184 (2017).

- Ito, K. et al. Deterioration of cerebral oxygenation by aortic arch calcification progression in patients undergoing hemodialysis: A cross-sectional study. BioMed Research International. 2852514 (2017).

- Ito, K. et al. Blood transfusion during haemodialysis improves systemic tissue oxygenation: A case report. Nefrologia. 37 (4), 435-437 (2017).

- Ito, K. et al. Improvement of bilateral lower-limb muscle oxygenation by low-density lipoprotein apheresis in a patient with peripheral artery disease undergoing hemodialysis. Nefrologia. 39 (1), 90-92 (2019).

- Kitano, T. et al. Changes in tissue oxygenation in response to sudden intradialytic hypotension. Journal of Artificial Organs. 23 (2), 187-190 (2020).

- Lemmers, P.M.A., Toet, M.C., van Bel, F. Impact of patent ductus arteriosus and subsequent therapy with indomethacin on cerebral oxygenation in preterm infants. Pediatrics. 121, 142–147 (2008).

- Ito, K. et al. Sleep apnea syndrome caused lowering of cerebral oxygenation in a hemodialysis patient: a case report and literature review. Renal Replacement Therapy. 4, 54 (2018).

- Minato, S. et al. Continuous monitoring of changes in cerebral oxygenation during hemodialysis in a patient with acute congestive heart failure. Journal of Artificial Organs. (2019).