Editorial comments:

You will find Editorial comments and Peer-Review comments listed below. Please read this entire email before making edits to your manuscript.
NOTE: Please include a line-by-line response to each of the editorial and reviewer comments in the form of a letter along with the resubmission.

Editorial Comments:

• Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammatical errors.

The manuscript has been proofread to ensure that there are no spelling or grammatical errors.
• Avoid punctuating the title.

The title has been edited slightly to ensure that there is not punctuation.
• Protocol Language: The JoVE protocol should be almost entirely composed of numbered short steps (2-3 related actions each) written in the imperative voice/tense (as if you are telling someone how to do the technique, i.e. "Do this", "Measure that" etc.). Any text that cannot be written in the imperative tense may be added as a brief “Note” at the end of the step (please limit notes). Please re-write your ENTIRE protocol section accordingly. Descriptive sections of the protocol can be moved to Representative Results or Discussion. The JoVE protocol should be a set of instructions rather a report of a study. Any reporting should be moved into the representative results.
o Section 2 will need to be edited so that all steps are in the imperative.
o Sections 3-4 will need to be entirely re-written as steps in the imperative voice.
Sections 2-4 have been edited and rewritten as necessary to ensure that they are written in the imperative tense.

• Protocol Detail: Please note that your protocol will be used to generate the script for the video, and must contain everything that you would like shown in the video. Please add more specific details (e.g. button clicks for software actions, numerical values for settings, etc) to your protocol steps. There should be enough detail in each step to supplement the actions seen in the video so that viewers can easily replicate the protocol.
The protocol has been amended accordingly.

• Protocol Numbering:
1. Please removing the numbering from the section headings (e.g., introduction, protocol), and start by numbering line 158 as 1.
2. Please adjust the numbering of your protocol section to follow JoVE’s instructions for authors, 1. should be followed by 1.1. and then 1.1.1. if necessary.
3. All steps should be lined up at the left margin with no indentations.
4. Please add a one-line space after each protocol step.
The numbering has been amended accordingly.

• Protocol Highlight: After you have made all of the recommended changes to your protocol (listed above), please re-evaluate the length of your protocol section. There is a 10-page limit for the protocol text, and a 3- page limit for filmable content. If your protocol is longer than 3 pages, please highlight ~2.5 pages or less of text (which includes headings and spaces) in yellow, to identify which steps should be visualized to tell the most cohesive story of your protocol steps.

1) The highlighting must include all relevant details that are required to perform the step. For example, if step 2.5 is highlighted for filming and the details of how to perform the step are given in steps 2.5.1 and 2.5.2, then the sub-steps where the details are provided must be included in the highlighting.
2) The highlighted steps should form a cohesive narrative, that is, there must be a logical flow from one highlighted step to the next.
3) Please highlight complete sentences (not parts of sentences). Include sub-headings and spaces when calculating the final highlighted length.
4) Notes cannot be filmed and should be excluded from highlighting.

The relevant parts of the protocol have been highlighted.

• Results: Please provide a distinct representative results section and a separate discussion section.

A distinct representative results section has been provided.

• Discussion:

1) JoVE articles are focused on the methods and the protocol, thus the discussion should be similarly focused. Please ensure that the discussion covers the following in detail and in paragraph form (3-6 paragraphs): 1) modifications and troubleshooting, 2) limitations of the technique, 3) significance with respect to existing methods, 4) future applications and 5) critical steps within the protocol.

The discussion has been amended and rewritten to ensure all these sections have been included.

2) Merge the conclusions into the discussion.

The conclusion has been merged with the discussion.

• Figures: Please remove the embedded figures from the manuscript. Figure legends (including those for supplementary files), however, should remain within the manuscript text, directly below the Representative Results text.

All figures have been removed from the manuscript but captions have been left below the representative results.

• Figure/Table Legends: Please expand the legends to adequately describe the figures/tables. Each figure or table must have an accompanying legend including a short title, followed by a short description of each panel and/or a general description.

Done.

• References:

1) Move all hyperlinks (e.g., line 141) to the table of materials.

Done.

2) Reference 13 is incomplete.

Reference 13 has been updated.

3) Please make sure that your references comply with JoVE instructions for authors. Citation formatting should appear as follows: (For less than six authors, list all authors. For more than 6 authors, list only the first author then et al.): [Lastname, F.I., LastName, F.I., LastName, F.I. Article Title. Source. Volume (Issue), FirstPage – LastPage, (YEAR).]

The JoVE style was downloaded and used.

• Commercial Language:JoVE is unable to publish manuscripts containing commercial sounding language, including trademark or registered trademark symbols (TM/R) and the mention of company brand names before an instrument or reagent. Examples of commercial sounding language in your manuscript are NanoSight, NanoSight NS300, Malvern Panalytical NS300

1) Please use MS Word’s find function (Ctrl+F), to locate and replace all commercial sounding language in your manuscript with generic names that are not company-specific. All commercial products should be sufficiently referenced in the table of materials/reagents. You may use the generic term followed by “(see table of materials)” to draw the readers’ attention to specific commercial names.

Commercial language has been removed.

2) Fig 1 should be removed.

Figure 1 has been removed.

• Table of Materials:

1) Please revise the table of the essential supplies, reagents, and equipment. The table should include the name, company, and catalog number of all relevant materials/software in separate columns in an xls/xlsx file.
2) Please sort in alphabetical order.

This has been updated and sorted in alphabetical order.

• If your figures and tables are original and not published previously or you have already obtained figure permissions, please ignore this comment. If you are re-using figures from a previous publication, you must obtain explicit permission to re-use the figure from the previous publisher (this can be in the form of a letter from an editor or a link to the editorial policies that allows you to re-publish the figure). Please upload the text of the re-print permission (may be copied and pasted from an email/website) as a Word document to the Editorial Manager site in the "Supplemental files (as requested by JoVE)" section. Please also cite the figure appropriately in the figure legend, i.e. "This figure has been modified from [citation]."

Comments from Peer-Reviewers:
____________________________________
Reviewers' comments:

Reviewer #1:

Manuscript Summary: The manuscript provides a detailed experimental description of the procedure that was used to measure 1 single sample in 7 different laboratories

Major Concerns:

The results prove that the described protocol works fine for this specific sample. However, the main question is how the user can know that he is making good measurements on any other sample. From own experience, I know that there are some critical factors, such as the degree of dilution. This was fixed in the protocol by specifying the dilution degree (i.e. 50 times dilution), rather than be proposing a criterium that should be met. In the latter case, the protocol could be applied to other samples, which is not the case now.

It is impossible to define “a” protocol for all samples. Protocol development and sample preparation are significant aspects to successfully using the instrument. The authors have however added further information throughout the paper to guide a user on how they can obtain a good measurement. Details regarding dilution factor, information with respect to shape, checking that the diluent is particle free and figures indicating a good and bad detection threshold setting have all been added.

Whereas the protocol was written out in great detail, I miss important information, such as the max jump distance (as mentioned on line 342), and the minimum number of frames over which a particle can be followed

These have been removed for this specific method as these parameters are not to be set by the user in this protocol. A note has been made in the protocol to leave all the processing parameters on their automatic settings.

The most interesting information would be to know what was different between round 1 (figure 4) and round 3 (figure 5). Line 365-366 only provide very general information, from which nothing can be learned: "not maintained as recommended", "inconsistency in sample preparation", "inconsistency in analysis set up". What do you mean ? What can be going wrong in sample preparation or analysis set up ?

This has been clarified further and now reads the following: ‘After the investigation it was found that the systems reporting biggest size variations were either not maintained as recommended or the analysis was affected by inconsistency in sample preparation whereby the dilution step can create variation caused by different pipetting equipment, user operation and technique, and/or measurement set up including the flow cell not being clean, the wrong camera level being used, the image not being focussed properly, and setting the analysis Detection Threshold incorrectly.’

Overall, this reads too much as a commercial leaflet: "look how well we can recover the specified size of a sample with our instrument". That is quite weak as take home message.

The authors would like to point out that the main message was that the ILC validated this specific NTA SOP to introduce confidence and clarity into nanomaterial risk assessment methods. Consistency amongst the labs and the collection of data ensured this. This is an important aspect when comparing results from different laboratories particularly in relation to nanomaterial risk assessment. This has been highlighted in more detail in the abstract, introduction and discussion.

Minor Concerns:

*) Figure 1 has no added value and can be removed

This Figure has been removed.

*) Lines 288-301 have no meaning for readers that do not have the same instrument or software; I would delete them

These lines have been deleted.

*) The Stokes -einstein equatuion was already introduced in line 113, and hence lines 313-317 can be removed

These lines have been removed.

*) I do not understand that Detection Threshold discussion at all: line 330 mentions as many visible particles (red crosses) as possible, whereas line 33 specifies less than 10 red crosses (and line 332 30 to 80 particles)

The reviewer is correct to indicate that this is confusing. Therefore in order to clarify this the authors have rewritten as follows:

The DT value chosen should identify and track as many visible particles as possible (marked as red crosses on the software image screen). As guidance for setting the detection threshold, the number of identified particles in an image should be in the range of 30-80 where no more than 10 red crosses should correspond to sites not considered to be particles by the observer. There should be no more than 5 blue crosses (indicative of noise) observed.

*) do not use a second title on top of the figures (for fig.4 and 5)

The second title on top of Figures 4 and 5 has been removed.

Reviewer #2:

Manuscript Summary: Very nice manuscript that sets out how to preform a measurement using the NS300 system which could be of great benefit to new users.
There have been some previous ILC studies on NTA, and it might be good to reference back to these. It would be nice to see how this version correlates to previous version (NS300 alone, versuses a mixture of LM10s, NS300, NS500 systems)
This paper shows nicely the expected outcome and repeatability of identical systems (assuming no differences in laser intensity, or alignment)

The authors thank the review for their comments. They particular appreciate the suggestion to make reference to previous ILC studies. The authors have done so referring to Roebben et al and Hole et al. in the introduction. This has strengthened the argument for the importance and need for well-developed SOPs for the characterisation of nanomaterials.
Major Concerns:
No major concerns

Minor Concerns:

It might be nice to specify that the inlet and outlet tubings are different diameters.
If these are swapped around does this impact the analysis? or can it damage the flow cell?
The inlet and outlet tubing are different diameters, with the inlet being smaller in diameter than the outlet. Swapping the inlet-outlet tubing connection may cause in over pressuring the flow cell and leaking. This explanation has been added as a note to section 1.3.3.1.

What would be the impact on changing the syringe flow rate, from 25 to say 30 or 5. Is there a load speed that is too fast? perhaps that can be investigated to see what the cuts would be.
It is recommended to set the pump flow speed so that the particles take 10 seconds to travel across the screen as mentioned in section 2.2.6.1. This guidance is the result of previous investigations into what the effect of flow speed on sizing is. Using flow rates faster than this will result in compromising the sizing accuracy of the measurement. The optimal pump speed setting may vary slightly from system to system but a difference of +/-5 units in pump speed will not greatly affect the particle flow speed.

Screen shots of good vs bad focus positions should be included, likewise for selection of detection thresholds.

Since the focus was set automatically, a bad focus will not be achieved. The authors however have included the following screen shots in section 1.4.1 to depict a bad and good detection threshold setting.

It should be pointed out that the results .csv file lists warning about vibration and noise, and that these should be checked to verify results quality.

These warnings also appear in the analysis tab. We added a line “To verify the measurement quality, look at the Analyse tab in the software or check the csv output file for any warnings message or alerts”.

Sections 2.2.4.3 runs into 2.2.5.2, and then back to 2.2.4.5 (page 6), this should be amended
Checking the assembled unit while filled with clean water is a good idea, but I query the necessity to then disassemble the system, clean it and refill it before sample loading
Perhaps this can be clarified, or amended

Yes, while this would not always be needed, for this ILC protocol this step was added as an extra precaution to further reduce any possible variation. This method minimises the risk of any cross contamination of the sample. A note has been added to the relevant section (now section 1.3.4.4 to explain this).

Please comment on the variability of the results for ILC 3, although minor. What is expected to be the main source of error? All systems are considered to be identical (405nm, high sensitivity systems, using autofocus) the difference between lab 4 and 5 for example in figure 5 is striking

For NTA technology, some result variation is always to be expected as a sample rather than the whole sample is seen. Flow helps to reduce this variation however some variation will always be present due to the low number analysis. Regardless, all results meet the ISO 19430 standard for particle sizing.

Reviewer #3:

Manuscript Summary: This paper studies how to use NTA characterize NPs in aqueous media. The authors utilize the technique to conduct a serious experiments and some information has been found. However; this study with the current data lacks the novelty required for publication in Journal of Visualized Experiments.

The authors thank the reviewer for their comments. The aim of this paper was not to have novel data but rather that the ILC validated this specific NTA SOP to introduce confidence and clarity into nanomaterial risk assessment methods. Furthermore to show the vitality of correct instrument set-up and running so as to obtain consistency amongst results. Consistency amongst the labs and the collection of data ensured this. This is an important aspect when comparing results from different laboratories particularly in relation to nanomaterial risk assessment. This has been highlighted in more detail in the abstract, introduction and discussion.

Major Concerns:

The design of the experiment was too simple and the result was obvious.

Although it was a simple measurement we have shown that it is still possible to make errors and not obtain reproducible results. We therefore wanted to highlight that adhering to a developed method and operating the NS300 system in the recommended way can improve the measurement results.

From my point of view, NTA experimental results strongly depend on the operator and software parameters for which the optimal choices must be empirically determined by the user. Therefore, the screen gain, camera level, detection threshold and focus need to be discussed in Results and Discussion.
This protocol aimed to keep the running of the instrument simple to avoid settings having to be determined by the user and ensure consistency amongst users. In the paper we have emphasised further in section 1.3.5.4 that the automatic camera and focus parameters were used. The DT setting is the only parameter that must be determined by the user – guidance on how to set the DT optimally is included in the paper.

(3) The most important is that this manuscript is more like the NanoSight 300 NTA Software Guide than a scientific paper. Most of the information provided in this manuscript can be found in the NanoSight 300 NTA Software Guide. I would suggest the authors to provide more data or a more informative discussion on their data to make it more suitable for the publication.

Novelty is not a requirement for JoVE publication. The JoVE mission is to increase the reproducibility of science so the focus is on the protocol itself.  The importance of this work was to introduce confidence and clarity into nanomaterial risk assessment methods. The discussion has been rewritten to highlight this and include more information regarding the technique’s advantages and limitations. Furthermore reference has been made to another NTA ILC which emphasised the need for a standard protocol. Publishing this specific protocol in JoVE would make it widely accessible for the scientific community to access for the visual characterisation of the size and concentration of spherical nanomaterials.