School of Geography, Earth                 & Environmental Sciences

Dr Andrew Chetwynd

School of Geography, Earth and

Environmental Sciences,

University of Birmingham,

Edgbaston, B15 2TT Birmingham, UK

4th September 2020

Dear Dr. Bajaj,

Please find attached our revised manuscript for consideration as a video article in the Journal of Visualized Experiments “Capillary Electrophoresis Mass Spectrometry approaches for characterization of the protein and metabolite corona acquired by nanomaterials” by Andrew J. Chetwynd, Wei Zhang, Klaus Faserl, James A. Thorn, Iseult Lynch, Rawi Ramautar, and Herbert Lindner.   This is an original contribution and is not under consideration anywhere else.  All co-authors have approved the submission.

We believe the amended manuscript and the point-by-point response to the editor and reviewers comments carefully and thoughtfully address the helpful questions and comments that arose during the review process.

We are confident that the manuscript will be of interest to JoVE readers/viewers, we look forward to receiving, and future comments in due course.

Thank you for your consideration of the manuscript.

Yours sincerely,

___________________________

School of Geography, Earth and

Environmental Sciences,

University of Birmingham

School of Geography, Earth and

Environmental Sciences,

University of Birmingham

Dr. Andrew Chetwynd

Response to editor and reviewers comments.

We would like to thank the editor and the 3 reviewers for their time and consideration of our manuscript and welcome their questions and suggestions that have helped us to improve our manuscript. We are particularly grateful for the quick turnaround of our manuscript by the editorial board and reviewers given the current ongoing worldwide pandemic. All changes in the manuscript have been marked using red text and responses to individual comments can be found below.

Editorial Comments:

• Protocol Detail: Please note that your protocol will be used to generate the script for the video, and must contain everything that you would like shown in the video. Please ensure that all specific details (e.g. button clicks for software actions, numerical values for settings, etc) have been added to your protocol steps. There should be enough detail in each step to supplement the actions seen in the video so that viewers can easily replicate the protocol.

1) Please include an ethics statement before your numbered protocol steps indicating that the protocol follows the guidelines of your institutions human research ethics committee.

We have now added this statement as requested and added that future work using this protocol on human or animal samples would require similar ethical approval.

2) 3.1: What is the source of human plasma?

When we developed the method we used two sources of plasma, one was a non-profit blood bank and sources from the Sanquin blood bank (Leiden, The Netherlands), while the other was sourced from an in house donor who had provided informed consent in line with the ethical guidelines of the institution. For the protocol, we are just splitting 1 sample into 2 halves for the proteomics and metabolomics workflows, and any source of plasma would work but the source to use would depend on the research question being addressed in the particular study.

3) 4.13: unclear what is being done here.

We have now added more detail to this step and removed the term Zip Tip which is a trade name. With the re-organization of the protocol this section is now 4.7 and 4.7.1.

• Protocol Numbering:
1) Please adjust the numbering of your protocol section to follow JoVE’s instructions for authors, 1. should be followed by 1.1. and then 1.1.1. if necessary.

We have changed the protocol to reflect this layout.

2) All steps should be lined up at the left margin with no indentations.

We have changed the protocol to reflect this layout.

3) Please add a one-line space after each protocol step.

We have changed the protocol to reflect the requested layout.

• Protocol Highlight: After you have made all of the recommended changes to your protocol (listed above), please re-evaluate the length of your protocol section. There is a 10-page limit for the protocol text, and a 3- page limit for filmable content. If your protocol is longer than 3 pages, please highlight ~2.5 pages or less of text (which includes headings and spaces) in yellow, to identify which steps should be visualized to tell the most cohesive story of your protocol steps.
1) The highlighting must include all relevant details that are required to perform the step. For example, if step 2.5 is highlighted for filming and the details of how to perform the step are given in steps 2.5.1 and 2.5.2, then the sub-steps where the details are provided must be included in the highlighting.
2) Merge shorter related steps if needed. There can be up to 4 related actions per step.
3) The highlighted steps should form a cohesive narrative, that is, there must be a logical flow from one highlighted step to the next.
4) Please highlight complete sentences (not parts of sentences). Include sub-headings and spaces when calculating the final highlighted length.
5) Notes cannot be filmed and should be excluded from highlighting.

We have made the requested corrections and provided the relevant highlighting.

• Discussion: JoVE articles are focused on the methods and the protocol, thus the discussion should be similarly focused. Please ensure that the discussion covers the following in detail and in paragraph form (3-6 paragraphs): 1) modifications and troubleshooting, 2) limitations of the technique, 3) significance with respect to existing methods, 4) future applications and 5) critical steps within the protocol.

We have made corrections to better suit these guidelines and in response to the reviewers comments and suggestions.

• Figures:
o Please remove the embedded figures from the manuscript. Figure legends, however, should remain within the manuscript text, directly below the Representative Results text.

The requested corrections have been made.

o Fig 2, 3: Define error bars.

These figures (now 3 and 4) are all box plots as such the middle line is the mean, the boxes the interquartile range and the “error bars” are the minimum and maximum values. Due to the reproducibility of the data for these results the typical box shape has been lost.  We have added a note to the figure capture to reflect this.

• Tables: Please remove the embedded Table from the manuscript. All tables should be uploaded to the Editorial Manager site in the form of Excel files. A description of the table should be included with the Figure legends.

We have now uploaded these as excel sheets 

• Commercial Language: JoVE is unable to publish manuscripts containing commercial sounding language, including trademark or registered trademark symbols (TM/R) and the mention of company brand names before an instrument or reagent. Examples of commercial sounding language in your manuscript are Rapigest SF™, MilliQ, ZipTip, nanospray, etc
1) Please use MS Word’s find function (Ctrl+F), to locate and replace all commercial sounding language in your manuscript with generic names that are not company-specific. All commercial products should be sufficiently referenced in the table of materials/reagents. You may use the generic term followed by “(see table of materials)” to draw the readers’ attention to specific commercial names.

This has now been accounted for in this manuscript, we do however have one remaining query for this. The Rapigest used to solubilize the peptides in step 4.4 is a tradename (Waters Ltd.) however other surfactants proved to be inadequate. Should we say surfactant here and then state in the materials tables that it’s Rapigest SF which is what we have changed it to for now?

The term nanospray is not a commercial name but a technical name for the type of source being used.

• Table of Materials:
1) Please revise the table of the essential supplies, reagents, and equipment. The table should include the name, company, and catalog number of all relevant materials/software in separate columns in an xls/xlsx file. Please include items such as
2) Please remove the registered trademark symbols TM/R from the table of reagents/materials.
3) Please sort in alphabetical order.

We have now amended our Table of Materials to reflect this request.

• If your figures and tables are original and not published previously or you have already obtained figure permissions, please ignore this comment. If you are re-using figures from a previous publication, you must obtain explicit permission to re-use the figure from the previous publisher (this can be in the form of a letter from an editor or a link to the editorial policies that allows you to re-publish the figure). Please upload the text of the re-print permission (may be copied and pasted from an email/website) as a Word document to the Editorial Manager site in the "Supplemental files (as requested by JoVE)" section. Please also cite the figure appropriately in the figure legend, i.e. "This figure has been modified from [citation]."

Reviewers' comments:
Reviewer #1: 
Manuscript Summary:
The manuscript is very interesting and describes the experimental procedures to prepare a plasma sample to evaluate proteins and metabolites adsorption on nanomaterial particles, in order to evaluate the formed corona. It requires English revision in terms of orthography, text composition, punctuation, and others. The results are interesting to be published in this journal and are an innovation in terms of concomitant proteome and metabolome evaluation on corona nanomaterials using CESI-MS - a technique that has not been previously been used for this purpose.

The authors thank the reviewer for their positive critique of our work, we believe that we have now addressed all of their comments and queries in our response below.

Major Concerns:
Line 134: where was this sample obtained? Were ethics issues considered? What was the criteria for selection of this single sample?

This is a protocol as such we didn’t obtain samples for this manuscript, the representative results come from previous studies we have carried out and they had relevant ethical approval which can be found in the cited research. We do however appreciate the concern over ethical approval and have now added a statement on the relevant ethical approval from the research institution. This information has been added now (see lines 119-120). The choice of human plasma was taken due to its prevalence in biomolecular corona studies and, therefore, used in this protocol for the characterization of the biomolecular corona. However, this protocol would also be applicable to serum or other biofluids such as CSF or urine which we now state in the introduction (see lines 111-112).

Line 140: but if the aim of this part of the experiment was the corona NM metabolome characterization, why was the supernatant analyzed? I suppose the non-adsorbed metabolites would be analyzed in the other plasma aliquot (which was not incubated with the NM), and a comparison between total metabolome with metabolome adsorbed on NM would be performed)

The reviewer is correct, unlike the proteins in the corona there is no “one method suits all” approach to characterize the metabolites adsorbed to the nanoparticles due to their very different chemistries. As such, we characterize the matrix before and after NM exposure. This is described in the introduction lines 81-85. We have also added to the discussion the relative drawbacks of characterizing the metabolite corona using this approach in response to a similar comment from Reviewer 2.

Lines 141 and 142: I understand that metabolome absorbed on NM should also be analyzed in this sample

In this work, we have examined which metabolites have been depleted from the matrix rather than eluting adsorbed metabolites from the NM surface. This approach also provides information on the extent of metabolite adsorption to the NMs when employing proper controls and suffers from less uncertainty regarding recovery efficiency for each individual metabolite and is less costly than analyzing the amounts recovered from the NMs and the amounts remaining in solution and performing mass balance to understand if some of each metabolite is still attached to the NM surface.

Figure 1: a discussion on the difference of different proteins classes on the surface of different NMs should have been performed

This is a representative result of the protocol rather than a research result per se; Our aim here is just to demonstrate its applicability, while the primary research on which this protocol is based thoroughly discusses the implications of these proteins in the corona. We have however added a sentence to the legend to state that this protocol allows for easy differentiation between the corona composition of different nanomaterials.

Figure 2: why proteome NM corona was evaluated on 6 NMs while metabolome was evaluated with 4 NMs?

As with Figure 1 these are representative results of the protocol, in the primary research these representative results come from all 6 materials were indeed studied but the PSNPs did not adsorb the metabolites being investigated. We have added a note to this effect to the figure legends to avoid confusion for the readers.

Minor Concerns:
1) It seems that the submitted version was not the final one, since some English mistakes have been found, and texts highlights are observed (as well some marked corrections)

We have carried out another proof read and corrected English language issues. The highlighted text in this version is at the request of JoVE for the segment to be filmed.

2) What was the equipment configuration (describe equipment details)? What was the nanomaterial composition?

The CE set up is described through steps 6 and 7 of the protocol, the nanosource being used is compatible with a wide range of mass spectrometer manufacturers and mass spectrometry types (i.e. time of flight, Orbitraps, FT-ICR, triple quads). Therefore, the technical aspects of the mass spectrometry side are not included as most are manufacturer and mass analyzer dependent. However, where these details are universal such as ionization mode, fragmentation and m/z range they have been included in steps 6-8.

We have added the details of the NMs we applied this too in the Note at the start of section of the protocol lines 120-126. We also make it clear that these can be substituted for any other NMs that may be of interest, we do not wish to limit the readers choice of material inadvertently by implying that only these materials are acceptable for this method as we do not require any NM specific steps. The NMs used to develop these methods have been detailed in the Materials Tables, in the case of the TiO2 these were bespoke particles though we provide details of the manufacturer,

Lines 193 - 195: why is it necessary to flush a neutral coated capillary with NaOH on the conditioning step?

The reviewer is correct - this is an error from when the manuscript was restructured prior to original submission and has now been deleted.

Figure 1: insert the color legend according to proteins abundance

We have added a color legend to the latest version of the figure.

Line 318: include "not" between "is" and "significant"

We are confused as to why the reviewer suggests this as this would mean that we are saying

that there is no scope to investigate the protein and metabolite coronas. The aim of this protocol is to enable this.

Reference number 14 has already been cited (10); I mean, it has been cited twice.

We thank the reviewer for their observation; we have now corrected this in the manuscript.

Reviewer #2: 
Manuscript Summary:
Andrew et al. present a protocol for the joint characterization of protein and metabolite coronas on nanomaterials (NMs). The authors briefly introduce state-of-the-art research and protocols to determine the protein and metabolite corona on NMs after incubation in biological fluids. The need for optimized protocols and various approaches for characterization of NM coronas are mentioned and discussed. The protocol is divided in eight sections, comprising steps of sample and buffer preparation, preconditioning of the CE instrument and capillary and data acquisition. The NMs are incubated in human blood plasma and protein and metabolite corona isolation are performed subsequently from the same sample. Preconditioning protocols to set up the CE-MS system for analysis are described in detail for both methods. Relevant instrumental parameters for data acquisition are provided.
The authors demonstrate the capabilities of their protocol with representative results for protein and metabolite coronas on NMs of various materials. Intra- and inter-day reproducibility for migration times and peak areas are provided. The authors discuss the impact and possible application of their protocol in the context of nanomedicine. The advantage of the joint characterization of proteins and metabolites on the NMs corona are mentioned and discussed. The results are compared with relevant publications in the field. The advantages of CE-MS over nanoLC-MS are briefly mentioned and discussed.

We thank the reviewer for their thoughtful and insightful comments on our manuscript. We have adjusted our work to address their comments and concerns. These changes are detailed below alongside their comments.

Major Concerns:
The authors should consider to briefly point out important steps in data analysis. This also includes strategies for quantification (I see that you have determined absolute concentrations for metabolites in plasma. What type of calibration are you using?)

We have not included data analysis, as the focus of this protocol is the sample preparation to isolate the corona and preparation of these samples for CE-MS analysis. Given the broad range of vendor specific and open access software describing a single approach would limit the applicability of this protocol. However, we have added to the discussion lines 340-343 that an absolute quantification for metabolites can be performed as well as an untargeted approach (though we do not have representative data to show for this with this particular application yet) whereas the protein corona aspect was developed using a bottom up approach. We have added a line in the introduction 101-103 to clarify this and explain why we took this approach.

Minor Concerns:
81-88:
The authors should discuss problems regarding this indirect determination of the metabolite corona. The approach only works when NMs are concentrated enough to result in a significant change in metabolites concentration by adsorption. Consequently, this method might be difficult for biologically/medicinally relevant NM concentrations. The same problem arises when metabolites with high concentrations (compared to the NMs) are considered.

This is an excellent point the reviewer has made, we have added a discussion of these limitations to the manuscript on lines 297 to 306.

105-107:
I agree in terms of throughput and carry-over with the authors. CE-MS offers fast separation and equilibration and due to the lack of stationary phase, carry-over is reduced. However, reproducibility has always been an issue in CE-MS. Having a look at your RSD data on migration times of cationic and anionic metabolites I agree that RSD values of ~2% are good for CE, however they are not comparable to RSD in retention times usually observed in (nano)LC (0.1-0.5%). The authors should cite relevant literature underlining their claim or discuss this point more detailed.

We have changed the sentence slightly to remove reference to analytical flow LC-MS and now refer to only nanoflow LC-MS to make a fairer comparison concerning the flow rate. While the reviewer is correct that CE-MS has previously had issues with RT reproducibility, the same is also true of nanoLC-MS; however recent advances in technology for both separation approaches have improved RT/MT reproducibility. We include a reference to a recent review of nLC-MS which finds that reported RT RSDs range from 0.2 – 3.4 % across 25 studies covering peptides and metabolites.

122/125:
Is the pH value determined/measured after dilution of the acid?

If the quality of acid and DI water being used is as stated in this manuscript then adding the stated volumes will reproducibly generate the same pH in our experience.

130-132:
Please provide a reference for the determination of particle size (ideally the protocol used by the authors)

We have added references to relevant methods; we suggested a range of approaches due to the diverse array of NMs being studied and certain sizing techniques are not suitable for all NMs.

136/137:
Please provide information on how to add the NMs to the plasma (are they solved in a stock solution or are they added as solid?)

This is explained in what is now step 2.1: particles are to be added to the plasma as a solution dispersed in water. We have also stated in what is now step 3.2 what concentration to use for the plasma incubation.

218/219:
Reading HCD" I suggest that you are using an Orbitrap instrument for data analysis. This should be mentioned here since the fragmentation parameters (28% normalized energy) cannot be directly transferred to instruments from other vendors.

The reviewer is correct that the proteomics portion of this approach was developed using an Orbitrap instrument. However, JoVE do not permit the use of brand names. As such we have changed this to state DDA/DIA fragmentation methods, as this CE set-up will work across a broad range of instrument manufacturers and the reviewer is correct that 28% normalisation energy does not translate to QToF instruments.

221:
Is HCl directly flushed towards the MS interface? Do you encounter any problems with contaminating the MS by flushing with non-volatile compounds?

Yes, the HCl is flushed towards the MS interface. However, a 90 cm capillary with internal diameter of 30 µm and just 100 PSI for 5 minutes pushes a little over 1 µL total volume of HCl through the capillary. At this time there is no ESI voltage being applied so the HCl does not reach the interface and evaporates at room temperature, thus it does not contaminate the MS.

232:
Please provide detailed information about internal standards used for the protocol.

These details have now been added to the protocol and to the materials table.

236-238:
I suppose that for cationic metabolite analysis you have residual EOF (despite the low pH value of the BGE) in your system since you do not need to support the separation with additional pressure. Can you comment on this?

Due to the use of a bare fused silica capillary for metabolomics we indeed retained a residual EOF at the employed BGE conditions and, therefore, an additional pressure step was not required. This was only needed with systems without an EOF, such as with the neutral coated capillary used for the proteomics.

Figure 1:
What does the percentages in the heat map represent? How were these values determined?

As suggested by the reviewer we have amended the caption for the table to state that the percentages reported refer to the percentage of overall protein abundance based upon the intensity of the 3 most intense peptides in the protein.

Table 1:
The numbers of (technical?) replicates used to calculated RSD should be provided with table 1.

We have revised as suggested and stated in the legend that these are technical replicates.

324-330:
I do not support the claim that CE-MS is a more green technology, especially compared to nanoLC. Solvent consumption in nanoLC-MS (commonly used in proteomics) is very low and in my view negligible compared to the high energy demands of the mass spectrometers and the high consumption of single-use plastic in laboratories.

We have tempered our statement over solvent usage only and stated the economic impact of this. Still when compared to nLC-MS for proteomics the “flow rate” is at least 10x lower in CESI and no organic solvents are used other than approximately 2 µL of methanol for the conditioning of the bare fused silica capillary. Compared to the flow rates used for metabolomics the CESI “flow rates” could be up to 100x lower and also use hardly any organic solvent. Reducing the environmental impact even in small ways can be regarded as an improvement even if power demands of the MS (which could be from reusable sources) or high use of single use plastic in labs have bigger environmental footprints overall.

Reviewer #3: 
Manuscript Summary:
The authors described a detailed protocol for characterization of proteins and metabolites adsorbed on nanoparticles using capillary electrophoresis mass spectrometry. The protocol is well written and provides detailed information, which will be useful for other researchers in the field. I have some minor comments to help the authors improve the manuscript further.

We thank the review for their positive appraisal of our work and we hope the following changes to our manuscript adequately address their comments.

Minor Concerns:
1) In terms of figures, some example electropherograms of peptides and metabolites will be useful for the readers.

As suggested, we have now included representative electrophoretograms for the proteins and metabolites as figure 1 and 2 respectively.

2) Lines 98-99: Two nice review papers about CE-MS for proteomics could be cited here. Zhang et al. Trends in Analytical Chemistry. 2018, 108, 23-37. Shen et al. Trends in Analytical Chemistry,

2019, 120, 115644.

We have added both of the suggested references to the manuscript.

3) Lines 157-158: Why alkylate the sample at the peptide level not at the protein level as usual?

Alkylating after digestion removes the need to quench the alkylation step and thus reduces the number of steps required.

4) The authors used 30-cm-long capillaries for CE-MS with the CESI instrument. Based on our experience, it is hard to use that short of capillaries for CE-MS. Please double check.

We thank the reviewer for highlighting this, the lengths stated were an error, which have now corrected to 90 cm.

5) Lines 193-195: Is it necessary to flush the neutral capillary with NaOH before CE-MS?

The reviewer is correct this is an error from when the manuscript was restructured during drafting of the manuscript. To clarify NaOH on a neutral capillary would strip the dynamic neutral coating and is thus not used.

6) Figure 1, are the percentages calculated by mass? Some explanation will be helpful.

As suggested by the reviewer we have amended the caption for the table to state that the percentage reported refers to the percentage of overall protein abundance based upon the intensity of the 3 most intense peptides in the protein.