# Submission ID #: 61765

# Scriptwriter Name: Bridget Colvin

Project Page Link: https://www.jove.com/account/file-uploader?src=18838373

Title: Obtaining Quality Extended Field-Of-View Ultrasound Images of Skeletal Muscle to Measure Muscle Fascicle Length

Authors and Affiliations: Amy N. Adkins1,4,5 and Wendy M. Murray 1,2,3,4,5

1Department of Biomedical Engineering, Northwestern University

2Department of Physical Medicine & Rehabilitation, Northwestern University Feinberg School of Medicine

3Department of Physical Therapy and Human Movement Sciences, Northwestern University Feinberg School of Medicine

4Shirley Ryan AbilityLab (formerly the RIC)

5Edward Hines, Jr. VA Hospital

☐   All author names are spelled correctly, and the affiliations are correct (city/state/country information not included in video title page).

# Corresponding Author:

Wendy M. Murray

w-murray@northwestern.edu

# Co-Authors:

adkins@u.northwestern.edu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or similar?  N

2. Software: Does the part of your protocol being filmed demonstrate software usage?  Y

If Yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps. Please upload all screen captured video files to your project page as soon as reasonably possible.

3. Interview statements: Considering the Covid-19-imposed mask-wearing and social distancing recommendations, which interview statement filming option is the most appropriate for your group? Please select one.

☐ 	Interviewees wear masks until the videographer steps away (≥6 ft/2 m) and begins filming. The interviewee then removes the mask for line delivery only. When the shot is acquired, the interviewee puts the mask back on. Statements can be filmed outside if weather permits.

☐ 	Interviewees self-record interview statements outside of the filming date. JoVE can provide support for this option.

☐ 	Interview Statements are read by JoVE’s voiceover talent.

☐ 	Author interview statement opt out (interview statements removed completely from paper).

4. Filming location: Will the filming need to take place in multiple locations (greater than walking distance)?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

- To ensure that your script can be filmed in one day, the Introduction, Protocol, and Conclusion sections are restricted to 55 shots (i.e., designated as 1.1.1., 1.2.1., etc).

Protocol Length

Number of Shots: 44

# Introduction

- Introductory Interview Statements

Your answers to these questions will become author interview statements, which will be incorporated into the script. Authors will memorize the statements and then deliver them on camera.

- Enter the full name of the author who will deliver the statement on camera.

- Each author should deliver no more than two statements.

- Fill out both required statements. One optional statement may also be selected.

- Please answer in full sentences in a style suitable for being spoken aloud.

- Limit the length of each statement to 30 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: Why is your protocol significant? OR What key questions can this method help answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

REQUIRED: What is the main advantage of this technique?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Do the implications of this technique extend toward the therapy (or diagnosis) of a particular disease, disability, or challenge? How so?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Are there any specific areas of research that this method could provide insight into? OR Can this method be applied to any other systems?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: How would you expect an individual who has never performed this technique to struggle? Do you have any advice to offer to somebody who is trying this technique for the first time?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Why is visual demonstration of this method critical?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Has any author been assigned more than two statements? ☐ No

# Introduction of Demonstrator on Camera

Complete this statement ONLY if any of the individuals who will be demonstrating the procedure on camera will not be delivering an Introductory Interview Statement.

  - Enter name of author who will introduce demonstrator: Demonstrating the procedure will be Click here to enter name of demonstrator(s)., a Click here to enter demonstrator job title. from my laboratory. Include additional demonstrators as needed.[1][2].

    - INTERVIEW: Author saying the above

    - The named demonstrator(s) looks up from workbench or desk or microscope and acknowledges the camera

Ethics Title Card

  - Procedures involving human subjects have been approved by the Institutional Review Board (IRB) or equivalent body at insert Institutional Name.

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The one-digit numbers represent sections of the video. The text will appear onscreen.

- The two-digit numbers (e.g. 2.1., 2.2.) represent steps of your protocol. The text will be recorded by a professional voiceover talent.

- The three-digit numbers (e.g. 2.1.1., 2.2.2.) represent the shots that our videographer will capture at your lab.

Please use this draft script to help you prepare for filming day.

- Filming should take no more than 10 minutes per step. If a step will take more than 10 minutes, prepare the product from that step in advance.

- Participant Preparation

  - After explaining the study protocol, seat the Participant in an adjustable chair that can be locked in place [1] and help the Participant adjust the chair so that they are as comfortable as possible while still providing access to the muscle of interest [2].

    - WIDE: Talent gesturing/Participant sitting in chair

    - Talent adjusting chair

  - Place the joint that the muscle of interest spans in a posture that can be controlled and repeated [1-TXT] and use clinical guidance to locate anatomical landmarks [2].

    - Talent adjusting joint position TEXT: See text for participant position and joint angle suggestion details

    - Talent locating landmark(s)

  - Using the ISB (I-S-B) standards for defining the joint coordinate system, measure the joint angle [1] and mark the anatomical landmarks with a skin safe marker [2].

    - Talent measuring joint angle

    - Talent marking landmark(s)

  - Align the center of a handheld goniometer with the axis of rotation of the joint [1] and align the arms of the goniometer with the joint segments [2].

    - Talent aligning center with joint

    - Talent aligning arms with segments

  - Then use cloth straps to secure the limb to minimize movement during the imaging protocol [1].

    - Talent securing limb

- Image Acquisition

  - For image acquisition, plug in and turn on the ultrasound system [1] and ensure that the exam is set to musculoskeletal, the transducer is set to the transducer in use, and the transmit frequency is set between 5-17 megahertz [2].

    - WIDE: Talent plugging in and turning on US system

    - Talent setting exam conditions

  - In the system software, set the footswitch to start-stop the imaging [1]. If the footswitch in use has multiple pedals, set additional pedals to “Freeze” or “Pause” and to “Print” or “Store” the image [2].

    - Talent selecting start/stop imaging

    - SCREEN: To be provided by Authors: Additional pedal options being set and Print and/or Store being selected

  - Next, apply a generous amount of ultrasound gel to the head of the transducer [1] and place the transducer onto the Participant’s skin in the approximate region of interest [2].

    - Talent applying gel to transducer

    - Talent placing transducer onto skin

  - Move the transducer across the short axis plane of the muscle, keeping the indicator pointed laterally [1].

    - Transducer being moved across short axis plane Video Editor: please emphasize indicator

  - Identify the muscle of interest in the short axis plane perpendicular to the muscle fiber direction [1] and move the transducer distally and proximally to acquire a full visualization of the muscle path [2], using the skin safe markers to mark important anatomical landmarks as they are identified [3].

    - SCREEN: To be provided by Authors: Shot of muscle in short axis plate

    - Talent moving probe, with monitor visible in frame as possible

    - Talent marking skin

  - Once the location of the muscle has been identified and marked, move the ultrasound transducer in the long axis plane with the indicator pointed distally [1].

    - Talent moving probe in long axis Video Editor: please emphasize indicator

  - When either the distal or proximal end of the muscle has been identified, rotate and tilt the transducer to identify the fascicle plane at that point [1] and mark the skin when the correct transducer position has been established [2].

    - Talent rotating and tilting transducer/identifying fascicle plane

    - Talent marking skin

  - Once the approximate fascicle plane has been established along the entire desired length to be scanned, practice following the muscle path [1] and set the ultrasound system to the extended field-of-view mode [2].

    - Talent practicing following muscle path

    - Talent setting system to EFOV

  - Starting at one end of the muscle, click the footswitch to start the image acquisition [1] and slowly and continuously move the ultrasound transducer along the long axis [2].

    - Talent clicking footswitch

    - Talent moving transducer, with monitor visible in frame

  - When the end of the muscle has been reached, click the footswitch to end image acquisition [1].

    - Talent clicking footswitch

- Image Optimization

  - To optimize image the visibility and clarity, if the image acquisition ends before the desired length of the muscle can be captured [1], increase the depth of the image [2].

    - WIDE: Talent checking image, with monitor visible in frame

    - SCREEN: To be provided by Authors: Depth being increased

  - To adjust the focus, place the focus arrow in the lower half of the Image just below the muscle of interest [1].

    - SCREEN: To be provided by Authors: Arrow being placed

  - Ensure that the gain is balanced through the depth of the image and use the indicator to determine the optimal speed for moving the transducer [1].

    - SCREEN: To be provided by Authors: Gain being checked, then shot of indicator speed

  - Once qualitatively good images have been collected, press the Print-Store footswitch pedal [1] and export the images as uncompressed DICOM (dye-com) images [2].

    - Talent pressing footswitch

    - SCREEN: To be provided by Authors: Image being exported

- Extended Field-of-View Ultrasound (EFOV-US) Quality Assessment

  - When all of the images have been acquired, open a long-axis image of interest in an appropriate image analysis software program [1] and confirm that the muscle can be identified as a dark hypoechoic shape with bright hyperechoic boundaries that represent the deep and superficial muscle fascia [2].

    - WIDE: Talent opening image, with monitor visible in frame

    - SCREEN: To be provided by Authors: Image of muscle Video Editor: please emphasize dark hypoechoic shape and bright hyperechoic boundaries when mentioned

  - Identify the connective tissue surrounding a muscle fascicle as bright hyperechoic lines between the muscle boundaries [1] and check that the image does not have excessive bending, as indicated by shadows or gaps in the image [2] or a jagged flexible ruler line over the image [3].

    - Use 5.1.2. Video Editor: please emphasize muscle fascicle when mentioned

    - SCREEN: To be provided by Authors: Shot of image with excessive being Video Editor: please emphasize excessive bending/shadows/gaps

    - SCREEN: To be provided by Authors: Shot of image with jagged flexible ruler line

  - If the image is missing one or more these tissue structures, deem the image “qualitatively poor” [1] and return to live 2D-mode [2].

    - SCREEN: To be provided by Authors: Image of “qualitatively poor” image Video Editor: please add “Qualitatively poor” text over image or similar

    - Talent selecting live 2D mode

- Muscle Fascicle Length

  - To quantify the muscle fascicle length, first open the image of interest in ImageJ [1] and use the Straight Line tool to draw a straight line from 0 to 1 centimeter on the ruler on the side of the ultrasound image [2].

    - WIDE: Talent opening image in ImageJ, with monitor visible in frame

    - SCREEN: To be provided by Authors: Tool being selected and line being drawn

  - Select Analyze and Measure to measure the line. If the image properties have been preserved, the length of the straight line should be 1 centimeter [1].

    - SCREEN: To be provided by Authors: Analyze and Measure being selected Video Editor: please indicate 1 cm measurement value

  - To measure the fascicle lengths in the image, right click on the Straight Line tool to select Segmented Line. Click one end of the fascicle to be measured and continuously click along the muscle path to ensure that the curvature in the fascicle has been captured [1].

    - SCREEN: To be provided by Authors: Tool being right clicked, Segmented Line being selected, then fascicle end and path being clicked

  - Then double-click at the end of the fascicle and select Analyze and Measure to measure the length of the line [1-TXT].

    - SCREEN: To be provided by Authors: End being double clicked, then Analyze and Measure being clicked TEXT: Repeat for each fascicle

## Protocol Script Questions

# Authors: Please use the step numbers from the script above (not step numbers from the manuscript) when answering the questions below. Please do not include steps that will be screen-captured and do not list entire sections.

A. Which steps from the protocol are the most important for viewers to see? Please list 4 to 6 individual steps.

Click here to list 4 to 6 individual steps, using the step numbers from the protocol section of the video script.

B. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1 or 2 individual steps from the script above.

Click here to list 1 or 2 individual steps, using the step numbers from the protocol section of the video script.

# Results

Please review this section to make sure that it accurately reflects your findings.

- Use Track Changes when making edits or revisions.

- If you would like the video to include different results, please revise this section.

- When revising, please keep the length of the voiceover below 200 words. Current word count: 167. (Voiceover is the text that follows the two-digit numbers.)

- Please note that the video cannot include voiceover without an accompanying visual.

- Notes in blue italics are for JoVE’s video editor.

- Results: Representative EFOV Muscle Imaging

  - In these representative analyses, extended field-of-view ultrasound [1] was used to obtain images from the long head of the biceps brachii and the tibialis anterior in 4 healthy volunteers [2].

    - LAB MEDIA: Table 1

    - LAB MEDIA: Table 1 Video Editor: please emphasize Bicep Side and Tibialis Anterior Side columns

  - Extended field-of-view ultrasound imaging of these muscles [1] was used to observe important aspects of the tissues, such the central tendon [2] and fascicle path [3].

    - LAB MEDIA: Figure 1 Video Editor: please emphasize dark blue outlines

    - LAB MEDIA: Figure 1 Video Editor: please emphasize light blue outlines

    - LAB MEDIA: Figure 1 Video Editor: please emphasize white lines

  - After imaging, qualitatively “good” images were then analyzed for each muscle in each individual [1].

    - LAB MEDIA: Figure 2 Good images

  - ImageJ was implemented to measure four fascicles [1] with paths that could be convincingly visualized from their origin to insertion and that were located in different portions of the muscle in eachy image [2].

    - LAB MEDIA: Figure 2 orange outline images

    - LAB MEDIA: Figure 2 orange outline images Video Editor: please emphasize white dashed lines

  - The average fascicle lengths [1] obtained in this study for the biceps brachii [2] and the tibialis anterior were within the range of previously reported fascicle lengths [3].

    - LAB MEDIA: Table 1

    - LAB MEDIA: Table 1 Video Editor: please emphasize Bicep Fascicle Length data column

    - LAB MEDIA: Table 1 Video Editor: please emphasize Tibialis Anterior Fascicle Length column

  - As it can be difficult to determine which images are appropriate for fascicle evaluation, here representative qualitatively “bad” images are shown [1], with the portions of the images that are specifically bad highlighted [2].

    - LAB MEDIA: Figure 2 Bad images

    - LAB MEDIA: Figure 2 Bad images Video Editor: please emphasize blue outlines

  - In these images, variations in the muscle quality and landmarks for different individuals can be observed [1].

    - LAB MEDIA: Figure 3

# Conclusion

- Conclusion Interview Statements

Below are prompts for interview statements that can be used to further emphasize the significance of your protocol.

- Please answer one or two of the statements.

- Each statement is limited to 30 words.

- Answer the questions in full sentences; you will need to memorize and deliver the sentences as spoken interview statements during filming.

- Indicate the full name of the author who will deliver the statement.

# What is the most important thing to remember when attempting this procedure? Please indicate the steps (e.g., 2.4., 2.5.) in the Protocol section of the script that this advice applies to.

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera (Enter step numbers referred to.)

# Following this procedure, what other methods can be performed? What questions would these additional methods answer?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

# After its development, did this technique pave the way for researchers to explore new questions within a specific scientific field? If so, how?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Thank you for addressing our questions. We will incorporate your answers and suggestions and send you the final script before your filming day. You will also receive detailed preparation instructions in the email accompanying the final script.