# Submission ID #: 61769

# Scriptwriter Name: Bridget Colvin

Project Page Link: https://www.jove.com/account/file-uploader?src=18839593

Title: A Novel Inhalation Mask System to Deliver High Concentrations of Nitric Oxide Gas in Spontaneously Breathing Subjects

Authors and Affiliations: Riccardo Pinciroli1, Lisa Traeger1, Anna Fischbach1, Stefano Gianni1, Caio C. A. Morais1, Bijan Safaee Fakhr1, Raffaele Di Fenza1, Ryan Carroll2, Warren M. Zapol1, and Lorenzo Berra1

1Department of Anesthesia, Critical Care and Pain Medicine, Massachusetts General Hospital and Harvard Medical School

2Division of Pediatric Critical Care Medicine, Department of Pediatrics, Massachusetts General Hospital and Harvard Medical School

☐   All author names are spelled correctly, and the affiliations are correct (city/state/country information not included in video title page).

# Corresponding Author:

# Lorenzo Berra

# lberra@mgh.harvard.edu

# Co-Authors:

rpinciroli@mgh.harvard.edu

ltrager1@mgh.harvard.edu

afischbach@mgh.harvard.edu

sgianni@mgh.harvard.edu

caraujomorais@mgh.harvard.edu

bsafaeefakhr@mgh.harvard.edu

rdifenza@mgh.harvard.edu

rcarroll4@mgh.harvard.edu

wzapol@mgh.harvard.edu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or similar?  N

2. Software: Does the part of your protocol being filmed demonstrate software usage?  N

3. Interview statements: Considering the Covid-19-imposed mask-wearing and social distancing recommendations, which interview statement filming option is the most appropriate for your group? Please select one.

☐ 	Interviewees wear masks until the videographer steps away (≥6 ft/2 m) and begins filming. The interviewee then removes the mask for line delivery only. When the shot is acquired, the interviewee puts the mask back on. Statements can be filmed outside if weather permits.

☐ 	Interviewees self-record interview statements outside of the filming date. JoVE can provide support for this option.

☐ 	Interview Statements are read by JoVE’s voiceover talent.

☐ 	Author interview statement opt out (interview statements removed completely from paper).

4. Filming location: Will the filming need to take place in multiple locations (greater than walking distance)?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

- To ensure that your script can be filmed in one day, the Introduction, Protocol, and Conclusion sections are restricted to 55 shots (i.e., designated as 1.1.1., 1.2.1., etc).

Protocol Length

Number of Shots: 23

# Introduction

- Introductory Interview Statements

Your answers to these questions will become author interview statements, which will be incorporated into the script. Authors will memorize the statements and then deliver them on camera.

- Enter the full name of the author who will deliver the statement on camera.

- Each author should deliver no more than two statements.

- Fill out both required statements. One optional statement may also be selected.

- Please answer in full sentences in a style suitable for being spoken aloud.

- Limit the length of each statement to 30 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: Why is your protocol significant? OR What key questions can this method help answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

REQUIRED: What is the main advantage of this technique?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Do the implications of this technique extend toward the therapy (or diagnosis) of a particular disease, disability, or challenge? How so?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Are there any specific areas of research that this method could provide insight into? OR Can this method be applied to any other systems?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: How would you expect an individual who has never performed this technique to struggle? Do you have any advice to offer to somebody who is trying this technique for the first time?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Why is visual demonstration of this method critical?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Has any author been assigned more than two statements? ☐ No

# Introduction of Demonstrator on Camera

Complete this statement ONLY if any of the individuals who will be demonstrating the procedure on camera will not be delivering an Introductory Interview Statement.

  - Enter name of author who will introduce demonstrator: Demonstrating the procedure will be Click here to enter name of demonstrator(s)., a Click here to enter demonstrator job title. from my laboratory. Include additional demonstrators as needed.[1][2].

    - INTERVIEW: Author saying the above

    - The named demonstrator(s) looks up from workbench or desk or microscope and acknowledges the camera

Ethics Title Card

  - Procedures involving human subjects have been approved by the Institutional Review Board (IRB) or equivalent body at insert Institutional Name.

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The one-digit numbers represent sections of the video. The text will appear onscreen.

- The two-digit numbers (e.g. 2.1., 2.2.) represent steps of your protocol. The text will be recorded by a professional voiceover talent.

- The three-digit numbers (e.g. 2.1.1., 2.2.2.) represent the shots that our videographer will capture at your lab.

Please use this draft script to help you prepare for filming day.

- Filming should take no more than 10 minutes per step. If a step will take more than 10 minutes, prepare the product from that step in advance.

- Patient Interface and Y-Piece Preparation

  - To set up the patient interface, connect the built-in elbow port of a snug-fitting, standard, non-invasive ventilation face mask of the appropriate size for the Subject to a high-efficiency particulate air filter through the 22-millimeter outer diameter-15-millimeter inner diameter connector [1-TXT].

    - WIDE: Talent connecting port to filter TEXT: e.g., highly hydrophobic bacterial/viral filter, HEPA class 13

  - To facilitate the Subject's movement and reduce the risk of disconnection, add a flexible 5-6.5-centimeter 15-millimeter outer diameter by 22-millimeter outer diameter-15-milimeter inner diameter patient connector for an endotracheal or tracheostomy tube between the mask interface and the HEPA filter [1].

    - Talent adding tube between mask interface and filter

  - To build a Y-piece, use two, opposite-sense, low-resistance, 22-millimeter male-female, one-way valves to create expiratory and inspiratory limbs on the two distal ends of a 22-millimeter to 22-millimeter and 15-French Y-piece connector with 7.6-millimeter ports [1].

    - Talent adding valve(s) to limb(s)

  - Position the one-way valve connector of the expiratory limb with the arrow pointing downward to allow a proximal-to-distal flow only [1] and position the connector for the inspiratory limb with the arrow pointing upward to allow a distal-to-proximal flow only [2].

    - Connector being positioned with arrow pointing down

    - Connector being positioned with arrow pointing up

  - Then connect the proximal end of the Y-piece to the HEPA filter [1] and use standard, kink-resistant, vinyl gas tubing with universal adaptors at both ends to connect the oxygen source to the inspiratory limb [2-TXT].

    - Talent connecting Y-piece of filter

    - Talent connecting tubing to oxygen source TEXT: Select tubing based on distance between subject and gas source

- Scavenging Chamber Preparation

  - To prepare the scavenging chamber, connect a 22- x 22-millimeter, flexible, silicon rubber connector adapter to the proximal end of a scavenger chamber containing 100 grams of calcium hydroxide [1-TXT].

    - WIDE: Talent connecting adapter to chamber TEXT: chamber internal diameter: 60 mm; internal length: 53 mm; volume: 150 mL

  - Attach a flexible, 5-6.5-centimeter 15-millimeter outer diameter by 22-millimeter outer diameter-15-millimeter inner diameter corrugated tube to the silicon rubber adapter [1] and connect another 22- x 22-millimeter flexible, silicon rubber connector adapter to the distal end of the scavenger [2].

    - Talent connecting tube to adapter

    - Talent connecting tube to scavenger

  - Then use a 15-22-millimeter two-step adapter to attach the scavenging chamber and tubing assembly to the inspiratory limb of the Y-piece [1].

    - Talent connecting chamber and Y-piece

- Nitric Oxide (NO) Reservoir and Medical Air Supply System Preparation

  - To prepare the nitric oxide reservoir system, connect a 3-liter, latex-free breathing reservoir bag and a 90-degree ventilator elbow connector without ports [1] and connect the other end of the elbow to the central opening of the aerosol T-piece [2].

    - WIDE: Talent connecting bag and connector

    - Talent connecting elbow to T-piece

  - Advance the T-piece into the distal end of the scavenging chamber until it tightly fits the silicon rubber connector [1] and attach two, consecutive 15-millimeter outer diameter by 22-millimeter outer diameter-15-millimeter inner diameter connectors with 7.6-millimeter sampling ports and flip-top caps [2].

    - Piece being advanced

    - Talent attaching connector(s)

  - Attach another one-way inspiratory valve with the arrow pointing up at the distal end of the nitric oxide-air supply system [1] and connect a 15-22-millimeter two-step adapter to the proximal end of the system [2].

    - Talent attaching valve

    - Talent connecting adapter

  - At the proximal end of the nitric oxide-air supply system, connect a 15-22-millimeter two-step adapter [1] and connect the proximal two-step adapter to the remaining free inlet of the green T-piece from the nitric oxide reservoir system [2].

    - Talent connecting adapter

    - Talent connect adapter

  - To attach the air and nitric oxide gas flow lines, use standard, kink-resistant, star-lumen vinyl oxygen gas tubing to connect the medical air flow to the most distal gas inlet port [1] and connect the nitric oxide gas flow from an 800-parts per million medical-grade nitric oxide tank to the next port downstream [2].

    - Talent connecting air to inlet port

    - Talent connecting tank to downstream port

- NO Delivery in Spontaneously Breathing Subjects

  - To administer nitric oxide to a spontaneously breathing Subject, set the air, oxygen, and nitric oxide gas flow according to the Table [1] and position the tight-fitting mask onto the Subject’s face, similar to a non-invasive ventilation interface setup [2].

    - LAB MEDIA: Table 1

    - WIDE: Talent positioning mask onto Subject’s face

  - Then start the inhalation session for the desired duration [1]. The nitric oxygen and nitrogen dioxide concentrations can be traced to assess the stability of the gas delivery during the treatment [2].

    - Talent starting inhalation session

    - LAB MEDIA: Figure 2

## Protocol Script Questions

# Authors: Please use the step numbers from the script above (not step numbers from the manuscript) when answering the questions below. Please do not include steps that will be screen-captured and do not list entire sections.

A. Which steps from the protocol are the most important for viewers to see? Please list 4 to 6 individual steps.

Click here to list 4 to 6 individual steps, using the step numbers from the protocol section of the video script.

B. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1 or 2 individual steps from the script above.

Click here to list 1 or 2 individual steps, using the step numbers from the protocol section of the video script.

# Results

Please review this section to make sure that it accurately reflects your findings.

- Use Track Changes when making edits or revisions.

- If you would like the video to include different results, please revise this section.

- When revising, please keep the length of the voiceover below 200 words. Current word count: 88. (Voiceover is the text that follows the two-digit numbers.)

- Please note that the video cannot include voiceover without an accompanying visual.

- Notes in blue italics are for JoVE’s video editor.

- Results: Representative NO and NO2 Concentration Tracings During a 160 ppm NO Inhalation in a Healthy Healthcare Worker

  - A 33-year-old respiratory therapist working at the ICU at Mass General Hospital during the surge of ICU admission for COVID-19 volunteered to receive nitric oxide as part of the trial involving healthcare workers [1].

    - LAB MEDIA: Figure 2

  - The resulting nitric oxide concentration was 160 parts per million at a total gas flow rate of 19.5 liters/minute as measured by three standard 15 liter/minute flowmeters [1] and remained stable throughout the entire period of inhalation [2].

    - LAB MEDIA: Figure 2 Video Editor: please emphasize red data line

    - LAB MEDIA: Figure 2

  - Nitrogen dioxide peaked at 0.77 parts per million and was therefore safely below the recommended toxicity threshold [1].

    - LAB MEDIA: Figure 2 Video Editor: please emphasize black data line

# Conclusion

- Conclusion Interview Statements

Below are prompts for interview statements that can be used to further emphasize the significance of your protocol.

- Please answer one or two of the statements.

- Each statement is limited to 30 words.

- Answer the questions in full sentences; you will need to memorize and deliver the sentences as spoken interview statements during filming.

- Indicate the full name of the author who will deliver the statement.

# What is the most important thing to remember when attempting this procedure? Please indicate the steps (e.g., 2.4., 2.5.) in the Protocol section of the script that this advice applies to.

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera (Enter step numbers referred to.)

# Following this procedure, what other methods can be performed? What questions would these additional methods answer?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

# After its development, did this technique pave the way for researchers to explore new questions within a specific scientific field? If so, how?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Thank you for addressing our questions. We will incorporate your answers and suggestions and send you the final script before your filming day. You will also receive detailed preparation instructions in the email accompanying the final script.