Dear Editor,

Please find our revised manuscript with substantial improvements according to yours and reviewers’ comments. As you have agreed, we removed the part of the protocol describing the construction of svFCS setup from the main manuscript and placed it as a supplementary material 1.

Below you will find our answers to your and reviewers’ specific comments and suggestions.

Editorial comments:

Changes to be made by the Author(s):

1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues. The JoVE editor will not copy-edit your manuscript and any errors in the submitted revision may be present in the published version.

The manuscript was proofread, and we believe that we corrected all spelling and grammar mistakes.

2. Please format the manuscript as: paragraph Indentation: 0 for both left and right and special: none, Line spacings: single. Please include a single line space between each step, substep and note in the protocol section. Please use Calibri 12 points.

It has been done.

3. Please make the title concise.

We want to keep the title in the present form since it contains all the necessary information and illustrates the best the topic of our article. To make the title more concise, we could use the abbreviation of the method. However, we try to avoid writing svFCS since we think that the technique's full name is more informative for a broader audience.

4. Please provide at least 6 keywords or phrases.

It is the case.

5. Please ensure that the long Abstract is within 150-300-word limit and clearly states the goal of the protocol.

The Abstract has been adjusted.

6. JoVE cannot publish manuscripts containing commercial language. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials and Reagents.
For example: SI254, Thorlabs; SM1D12D, Thorlabs; 155411PK, Thermo Fisher; LabVIEW 2017 (National Instruments, Texas); (PCI-6602, National Instruments; Axiovert 200M (Carl Zeiss, Germany; Physik Instrumente (PI, Germany); Zeiss website; Uniblitz shutter (Vincent Associates, NY; Lab-Tek chambers (155411PK Thermo 432 Fisher); PolyJet Complexes; Gibco; Sigma-Aldrich; ROTI Nanoquant working solution (Roth); GloMax Discover Microplate Reader (Promega); etc.

It has been corrected.

7. Please ensure that all text in the protocol section is written in the imperative tense as if telling someone how to do the technique (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “Note.”

It has been corrected.

8. The Protocol should contain only action items that direct the reader to do something.

Yes, it does.

9. Please ensure you answer the “how” question, i.e., how is the step performed?

Yes, we did.

10. The Protocol should be made up almost entirely of discrete steps without large paragraphs of text between sections. Please simplify the Protocol so that individual steps contain only 2-3 actions per step. e.g., Step 2.

It has been corrected.

11. Please revise the protocol text to avoid the use of any personal pronouns in the protocol (e.g., "we", "you", "our" etc.).

It has been done.

12. Lines 288-294: We cannot have non numbered paragraphs in the protocol section. Please convert to numbered action step or a note under section 4.

It has been corrected.

13. In the JoVE Protocol format, “Notes” should be concise and used sparingly. They should only be used to provide extraneous details, optional steps, or recommendations that are not critical to a step. Any text that provides details about how to perform a particular step should either be included in the step itself or added as a sub-step.

Yes, we adapted to this format.

14. There is a 10-page limit for the Protocol (including single line space between each step), but there is a 2.75-page limit for filmable content. Please highlight 2.75 pages or less of the Protocol (including headings and spacing) that identifies the essential steps of the protocol for the video, i.e., the steps that should be visualized to tell the most cohesive story of the Protocol.

The limits are now respected.

15. Please obtain explicit copyright permission to reuse any figures from a previous publication. Explicit permission can be expressed in the form of a letter from the editor or a link to the editorial policy that allows re-prints. Please upload this information as a .doc or .docx file to your Editorial Manager account. The Figure must be cited appropriately in the Figure Legend, i.e. “This figure has been modified from [citation].”

According to the information on the Annual Reviews webpage (https://www.annualreviews.org/page/about/copyright-and-permissions), for the Annual Reviews Authors (dr D. Marguet in this case): There is no need to obtain permission from Annual Reviews for the use of your own work(s). Our copyright transfer agreement provides you with all the necessary permissions. Our copyright transfer agreement provides: “…The nonexclusive right to use, reproduce, distribute, perform, update, create derivatives, and make copies of the work (electronically or in print) in connection with the author’s teaching, conference presentations, lectures, and publications, provided proper attribution is given…”

16. Each Figure Legend should include a title and a short description of the data presented in the Figure and relevant symbols.

It has been corrected.

17. As we are a methods journal, please ensure that the Discussion explicitly cover the following in detail in 3-6 paragraphs with citations:

a) Critical steps within the protocol

b) Any modifications and troubleshooting of the technique

c) Any limitations of the technique

d) The significance with respect to existing methods

e) Any future applications of the technique

Additional improvements were made in the Discussion section to consider the aforementioned points.

18. Please do not abbreviate the journal titles in the references section.

It has been corrected.

19. Please sort the materials table in alphabetical order.

It has been done.

20. Table 1: Please upload in .xlsx format separately to your Editorial Manager account.

Yes.

Reviewer #1:

Manuscript Summary:

The article by Mailfert et al. describes the construction and usage of a confocal fluorescence correlation spectroscopy (FCS) setup with variable detection area to perform spot variation FCS. The setup is clearly described and should make it possible for the reader to construct such a system and operate it. I have some remarks, clarification of which might help reader.

Major Concerns:

One page 5 the authors write "The optimal is reached when the autocorrelation function shows the higher intensity count with the lower number of molecules present in the excitation volume (meaning the higher G0)."

Essentially the authors are saying the want to maximize the number of photon counts per particle and second. I would rewrite the phrase and say "The optimum is reached when one obtains the maximum counts per particle and second, i.e. a maximum correlation function amplitude with high intensity." One also could point out that at the same time the diffusion time should be smallest.

We agree with the reviewer comment. The phrase has been changed accordingly (page 9 of the supplementary material 1).

Point 2.1.3: The authors mention that the reader should use a tube lens free microscope. But they also need to use lens L2 to create an intermediate image plane for the pinhole. Can one not use the microscope image plane after the tube lens to place the pinhole?

We thank the reviewer for this interesting remark. One can use the microscope image plane after the tube lens to place the pinhole. Two possibilities are then available: (1) a first requires separating the excitation and detection beams by using two microscope ports (e.g., rear and left port) or (2) if one wants to use a tube lens configuration, another possibility is to add an extra lens in the excitation path to keep the beam collimated and the back-pupil objective size. A paragraph has been added accordingly (page 7 of the supplementary material 1).

Some of the installation instructions are rather restrictive and could be loosened to give the protocol a wider range of applicability. E.g. under point 4 the authors ask for installations of 32 bit software on the basis that the microscope is controlled by 32 bit libraries. But this is not a state to exist for long and one wonders if it wouldn't be better to write this part more generally and that the user needs to adapt his software to the microscope with some details on what to watch out for?

We agree. The restriction about using a 32-bit version of Labview comes only from the fact that we use an old microscope driven by old 32-bit libraries. It was stated at page 1 of the supplementary material.

Under 10.6 the authors mention a software written for IGOR Pro. Is this software available to the readers? If yes, any restrictions on the version (32 bit or 63 but, 6,7,8 etc.)? If not, can it be any self-written software? In addition, one could point to already freely available software.

The fitting software was developed with IGOR Pro 6.37 32-bit and is available upon request through an MTA (Material Transfer Agreement). The 64-bit is not recommended by Wavemetrics as we are not working with extremely large data sets. The newer version of IGOR Pro (7 and 8) are also compatible with our software but request minor code modifications.

As the file provided by the correlator is in an easily readable format (text), the user can use any free software to load, discard, average fit, and analyse data.

Minor Concerns:

Line 50: replace "At first" with "First"

Line 74: replace "principal" with "important" or a synonym

Line 98: remove "a" before "a rapidly developing …"

Line 121: replace "… scratch custom …" with "… scratch a custom …"

Line 137: replace "optical path" with "optical axis"

Line 271: replace "eliminates" with "eliminate"

Line 385: remove the redundant "to the computer"

Line 487: delete "one of" in "at one of the maximal"

Line 566 onwards: Decimal points in numbers should be "." not ","

All have been corrected.

Reviewer #2:

Manuscript Summary:

The manuscript provides a useful protocol to assemble and employ a custom-built optical set up for spot variation FCS (svFCS) that is underpinned by a continuous wave laser, motorised inverted microscope, high NA water objective, an avalanche photodiode detector and a correlator, for live cell membrane studies. The microscopy set up is well described, however, the svFCS data acquisition and ACF analysis require more explanation. In particular, the details surrounding: (1) how the recorded fluctuations in fluorescence intensity are temporally segmented, (2) how the resulting autocorrelation functions (ACFs) are selected for averaging, and (3) how the ACF shape (diffusion time) changes as a function of spot radius; all need to be described and preferably visualised in a dedicated figure if the users of this protocol are going to be able to establish a diffusion law for molecules in the plasma membrane.

We agree with this comment, and we took into consideration the remarks when preparing the revised version of this manuscript. However, all the data acquisition steps were described in detail in our previous publication (Mailfert et al., 2017) together with appropriate figures, and we would not like to copy it in this manuscript to avoid redundancy.

In the Protocol under spot size calibration 10.5 and svFCS data acquisition 11.6 more detail should be provided on what is meant by '10 runs lasting for 20 seconds each' and '20 runs lasting for 5 seconds each'?

It has been addressed in point 8 of the revised Protocol.

For example what is the sampling frequency of these acquisitions (to emphasise what temporal resolution is required for a fast freely diffusing molecule like Rho6G versus a membrane protein like Thy-1-eGFP) and how many points are acquired in a trace (to emphasise the statistics required and the manner in which temporally you cut the data for the ACF analysis)?

The sampling frequency is determined by the hardware correlator, which provides a minimum sample time of 12.5ns (i.e., a sampling frequency of 80MHz), which is at least one thousand smaller than the resident time of freely diffusing Rho6G molecules (few tens of µs) and one million smaller than the diffusion of membrane protein within a 200 nm beam radius (discussed in the manuscript, section 3).

Also, what is the different between a run and a recording? If the run lasts for 5 seconds, how many points in time are acquired in this run? Is the 'recording' the acquisition time and then the 5 second 'runs' are the temporal segments used for the correlation analysis? If this is the case it would be good to provide this detail more accurately, and more importantly, provide commentary on why these parameters were selected (e.g. why 10 runs for 20 s for Rho6G and 20 runs for 5 s for Thy1-eGFP?).

We call a run a single measurement, providing a single ACF and recording a set of multiple runs. According to the multiple-tau correlation algorithm, the ACF provided by the real-time hardware correlator is composed of more than one thousand delay times logarithmically spaced between 12.5 ns to few seconds. It has been addressed in points 7 and 8 of the revised Protocol.

2. The information provided in 11.8 (four waste sizes varying between 200-400 nm) should be provided earlier in the protocol under spot size calibration 10.8 / svFCS data acquisition 11.7 and then further elaborated upon in terms of: (1) why were these parameters selected and (2) is there any point in going lower or higher than these waste values? In section 10.8 and 11.7 you are left wondering what sizes you should actually modulate the beam waste between and at what step size - for example what are the four wastes sizes between 200-400 nm that were employed and was this only because of optical limitations?

The waist size ranges from 200 to 400 nm only because of optical limitations (diffraction). By using different microscope objectives with different numerical apertures, we can modulate this range slightly.

3. In the Protocol under svFCS data analysis 12.1 it is stated that the users should 'select the individual autocorrelation functions characterised by the appropriate curve shape'. What is considered an 'appropriate' curve shape? Presumably, this point is meant to highlight that autocorrelation functions suffering from photobleaching or z drift will give rise to a slow a decay that extends out to 10 s and autocorrelation functions that don't detect the molecules will give rise to a correlation function that sharply decreases and potentially only contains correlation due to instrumentation effects. In either case what is 'appropriate' should be described more technically for users unfamiliar with ACF analysis and potentially there should be a figure dedicated to this since no ACF is shown in the protocol.

As stated previously, this critical step has been addressed in our previous publication (Mailfert et al., 2017). In the manuscript, we wrote the significant information according to the reviewer’s comment (7.5, 7.11, 7.14, 8.9).

4. Given the points raised above (#2 and #3) it would be useful to generate a figure that describes the svFCS data acquisition and analysis. In particular present an example fluorescence fluctuation that was recorded, how this data is temporally segmented for ACF analysis, example ACF curves (appropriate and not appropriate), the average ACF curve fit to a 2D later diffusion model and extraction of the various parameters for establishing the diffusion law. In this figure it would also be particularly powerful to provide the average ACF curve derived as a function of spot waste so that users can see visually the impact this change in observation volume has on diffusion time.

As mentioned before, detailed figures describing these analyses were published within our previous publication (Mailfert et al., 2017). We would not like to copy it in this manuscript because of space limitations and to avoid redundancy.

5. In the protocol under svFCS data analysis there is no mention of how the ACF function amplitudes (G(0)) can be useful for determining the number or particles present in a selected single point measurement and how this information could be used in the context of svFCS to generate a concentration gradient. Potentially this avenue should be explored / presented or at least mentioned in the protocol text.

We added a comment in the manuscript's introduction to account for the comment about concentration measurement and in section 7.5.

Minor Concerns:

1. In the Protocol under svFCS data acquisition 11.5 it is unclear what is meant by the instructions on how to position the observation volume for the svFCS data measurement in the membrane of a selected cell. Why would it be preferable to run the scan in the nuclear area of the cell when the introduction of the protocol suggests the svFCS application here is for membrane-based studies? Potentially does this point mean to say the user should focus on the middle plane of the nucleus and then bring the focus down to where the nuclear envelope is no longer visible and Thy-1 eGFP extends to the perimeter of the cell.

It has been addressed in the revised Protocol, point 8.8.

2. In the Introduction it would be good to also mention spatiotemporal correlation spectroscopy (e.g. RICS, STICS, the pair correlation function, iMSD) alongside FRAP and FCS - as this family of methods also probe membrane diffusion with a spatial component introduced into the correlation function and like spot variant FCS (svFCS) that does this experimentally, reveal information concerning the underlying structure of the membrane. Then the high temporal resolution / statistics of svFCS compared to these spatiotemporal correlation-based approaches could be highlighted.

We agree that maybe it would be nice to introduce other methods. Still, we consider that this publication focuses on svFCS and its practical implementation. Taking into account the manuscript lenght, we did not introduced different methods in detail. In our opinion, there is no particular need to do so. However, we have mentioned some of the methods in the Discussion section.

3. In the introduction it should also be mentioned that the amplitude of the autocorrelation function can provide the number of particles present in the observation volume and how this parameter in the context of svFCS could provide information on the concentration gradient at any position interrogated.

It has been mentioned in the revised manuscript.

4. In the introduction at the top of page 2 the fluorescence signal is not 'induced' by diffusing fluorescent molecules but rather generated.

It has been corrected, thank you.