Rebuttal letter

JoVE61890R1 - “Single-cell optical action potential measurement in human induced pluripotent stem cell-derived cardiomyocytes”

Editorial Reviewer Comments:

We thank the editors for giving such precise instructions for shaping the manuscript into a text that is appropriate for the Journal of Visualized Experiments. A point-by-point response is detailed below.

1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues. The JoVE editor will not copy-edit your manuscript and any errors in the submitted revision may be present in the published version. Please use American English throughout.

The text has been thoroughly checked by all authors and all ‘british-isms’ changed to American English.

2. Please format the manuscript as: paragraph Indentation: 0 for both left and right and special: none, Line spacings: single. Please include a single line space between each step, substep and note in the protocol section. Please use Calibri 12 points

Formatting has been followed as instructed.

3. As one of the authors is from UK please agree to UK ALA in the additional details page of the submission site.

This has been agreed to on the submission site.

4.Please provide at least 6 keywords or phrases.

5. We cannot have commercial terms in the keyword. Please remove FluoVolt and use generic term instead.

Keyword number has been updated to 6 and no longer contains commercial language.

6. JoVE cannot publish manuscripts containing commercial language. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials and Reagents.
For example: FluoVolt, Invitrogen, TrypLE, Accutase, falcon, Axon Instruments Digidata, Eppendorf, PowerLoad, etc.

7. JoVE policy states that the video narrative is objective and not biased towards a particular product featured in the video. The goal of this policy is to focus on the science rather than to present a technique as an advertisement for a specific item. To this end we ask you remove the term FluoVolt and use generic term instead.

Commercial language in the protocol has been removed. FluoVolt has been replaced with the generic term VF2.1Cl (VoltageFluor2.1Cl).

8. Please ensure the Introduction include all of the following with citation:
a) A clear statement of the overall goal of this method
b) The rationale behind the development and/or use of this technique
c) The advantages over alternative techniques with applicable references to previous studies
d) A description of the context of the technique in the wider body of literature
e) Information to help readers to determine whether the method is appropriate for their application

We have made major revisions to the introduction which now covers each of these points.

9. Please ensure that all text in the protocol section is written in the imperative tense as if telling someone how to do the technique (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “Note.” However, notes should be concise and used sparingly.

Text has been changed to the imperative tense throughout the protocol. Notes have been used where appropriate.

10. The Protocol should be made up almost entirely of discrete steps without large paragraphs of text between sections. Please simplify the Protocol so that individual steps contain only 2-3 actions per step.

11. We cannot have non numbered paragraph as well. Please convert these to notes instead.

Each step now contains a maximum of three actions and non-numbered paragraphs have been removed.

12. The Protocol should contain only action items that direct the reader to do something.

13. Please ensure you answer the “how” question, i.e., how is the step performed?

We have taken care to ensure that the protocol uses direct language and details how each step is to be performed.

14. There is a 10-page limit for the Protocol, but there is a 3-page limit for filmable content. Please highlight 3 pages or less of the Protocol (including headings and spacing) that identifies the essential steps of the protocol for the video, i.e., the steps that should be visualized to tell the most cohesive story of the Protocol.

15. Please include all relevant details that are required to perform the step in the highlighting. For example: If step 2.5 is highlighted for filming and the details of how to perform the step are given in steps 2.5.1 and 2.5.2, then the sub-steps where the details are provided must be highlighted.

Appropriate and filmable protocol content has now been highlighted. It is well within a 3 page limit.

16. Please obtain explicit copyright permission to reuse any figures from a previous publication. Explicit permission can be expressed in the form of a letter from the editor or a link to the editorial policy that allows re-prints. Please upload this information as a .doc or .docx file to your Editorial Manager account. The Figure must be cited appropriately in the Figure Legend, i.e. “This figure has been modified from [citation].”

As our figure is entirely self-made and every graphical element is completely original, we have removed this citation.

17. As we are a methods journal, please ensure that the Discussion explicitly cover the following in detail in 3-6 paragraphs with citations: a) Critical steps within the protocol

b) Any modifications and troubleshooting of the technique

c) Any limitations of the technique

d) The significance with respect to existing methods

e) Any future applications of the technique

We have made major revisions to the discussion it now satisfactorily covers all of these points.

18. Please sort the materials table in alphabetical order.

The materials have been sorted into alphabetical order in their respective groups.

19. Please check if the attached supplementary file is correct. If this needs filming please include this with the actual protocol and highlight accordingly.

The supplemental file is correct and does not require filming. We feel it is still necessary to include in order to make clear to the readers that many cellular and animal constructs can be assessed with this multi-purpose protocol.

Reviewer 1 Comments:

This manuscript provides a detailed protocol to acquire optical recordings of action potentials in hiPSC-derived cardiomyocytes using a commercially available voltage-sensitive dye, FluoVolt. This is a method that is gaining interest for researchers studying hiPSC-cardiomyocytes due to its high throughput. The method is also technically less challenging than the traditional method of measurement using the patch-clamp technique. Hence, the publication is timely and of value to the readers. There are no major concerns.

We thank this reviewer for his/her positive evaluation of our manuscript.

Minor Concerns: The reason for the addition of probenecid should be included in the text.

We are incredibly grateful to this reviewer for pointing out this sloppy mistake. Probenecid is used in other protocols we employ for imaging of cytosolic calcium and ensuring the indicator is not extruded from the cell. This has been removed from the manuscript. The Tyrode’s solution constituents have been double checked and now are listed correctly.

Reviewer 2 Comments:

We thank this reviewer for his/her detailed evaluation of our protocol and their incredibly helpful comments which have guided us to substantially improve our revised manuscript. We have made major adjustments, particularly to the discussion, based on this reviewers suggestions. A point-by-point response is detailed below.

The authors describe a well built system for optical mapping of single cells. The VSD chosen has been in the market for a few years and has all the qualities described in the paper. The use of the fluovolt has been well described and results documented in a fairly large number of publications. The main contribution of this manuscript is the description of how to assemble a PMT-based optical mapping system. Although Cairn contributed financially and provided materials for assembly of the system, a PMT-based optical mapping system can theoretically be assembled with parts from other brands. It would be desirable to have that clearly stated in the manuscript, otherwise this would be a simple assembly manual for a Cairn system (which by the way manufactures excellent and high quality equipment).

Firstly, we agree and wholly support Cairn Products! Secondly we thank the reviewer for highlighting this point. To our knowledge, Cairn is only distributor that provides a modular system which does not directly integrate the PMT into their acquisition hardware. Respectively, this allows for simultaneous measurements of multiple parameters and gives the operator more control over analogue signal processing prior to digitization. We have incorporated these points into our discussion and reduced our commercial plugs for Cairn equipment.

In addition we have added this sentence to our discussion to make this clear for the readers:

“Many optics-based setups are commercially available and can be optimized for a wide range of signal recording requirements”

Furthermore, it is imperative to indicate the limitations of this system in regards to the loss of spatial resolution. It is particularly important because the CiPA and JICSA efforts to validate the use of hiPSC-CMs for cardiotoxicity screening was performed on syncytia of hiPSC-CMs. Being direct to the point, please discuss the inability to assess conduction velocity (an important contributor to the establishment of arrhythmia) with a PMT.

We agree that this is a valuable point. We have indeed directly included this limitation into the discussion and devoted a full paragraph to discussing the advantages and disadvantages of alternative imaging systems and their implications.

We have included the following paragraph into our discussion:

“Whilst the temporal resolution of our system brings many advantages, it should be noted that parameters requiring analysis of the spatial distribution of fluorescence signals, such as excitation patterns or conduction velocity, are not suitable for investigation by the photometry techniques that we used here and are in the realms of optical mapping. The hardware described here can be readily converted to an optical mapping configuration by exchange of the photomultiplier for a suitable specialist camera with high frame rates and large well capacity. Our described configuration and protocol can offer extremely detailed temporal information at a fraction (up to 1/10th) of the commercial price of advanced large area cameras with equivalent capabilities. Confocal line scan techniques are also more spatially detailed than our method, however limitations in z restrict fluorescence acquisition from multiple transverse planes at once. This is not ideal for imaging membrane bound reporters held by iPSC-CMs with typically heterogeneous morphologies. Photometry measurements using a standard epifluorescence microscope avoid this issue by accessing the entire cellular surface instantaneously, again at a much lower cost per data point.”

In addition, we have discussed this method in the context of proarrhythmia screening throughout the manuscript. This is best highlighted in the introduction:

“A recent study has demonstrated the suitability of ratiometric voltage sensitive probe photometry to accurately quantify the cardiac action potential (Hortigon-Vinagre et al., 2016). Furthermore the ability to readily scale up optical photometry approaches lends this technique to large scale cardiotoxicity screens critical in therapeutic drug development (e.g. CiPA). Development of standardized cardiotoxicity protocols in a blinded multi-site study using microelectrode array and voltage-sensing optical techniques has demonstrated the key value of this approach (Blinova et al. 2018).”

Additionally, it would increase the strength of this manuscript to indicate that the over-killing acquisition rate might not be required for obtaining precise measurements of upstroke velocity, but would be incredibly useful for measuring APD in other excitable cells that have incredibly short APDs, like neurons.

We thank the reviewer for this useful comment which prompted us to fully investigate and review our digitization procedure. We agree that we are oversampling. Reliable analysis of upstroke gradient will require 5 or more data points. On the assumption that a typical cardiac AP rise time is 10 ms, an analogue signal cut off at 500 Hz is appropriate. We double this to 1 kHz, both to ensure a detailed recording, and because our PMT amplifier only allows for filtering at exponents of 10 (1, 10, 100 etc..). To satisfy Nyquist requirements and prevent aliasing we would therefore only need a digitization rate of 2 – 3 kHz.

In the protocol itself, we have directed attention to these steps, and we have added the following sentence to the discussion:

“High sampling rates are required to ensure reliable measurements of action potential upstroke in cardiomyocytes and can be critical in other excitable constructs (i.e. neurons) with extremely fast excitation kinetics”

It would be interesting to discuss that the system described may be suitable for measuring duration of intra-cellular calcium release and perhaps amplitude or relative amplitude.

We completely agree with this comment. We have now added a section directly discussing this systems ability to provide readouts of intracellular calcium. This is helpful on multiple levels, because describing the multi-functionality of such a system of course increases its desirability.

The following sentences have been added to the discussion:

“Addition of a second photomultiplier channel, with appropriate filter sets to avoid spectral overlap, will enable true simultaneous measurement of the fluorescent intracellular calcium indicator CAL-590 alongside VF2.1Cl. It is also important to note that this multipurpose system can be used alternatively for deep assessment of intracellular calcium handling as previously described (Voigt et al., 2013, Voigt et al., 2012, Voigt et al., 2014, Fakuade et al., 2020).”

Mapping is yet far from being able to measure specific currents. Nevertheless, it still can significantly reduce the amount of patch clamp work if it is used for assessment of APD morphology in order to do educated decisions on the ionic currents that should be assessed in following experiments.

We would like to clarify that we are only comparing our method to traditional current-clamp protocols. We thoroughly agree with the reviewer that our method will never replace voltage-clamp based assessment of individual ionic currents.

To address this, we have clarified this comparison with current-clamp in the introduction and included this brief reminder in the discussion:

“This method allows for rapid and precise quantification of repolarization mechanics, which can provide valuable insights into cellular ionic abnormalities.”