# Submission ID #: 61907

# Scriptwriter Name: Bridget Colvin

Project Page Link: https://www.jove.com/account/file-uploader?src=18881683

Title: Magnetic Field Mapping Using Off-Axis Electron Holography in the Transmission Electron Microscope

Authors and Affiliations: Fengshan Zheng1, András Kovács1, Thibaud Denneulin1, Jan Caron1, Teresa Wessels1, and Rafal E. Dunin-Borkowski1

1Ernst Ruska-Centre for Microscopy and Spectroscopy with Electrons and Peter Grünberg Institute

☐   All author names are spelled correctly, and the affiliations are correct (city/state/country information not included in video title page).

# Corresponding Author:

# Fengshan Zheng

# f.zheng@fz-juelich.de

# András Kovács

# a.kovacs@fz-juelich.de

# Co-Authors:

t.denneulin@fz-juelich.de

j.caron@fz-juelich.de

t.wessels@fz-juelich.de

r.dunin-borkowski@fz-juelich.de

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or similar?  N

2. Software: Does the part of your protocol being filmed demonstrate software usage?  Y

If Yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps. Please upload all screen captured video files to your project page as soon as reasonably possible.

3. Interview statements: Considering the Covid-19-imposed mask-wearing and social distancing recommendations, which interview statement filming option is the most appropriate for your iron germanium group? Please select one.

☐ 	Interviewees wear masks until the videographer steps away (≥6 ft/2 m) and begins filming. The interviewee then removes the mask for line delivery only. When the shot is acquired, the interviewee puts the mask back on. Statements can be filmed outside if weather permits.

☐ 	Interviewees self-record interview statements outside of the filming date. JoVE can provide support for this option.

☐ 	Interview Statements are read by JoVE’s voiceover talent.

☐ 	Author interview statement opt out (interview statements removed completely from paper).

4. Filming location: Will the filming need to take place in multiple locations (greater than walking distance)?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

- To ensure that your script can be filmed in one day, the Introduction, Protocol, and Conclusion sections are restricted to 55 shots (i.e., designated as 1.1.1., 1.2.1., etc).

Protocol Length

Number of Shots: 34

# Introduction

- Introductory Interview Statements

Your answers to these questions will become author interview statements, which will be incorporated into the script. Authors will memorize the statements and then deliver them on camera.

- Enter the full name of the author who will deliver the statement on camera.

- Each author should deliver no more than two statements.

- Fill out both required statements. One optional statement may also be selected.

- Please answer in full sentences in a style suitable for being spoken aloud.

- Limit the length of each statement to 30 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: Why is your protocol significant? OR What key questions can this method help answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

REQUIRED: What is the main advantage of this technique?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Do the implications of this technique extend toward the therapy (or diagnosis) of a particular disease, disability, or challenge? How so?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Are there any specific areas of research that this method could provide insight into? OR Can this method be applied to any other systems?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: How would you expect an individual who has never performed this technique to struggle? Do you have any advice to offer to somebody who is trying this technique for the first time?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Why is visual demonstration of this method critical?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Has any author been assigned more than two statements? ☐ No

# Introduction of Demonstrator on Camera

Complete this statement ONLY if any of the individuals who will be demonstrating the procedure on camera will not be delivering an Introductory Interview Statement.

  - Enter name of author who will introduce demonstrator: Demonstrating the procedure will be Click here to enter name of demonstrator(s)., a Click here to enter demonstrator job title. from my laboratory. Include additional demonstrators as needed.[1][2].

    - INTERVIEW: Author saying the above

    - The named demonstrator(s) looks up from workbench or desk or microscope and acknowledges the camera

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The one-digit numbers represent sections of the video. The text will appear onscreen.

- The two-digit numbers (e.g. 2.1., 2.2.) represent steps of your protocol. The text will be recorded by a professional voiceover talent.

- The three-digit numbers (e.g. 2.1.1., 2.2.2.) represent the shots that our videographer will capture at your lab.

Please use this draft script to help you prepare for filming day.

- Filming should take no more than 10 minutes per step. If a step will take more than 10 minutes, prepare the product from that step in advance.

- Electron Microscope Alignment

  - To set the microscope to Lorentz mode, load a dedicated alignment file [1].

    - WIDE: Talent loading file, with monitor visible in frame

  - Load the specimen into a TEM (T-E-M) specimen holder [1-TXT] and set the appropriate microscope imaging and alignment parameters [2].

    - Talent loading specimen into holder TEXT: transmission electron microscope

    - Talent setting microscope parameters

  - To correct the two-fold astigmatism of the Lorentz lens, use the detector image recording and processing software to monitor the Fourier transform of a thin, amorphous region on the specimen in real time [1-TXT].

    - SCREEN: To be provided by Authors: Thin, amorphous region being monitored TEXT: Alternative: monitor separate cross-grating or amorphous carbon thin film

  - After adjusting the magnification of the microscope for the desired field of view to ideally include a region of vacuum across at least 10% of the image [1], move the specimen away from the field of view [2].

    - Talent adjusting magnification, with monitor visible in frame

    - Talent moving specimen away from FOV

  - Insert the biprism [1] and orient the biprism with respect to the specimen [2].

    - Talent inserting the biprism

    - Talent orienting biprism

  - Apply a desired voltage to the biprism, typically at a rate of no more than 1 volt/second to avoid damage [1] and tune the condenser lens astigmatism to set up an elliptical electron beam condition [2].

    - Talent applying voltage OR SCREEN: To be provided by Authors: Voltage being applied

    - Talent tuning astigmatism OR SCREEN: To be provided by Authors: Astigmatism being tuned

  - Center the electron beam and fine-tune the gun and condenser lens astigmatism settings to maximize the holographic interference fringe contrast [1-TXT].

    - SCREEN: To be provided by Authors: Beam being centered, then settings being fine-tuned TEXT: Do not focus bean onto biprism

  - When the microscope, electron beam, biprism, and specimen have stabilized, set the magnification according to the size of the region of interest [1] and adjust the biprism voltage according to the desired overlap width and spatial resolution. Then move the region of interest into the field of view [2].

    - Talent setting/selecting magnification

    - SCREEN: To be provided by Authors: Voltage being adjusted, then ROI being moved into FOV

- Lorentz TEM (Fresnel Defocus Imaging)

  - For Loretnz defocus imaging, switch back to round beam illumination [1-TXT] and change the defocus of the Lorentz lens to record Fresnel defocus images [2].

    - WIDE: Talent switching to round beam illumination TEXT: Move biprism from FOV as necessary

    - SCREEN: To be provided by Authors: Lortentz lens being changed to record Fresnel defocus images

  - Then set the desired exposure time and use the camera control software to record in-focus, under focus, and over focus images [1].

    - SCREEN: To be provided by Authors: Exposure time being set, then in-focus, under focus, and over focus images

- TEM Field Evolution of the Specimen

  - To change the magnetic field applied to the specimen, tune either the current of the conventional microscope objective lens while remaining in Lorentz mode [1] or, if applicable, tune the current in the coils of a magnetizing TEM specimen holder [2].

    - WIDE: Talent tuning lens

    - Talent tuning current

  - Adjust the eucentric height of the specimen and the defocus [1] and check the alignment of the microscope [2].

    - Talent adjusting height and/or defocus

    - Talent checking alignment

  - Set the exposure time and record in-focus, underfocus, and overfocus images as demonstrated [1].

    - SCREEN: To be provided by Authors: Exposure time being set, then images being recorded

  - Then change the applied magnetic field, specimen tilt, and/or temperature as necessary to follow the magnetic response of the specimen and to select a suitable condition for off-axis electron holography [1].

    - SCREEN: To be provided by Authors: Field, tilt, and/or temperature being changed

- Off-Axis Electron Holography

  - For off-axis electron holography of the specimen, switch back to an elliptical beam condition [1] and bring the biprism to the center of the field of view [2].

    - WIDE: Talent switching to elliptical beam condition, with monitor visible in frame

    - SCREEN: To be provided by Authors: Biprism being centered in FOV

  - Check the alignment of the specimen, biprism, and microscope [1]. After the specimen has stabilized for 10-30 minutes, place the region of interest on the specimen within the field of view [2].

    - Talent checking alignment

    - Talent placing specimen into FOV

  - Set the desired exposure time and record single or multiple electron holograms of the specimen [1]

    - SCREEN: To be provided by Authors: Exposure time being set, then hologram(s) being recorded

  - Then translate the region of interest on the specimen from the field of view and record single or multiple reference holograms [1].

    - SCREEN: To be provided by Authors: ROI being translated and hologram(s) being recorded

- Off-Axis Electron Holography Field Evolution of the Specimen

  - Change the magnetic field applied to the specimen as demonstrated [1] and set up the elliptical beam condition according to the objective lens parameters [2].

    - WIDE: Talent turning lens

    - SCREEN: To be provided by Authors: Elliptical beam condition being set

  - Realign the specimen, biprism, and microscope [1] and bring the specimen into focus [2].

    - Talent realigning specimen, biprism, and/or microscope

    - SCREEN: To be provided by Authors: Specimen coming into focus

  - Place the region of interest within the field of view and set the time of exposure [1].

    - SCREEN: To be provided by Authors: ROI being moved into FOV, then exposure time being set

  - Then record single or multiple electron holograms of the specimen as demonstrated before translating the region of interest on the specimen from the field of view and recording a second set of single or multiple reference holograms [1-TXT].

    - SCREEN: To be provided by Authors: Hologram(s) being recorded, ROI being translated, then hologram(s) being recorded TEXT: Repeat analyses for each value of applied field and/or voltage or specimen temperature before and after turning specimen over

## Protocol Script Questions

# Authors: Please use the step numbers from the script above (not step numbers from the manuscript) when answering the questions below. Please do not include steps that will be screen-captured and do not list entire sections.

A. Which steps from the protocol are the most important for viewers to see? Please list 4 to 6 individual steps.

Click here to list 4 to 6 individual steps, using the step numbers from the protocol section of the video script.

B. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1 or 2 individual steps from the script above.

Click here to list 1 or 2 individual steps, using the step numbers from the protocol section of the video script.

# Results

Please review this section to make sure that it accurately reflects your findings.

- Use Track Changes when making edits or revisions.

- If you would like the video to include different results, please revise this section.

- When revising, please keep the length of the voiceover below 200 words. Current word count: 203. (Voiceover is the text that follows the two-digit numbers.)

- Please note that the video cannot include voiceover without an accompanying visual.

- Notes in blue italics are for JoVE’s video editor.

- Results: Representative Magnetic Field Mapping

  - Here a representative lamella attached to a copper grid can be observed [1].

    - LAB MEDIA: Figure 1C Video Editor: please add/emphasize text and arrow

  - The magnetic state of the iron germanium specimen is expected to follow the illustrated magnetic field versus temperature phase curve [1].

    - LAB MEDIA: Figure 4A Video Editor: please emphasize curved data line

  - Below the transition temperature, iron germanium typically contains a helical magnetic structure in a zero applied magnetic field [1] that produces maze-like lines of black and white contrast in Lorentz images of thin TEM specimens [2].

    - LAB MEDIA: Figure 4B

    - LAB MEDIA: Figure 4B Video Editor: please emphasize maze-like lines in FeGe image

  - Here Fresnel defocus images of Bloch-type skyrmions recorded at a specimen temperature of 100 Kelvin in the presence of an out-of-plane magnetic field of 100 millitonne can be observed [1].

    - LAB MEDIA: Figure 2

  - Cooling of the specimens in the presence of an applied magnetic field [1] results in the formation of a regular close-packed lattice of Bloch-type skyrmions as an intensity maximum or minimum, depending on the sign of the defocus [2].

    - LAB MEDIA: Figure 2 Video Editor: please emphasize Figure 2A magnified inset image

    - LAB MEDIA: Figure 2 Video Editor: please emphasize Figure 2B magnified inset image

  - A visual representation of the projected in-plane magnetic induction can be obtained by adding contours to the magnetic phase image [1] and its derivatives can also be used to generate colors that can be used to represent the direction and magnitude of the projected in-plane magnetic induction [2].

    - LAB MEDIA: Figure 5A

    - LAB MEDIA: Figure 5B

  - The magnetic phase image can be analyzed further to determine the projected in-plane magnetization in the specimen using either a model-independent or a model-based algorithm [1].

    - LAB MEDIA: Figure 5C

# Conclusion

- Conclusion Interview Statements

Below are prompts for interview statements that can be used to further emphasize the significance of your protocol.

- Please answer one or two of the statements.

- Each statement is limited to 30 words.

- Answer the questions in full sentences; you will need to memorize and deliver the sentences as spoken interview statements during filming.

- Indicate the full name of the author who will deliver the statement.

# What is the most important thing to remember when attempting this procedure? Please indicate the steps (e.g., 2.4., 2.5.) in the Protocol section of the script that this advice applies to.

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera (Enter step numbers referred to.)

# Following this procedure, what other methods can be performed? What questions would these additional methods answer?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

# After its development, did this technique pave the way for researchers to explore new questions within a specific scientific field? If so, how?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Thank you for addressing our questions. We will incorporate your answers and suggestions and send you the final script before your filming day. You will also receive detailed preparation instructions in the email accompanying the final script.