TITLE:

Maximum Isometric Tetanic Force Measurement of the Tibialis Anterior Muscle in the Rat

AUTHOR:
Meiwand Bedar1, 2, Tiam M. Saffari1, 2, Patricia F. Friedrich1, Guilherme Giusti1, Allen T. Bishop1, Alexander Y. Shin1

AFFILIATIONS:

- Department of Orthopedic Surgery, Mayo Clinic, Rochester, MN, USA

- Department of Plastic-, Reconstructive- and Hand Surgery, Radboud University Medical Center, Nijmegen, the Netherlands

Study performed at: Microvascular Research Laboratory, Department of Orthopedic Surgery Mayo Clinic, Rochester, Minnesota, USA

E-mail addresses of co-authors:

Meiwand Bedar 		(bedar.meiwand@mayo.edu)

Tiam M. Saffari 		(saffari.tiam@mayo.edu)

Patricia F. Friedrich		(friedrich.patricia@mayo.edu)

Guilherme Giusti 		(giusti.guilherme@mayo.edu)

Allen T. Bishop			(bishop.allen@mayo.edu)

Corresponding author:

Alexander Y. Shin

shin.alexander@mayo.edu

KEYWORDS:

Nerve injury, nerve regeneration, sciatic nerve, functional recovery, motor function, tetanic muscle force, rat model.

SUMMARY:

Evaluation of motor recovery remains the benchmark outcome measure in experimental peripheral nerve studies. The isometric tetanic force measurement of the tibialis anterior muscle in the rat is an invaluable tool to assess functional outcomes after reconstruction of sciatic nerve defects. The methods and nuances are detailed in this article.

ABSTRACT:

Traumatic nerve injuries result in substantial functional loss and segmental nerve defects often necessitate the use of autologous interposition nerve grafts. Due to their limited availability and associated donor side morbidity, many studies in the field of nerve regeneration focus on alternative techniques to bridge a segmental nerve gap. In order to investigate the outcomes of surgical or pharmacological experimental treatment options, the rat sciatic nerve model is often used as a bioassay. There are a variety of outcome measurements used in rat models to determine the extent of nerve regeneration. The maximum output force of the target muscle remains the most relevant outcome for clinical translation of experimental therapies. Isometric force measurement of tetanic muscle contraction has previously been described as a reproducible and valid technique for evaluating motor recovery after nerve injury or repair in both rat and rabbit models. In this video, we will provide a step-by-step instruction of this invaluable procedure for assessment of functional recovery of the tibialis anterior muscle in a rat sciatic nerve defect model using optimized parameters. We will describe the necessary pre-surgical preparations in addition to the surgical approach and dissection of the common peroneal nerve and tibialis anterior muscle tendon. The isometric tetanic force measurement technique will be detailed. Determining the optimal muscle length and stimulus pulse frequency is explained and measuring the maximum tetanic muscle contraction is demonstrated.

INTRODUCTION:

Loss of motor function following traumatic peripheral nerve injury has a significant impact on the quality of life and socioeconomic status of patients1-3. The prognosis of this patient population remains poor due to minimal improvements in surgical techniques over the years4. Direct end-to-end tension-free epineural repair forms the gold standard surgical reconstruction. However, in cases with extended nerve gaps interposition of an autologous nerve graft has proven to be superior5,6. The associated donor site morbidity and limited availability of autologous nerve grafts have imposed the need for alternative techniques7,8.

Experimental animal models have been used to elucidate the mechanism of peripheral nerve regeneration and to evaluate outcomes of a variety of reconstructive and pharmacological treatment options8,9. The rat sciatic nerve model is the most frequently used animal model10. Their small size makes them easy to handle and house. Due to their superlative neuroregenerative potential, the diminished time between intervention and evaluation of outcomes can result in relatively lower costs11,12. Other advantages of its use include morphological similarities to human nerve fibers and the high number of comparative/historic studies13. Although the latter should be approached cautiously, as a wide variety of different outcome measures between studies makes it difficult to compare results14-18.

Outcome measures to assess nerve regeneration range from electrophysiology to histomorphometry, but these methods imply a correlation but do not necessarily directly measure the return of motor function14,15. Regenerating nerve fibers might not make appropriate connections which can cause an overestimation of the number of functional connections14,15,19,20. The best and clinically most relevant measurement to demonstrate correct reinnervation of end organs remains assessment of muscle function21-23. Creating motor function assessment tools for animal models is, however, challenging. Medinaceli et al. first described the walking track analysis, which has since been the most frequently used method to evaluate functional recovery in experimental peripheral nerve studies21,24-28. The walking track analysis quantifies the sciatic functional index (SFI) based on measurements of pawprints from walking rats21,29. Major limitations of the walking track analysis, such as toe contractures, automutilation, smearing of the print and poor correlation with other measures of reinnervation, have necessitated the use of other parameters for quantification of functional recovery30,31.

In previous studies in Lewis rats32 and New Zealand rabbits33, we validated the isometric tetanic force (ITF) measurement for the tibialis anterior (TA) muscle and demonstrated its effectiveness in the evaluation of muscle recovery after different types of nerve repair34-39. The TA muscle is well suited because of its relatively large size, innervation by the peroneal branch of the sciatic nerve and well elucidated biochemical properties40-43. When muscle length (preload force) and electrical parameters are optimized the ITF provides a side-to-side variability of 4.4% and 7.5% in rats32 and rabbits33, respectively.

This article provides a detailed protocol of the ITF measurement in the rat sciatic nerve model, including a thorough description of the necessary pre-surgical planning, surgical approach and dissection of the common peroneal nerve and the distal TA muscle tendon. Using predetermined values for the stimulus intensity and duration, the optimal muscle length and stimulus pulse frequency will be defined. With these four parameters, the ITF can subsequently be consistently and accurately measured.

PROTOCOL:

All animal procedures were performed with approval of the Institutional Animal Care and Use Committee (IACUC A334818).

- Calibration of the force transducer

  - Ensure that the computer is properly connected to the USB-6009 multifunctional I/O data acquisition (DAQ) device, which in turn should be connected to the force transducer.

NOTE: Other rat strains and species may require a different load-cell force transducer as higher forces are to be expected 44.

  - Attach a custom clamp fashioned from a modified surgical hemostat to the force transducer that is mounted to a vacuum base adjustable lever arm.

NOTE: The custom-made clamp consists of a surgical hemostat modified with a tightening screw that allows for adjustment of the tension (Figure 1).

  - Position the custom-made acrylic glass testing platform, which contains two wooden blocks for fixation of the rat hind limb, on the table.

NOTE: Other materials such as urethane can also be used instead of wood as long as the K-wires are able to penetrate and fixate.

  - Attach the clamp, force transducer and adjustable lever arm combination vertically to the testing platform using its vacuum base.

  - Fasten a hook or loop to the clamp for the calibration weights.

  - Turn on the computer and open the software (e.g., Labview).

  - Once the software is opened, start the custom-made virtual instrument (VI) for ITF measurement (Figure 2).

NOTE: Figure 2 contains the LabVIEW code in a VI snippet. This VI snippet can be dragged onto the block diagram in LabVIEW. It will automatically be transformed into a graphical code. For this experiment the sampling rate was set at 2000 Hz with 25 samples to read for each iteration.

  - Run the VI by pressing the white arrow in the left upper corner and select New calibration. A new window will open.

  - Start the calibration process with zero weight (only the clamp with an attached hook or loop) and press OK.

  - Consecutively, add 10, 20, 30 and 50 grams of weight and press OK in between each weight measurement.

  - Once all five measurements are collected, click on Process.

  - Only accept the values if the graph on the VI displays a positive linear curve (Figure 3).

  - Reposition the clamp, force transducer and adjustable lever arm combination horizontally on the testing platform. This will be the position used for measuring the ITF.

  - Click on Zero and the window will automatically close.

- Animal subjects

  - Use male Lewis rats weighing between 300-500 g.

NOTE: For comparison of nerve regeneration, it is imperative to use the same rat strain in both the control and experimental groups, since weight and incidence of autotomy are strain dependent and can tremendously influence the results of the ITF10,32,45-47.

- Surgical preparation

  - Prepare all required surgical instruments prior to surgery (Table of Materials).

  - Weigh the animals to determine the required amount of anesthesia.

  - Induce anesthesia by placing the rat in a chamber gassed with 3% isoflurane in oxygen.

  - Deeply anesthetize the rat using a cocktail of ten-parts ketamine (100 mg/mL) and one-part xylazine (100 mg/mL) at a dosage of 1 mL/kg body weight via an intraperitoneal injection. Monitor the depth of anesthesia based on the response to a toe pinch and by observing the respiratory rate.

  - Approximately 30 minutes after the initial dosage of the ketamine/xylazine cocktail, administer a supplementary dose of 0.3-0.6 mL/kg body weight of only ketamine (100 mg/mL) intraperitoneally to maintain adequate anesthesia throughout the entire procedure, which is defined as a low respiratory rate and an absent response to a toe pinch.

CAUTION: It is important to meticulously administer the required anesthesia as an overdose cannot be counteracted.

  - Carefully shave the hind limbs of the rat using electric clippers.

  - Place the rat in prone position on a heating pad to maintain the body temperature at 37 °C. Optionally, the body temperature can be monitored using a rectal thermometer.

  - Inject 5 mL of 0.9% sodium chloride (NaCl) subcutaneously into the loose skin over the neck of the rat to preserve an adequate hydration status throughout the procedure.

  - Due to the non-survival nature of this procedure, the surgical field and instruments do not require to be sterile. The surgeon should use personal protective equipment (PPE) and surgical loupes are advised for proper visualization of the anatomical structures.

- Surgical approach to the common peroneal nerve

  - Place the rat in either the right or left lateral recumbent position depending on which side will be measured first.

  - Create a 2-3 cm incision in the skin of the posterolateral thigh parallel to the femur starting at the greater trochanter using a surgical no. 15 blade.

  - Identify the plane between the biceps femoris muscle and the gluteus maximus and vastus lateralis muscles and perform a blunt dissection using tenotomy scissors to separate these muscles and expose the underlying sciatic nerve.

  - Locate the trifurcation of the sciatic nerve and place a retractor to acquire better access. The three branches of the sciatic nerve include the common peroneal nerve, the tibialis nerve and the sural nerve.

  - Isolate the common peroneal nerve branch (usually the most ventral branch) of the sciatic nerve using a curved microsurgical forceps.

NOTE: In case of uncertainty, gently stimulate the isolated nerve with a surgical nerve stimulator and observe the motor response. Stimulation of the common peroneal nerve results in dorsiflexion of the paw.

- Dissection of the distal tibialis anterior muscle tendon

  - In order to expose the TA muscle and its insertion, incise the skin at the anterolateral aspect of the lower leg, starting at the knee joint and descending to the mediodorsal side of the hind paw.

  - Dissect the distal TA muscle tendon from the surrounding tissue using a scalpel with a surgical blade no. 15.

  - Using a mosquito forceps, bluntly dissect the TA muscle tendon towards the insertion and cut the tendon as distal as possible. Leave the proximal TA muscle undisturbed, preserving the neurovascular pedicle.

NOTE: Regularly (approximately every 5 minutes), moist the TA muscle with heated 0.9% NaCl (37 °C) to prevent cooling and desiccation.

- Isometric tetanic force measurement

  - Connect the bipolar electrode cables and the ground cable according to their color to a bipolar stimulator device.

  - Attach the other end of the bipolar electrode cables to a subminiature electrode.

NOTE: The reference electrode (red, anode) should be placed distal and the active electrode (black, cathode) proximal.

  - Transfer the animal together with the heating pad to the testing platform.

  - Fixate the hind limb of the rat to the wooden block using two 1 mm Kirschner wires through the ankle and the lateral condyle of the distal femur avoiding the posterior aspect of the knee.

CAUTION: Avoid vascular damage to the popliteal artery and vein which are located dorsally to the femur condyle.

  - Attach a holder with a custom clamp to the testing platform using its vacuum base.

  - Secure the distal TA muscle tendon to the clamp attached to the force transducer.

NOTE: The clamp and force transducer should be positioned parallel to the course of the TA muscle.

  - Place the retractor at the posterolateral thigh of the rat in order to access the common peroneal nerve.

NOTE: The sciatic nerve and its branches should be kept moist with heated 0.9% NaCl (37 °C) to prevent cooling and desiccation.

  - Insert the ground cable in the surrounding muscles (e.g., the vastus lateralis muscle).

NOTE: The Grass SD9 stimulator requires a ground cable to reduce electrical artifacts. Newer stimulators might not require an extra ground cable.

  - Hook the common peroneal nerve to the subminiature electrode and fix its position using the holder on the platform (Figure 4).

NOTE: Ensure that only the common peroneal nerve is hooked to the subminiature.

  - Optimization of the muscle length

    - Turn the bipolar stimulator device on and adjust the settings as follow: square monophasic pulse, delay 2 ms, stimulus pulse duration 0.4 ms, stimulus intensity 2 V.

NOTE: The delay determines the time between the sync out pulse and the delivery of the leading edge of the pulse.

    - Select Parameter test and turn on Trigger collection in the VI.

    - Increase the muscle length (preload) by adjusting the lever arm attached to the force transducer.

    - Start at 10 g of preload and use increments of 10 g until the maximum active muscle force is determined.

    - For each preload, apply two single twitches directly after each other using the button on the bipolar stimulator device. The output will be visible on the screen and the rat should show dorsiflexion of the paw.

NOTE: Before stimulating the nerve, always remove any excess 0.9% NaCl surrounding the nerve using cotton tipped applicators to ensure the signal is not conducted to the surrounding tissue.

    - To stop the measurement, hit Trigger collection again in the VI.

    - If the program automatically detects the two peak output forces click on Accept. In case the program does not automatically select these output forces, press Decline and select the peaks manually. The two peak output forces will be averaged to a mean peak output force (Figure 5).

    - Calculate the active muscle force by subtracting the preload from the mean peak output force.

    - Write down the active force for each preload to visualize the trend and recognize the maximum active force (Figure 6). A spreadsheet can also be used.

  - Measurement of isometric tetanic force

    - After determining the ideal muscle length, let the muscle rest at zero preload for 5 minutes prior to starting the tetanic muscle contractions.

    - Meanwhile, switch from Parameter test to Frequency test on the VI and adjust the stimulus intensity to 10 V on the bipolar stimulator device.

    - Keep the delay and stimulus pulse duration at 2 ms and 0.4 ms, respectively.

    - Measure the isometric tetanic muscle force using increasing stimulus frequencies starting at 30 Hz with increments of 30 Hz until the maximum force plateau is observed.

    - Click on Trigger collection and set to the predetermined optimal muscle length.

    - Press the Repeat button on the bipolar stimulator device to induce a tetanic stimulation for a maximum of 5 seconds or until a force peak is clearly observed.

NOTE: Before stimulating the nerve, always remove any excess 0.9% NaCl surrounding the nerve using cotton tipped applicators to ensure the signal is not conducted to the surrounding tissue.

    - To collect the data, press Trigger collection again and document the maximum output force. In case the program does not automatically detect the peak maximum output force, press Decline and select the peak manually.

    - Let the muscle rest again at zero preload for 5 minutes prior to starting the next tetanic muscle contractions.

NOTE: Regularly (approximately every 5 min), moist the TA muscle with heated 0.9% NaCl (37 °C) to prevent cooling and desiccation.

    - Continue increasing the stimulus frequency until the maximum force plateau is reached. The force plateau will be defined as the maximum isometric tetanic force.

NOTE: After this step, remove the K-wires, staple or suture the skin and repeat the entire procedure to the contralateral hind limb, starting at step 4.

REPRESENTATIVE RESULTS:

Five parameters are used to measure the ITF measurement. These include muscle tension (preload force), stimulus intensity (voltage), stimulus pulse frequency, stimulus duration of 0.4 ms and a delay of 2 ms. Prior to measuring the ITF, the optimal muscle tension has to be determined using two single twitch muscle contractions at an intensity of 2 V during the parameter test. These stimuli cause dorsiflexion of the paw and produce an output signal on the graph in the VI (Figure 5). These single twitch curves ideally have a rapid vertical upswing representing the contraction period directly followed by a slower vertical decrease period demonstrating the relaxation period. The program will average these two peak output forces, but the active force has to be manually calculated by subtracting the preload force from the mean output force. In the example in Figure 5, a preload of 10 g results in two peak output forces of 411.09 g (4.03 N) and 379.78 g (3.73 N), which is averaged to a mean peak output force of 395.43 g (3.88 N). When the active forces of each preload are plot in a graph, the maximum active force can be identified. These active forces usually produce a bell-shaped curve and the maximum active force for Lewis rats weighing 300-500 g should be around 30-40 g (0.29-0.39 N) (Figure 6).

For the tetanic stimulations during the frequency test, the stimulus intensity is increased to a supra-maximal voltage (10 V) to ensure maximal activation of all TA muscle motor units using increasing frequencies. The optimal tetanic curve increases and decreases sharply and has a slowly decreasing plateau phase with minimal oscillations. Figure 7 depicts an example of a tetanic curve at a stimulus frequency of 30 Hz with an isometric tetanic force of 803.25 g (7.88 N). The highest force plateau is defined as the maximum ITF.

FIGURE LEGENDS:

Figure 1: Image of customized clamp fashioned from a surgical hemostat and modified with a tightening screw that allows for adjustment of the tension.

Figure 2: Graphical code for virtual instrument for isometric tetanic force measurement on LabVIEW.

Figure 3: Calibration of the force transducer. Successful calibration of the force transducer with five weights (0, 10, 20, 30 and 50 g) should result in a positive linear curve.

Figure 4: Schematic overview of experimental setup for isometric tetanic force measurement. (Copyrighted and used with permission of the Mayo Foundation for Medical Education and Research; all rights reserved. Reprinted from: Shin, R. H. et al. Isometric tetanic force measurement method of the tibialis anterior in the rat. Microsurgery. 28 (6), 452-457 (2008)).

Figure 5: Representative single twitch curves for optimization of muscle length. For each preload measurement, two single twitches are applied. These single twitch curves have a rapid vertical upswing (contraction period) followed by a vertical decrease (relaxation period). The two peak output forces will be averaged to a mean peak output force. In this example with a Lewis rat, a preload of 10 g results in two peak output forces of 411.09 g (4.03 N) and 379.78 g (3.73 N), which is averaged to a mean peak output force of 395.43 g (3.88 N).

Figure 6: Optimal muscle length (preload). The active muscle force can be calculated by subtracting the preload from the mean peak output force. The active muscle force for each preload should be documented until a drop in active muscle force is visible. The preload yielding the highest active muscle force will be used to measure the isometric tetanic force. The optimal preload for Lewis rats weighing 300-500 g should be around 30-40 g (0.29-0.39 N) (N=10).

Figure 7: Representative isometric tetanic force curve. The optimal tetanic curve increases sharply, then has a slowly decreasing plateau phase followed by a sharp decrease. The highest force plateau is defined as the maximum ITF. This example depicts the tetanic curve at a stimulus frequency of 30 Hz with an isometric tetanic force of 803.25 g (7.88 N).

DISCUSSION:

This protocol describes a previously validated method for acquiring accurate maximum ITF measurements of the TA muscle in the rat model32. The recovery of maximum strength after experimental nerve reconstruction treatments is of primary interest in the clinical setting as it proves that the nerve not only regenerated, but also made working connections with the target muscle. The ITF can be used in a small nerve gap model, such as the rat sciatic nerve model32, and with a few modifications to the protocol, it can also be used in a larger nerve gap rabbit model33.

There are several critical steps that should be considered to ensure consistent and reliable maximum isometric muscle force measurements. The importance of carefully selecting the type of anesthesia to prevent skeletal muscle side effects has previously been described32,33. The use of isoflurane has demonstrated a time dependent decrease in muscle force, which can be explained by its ability to induce sarcoplasmic reticulum stimulated release of calcium33,48. The effect of ketamine/xylazine on the muscle force has proven to be minimal based on our experience and previous study32. Secure attachment of the distal TA muscle tendon to the force transducer is also of great importance for accurate measurements. Slippage or tearing of the tendon should be prevented or directly corrected. Therefore, a custom-made clamp was created from a surgical hemostat and modified with a tightening screw. Other research groups have described a technique of drying the tendon for about 30 minutes to mechanically strengthen the interface between the tendon and a clamp49. In order to maintain endurance of the muscle it is critical to avoid desiccation of the TA muscle and tendon with warm 0.9% NaCl and implement a 5-minutes resting period between each tetanic stimulation. The resting period is based on the activity of the phosphagen system, also known as the immediate energy source, which is important for explosive muscle contractions. It consists of adenosine triphosphate (ATP) and creatine phosphate activity and provides energy for less than 10 seconds of maximal activity. It requires approximately 3-5 minutes to replenish 100% of the phosphagens50.

We recognize the limitations of the method described in this video. The non-survival nature of the procedure does not allow for serial measurements over time. Additionally, it is a detailed and time-consuming testing protocol. During the 1 to 2 hour testing time, the nerve and muscle undergo a significant number of stimulations which may result in muscle fatigue with potential decrease in ITF. This has, however, proven to be less prominent in the rat model compared to the rabbit33.

In conclusion, the ITF measurement described in this video is an invaluable tool in experimental peripheral nerve studies to quantify motor recovery. When presented with other outcome measures such as electrophysiology and histomorphometry, a global assessment of nerve function can be provided.

ACKNOWLEDGMENTS:

Research reported in this publication was supported by the National Institute of Neurological Disorders and Stroke of the National Institutes of Health under Award Number RO1 NS 102360. The content is solely the responsibility of the authors and does not necessarily represent the official views of the National Institutes of Health.

DISCLOSURES:

The authors have nothing to disclose.

REFERENCES:

1	Taylor, C. A., Braza, D., Rice, J. B., Dillingham, T. The incidence of peripheral nerve injury in extremity trauma. American Journal of Physical Medicine & Rehabilitation. 87 (5), 381-385 (2008).

2	Huckhagel, T., Nuchtern, J., Regelsberger, J., Lefering, R., TraumaRegister, D. G. U. Nerve injury in severe trauma with upper extremity involvement: evaluation of 49,382 patients from the TraumaRegister DGU(R) between 2002 and 2015. Scandinavian Journal of Trauma, Resuscitation and Emergency Medicine. 26 (1), 76 (2018).

3	Tapp, M., Wenzinger, E., Tarabishy, S., Ricci, J., Herrera, F. A. The Epidemiology of Upper Extremity Nerve Injuries and Associated Cost in the US Emergency Departments. Annals of Plastic Surgery. 83 (6), 676-680 (2019).

4	Grinsell, D., Keating, C. P. Peripheral nerve reconstruction after injury: a review of clinical and experimental therapies. BioMed Research International. 2014, 698256 (2014).

5	Terzis, J., Faibisoff, B., Williams, B. The nerve gap: suture under tension vs. graft. Plastic and Reconstructive Surgery. 56 (2), 166-170 (1975).

6	Millesi, H. Forty-two years of peripheral nerve surgery. Microsurgery. 14 (4), 228-233 (1993).

7	Wood, M. D., Kemp, S. W., Weber, C., Borschel, G. H., Gordon, T. Outcome measures of peripheral nerve regeneration. Annals of Anatomy-Anatomischer Anzeiger. 193 (4), 321-333 (2011).

8	Alvites, R. et al. Peripheral nerve injury and axonotmesis: State of the art and recent advances. Cogent Medicine. 5 (1), 1466404 (2018).

9	Diogo, C. C. et al. The use of sheep as a model for studying peripheral nerve regeneration following nerve injury: review of the literature. Journal of Neurology Research. 39 (10), 926-939 (2017).

10	Irintchev, A. Potentials and limitations of peripheral nerve injury models in rodents with particular reference to the femoral nerve. Annals of Anatomy. 193 (4), 276-285 (2011).

11	Brenner, M. J. et al. Role of timing in assessment of nerve regeneration. Microsurgery. 28 (4), 265-272 (2008).

12	Vleggeert-Lankamp, C. L. The role of evaluation methods in the assessment of peripheral nerve regeneration through synthetic conduits: a systematic review. Laboratory investigation. Journal of Neurosurgery. 107 (6), 1168-1189 (2007).

13	Deumens, R., Marinangeli, C., Bozkurt, A., Brook, G. A. Assessing motor outcome and functional recovery following nerve injury. Methods in Molecular Biology. 1162, 179-188 (2014).

14	Dellon, A. L., Mackinnon, S. E. Selection of the appropriate parameter to measure neural regeneration. Annals of Plastic Surgery. 23 (3), 197-202 (1989).

15	Munro, C. A., Szalai, J. P., Mackinnon, S. E., Midha, R. Lack of association between outcome measures of nerve regeneration. Muscle Nerve. 21 (8), 1095-1097 (1998).

16	Varejao, A. S., Melo-Pinto, P., Meek, M. F., Filipe, V. M., Bulas-Cruz, J. Methods for the experimental functional assessment of rat sciatic nerve regeneration. Journal of Neurology Research. 26 (2), 186-194 (2004).

17	Hadlock, T. A., Koka, R., Vacanti, J. P., Cheney, M. L. A comparison of assessments of functional recovery in the rat. Journal of the Peripheral Nervous System. 4 (3-4), 258-264 (1999).

18	Kanaya, F., Firrell, J. C., Breidenbach, W. C. Sciatic function index, nerve conduction tests, muscle contraction, and axon morphometry as indicators of regeneration. Plastic and Reconstructive Surgery. 98 (7), 1264-1271, discussion 1272-1264 (1996).

19	Nichols, C. M. et al. Choosing the correct functional assay: a comprehensive assessment of functional tests in the rat. Behavioural Brain Research. 163 (2), 143-158 (2005).

20	Terzis, J. K., Smith, K. J. Repair of severed peripheral nerves: comparison of the "de Medinaceli" and standard microsuture methods. Experimental Neurology. 96 (3), 672-680 (1987).

21	de Medinaceli, L., Freed, W. J., Wyatt, R. J. An index of the functional condition of rat sciatic nerve based on measurements made from walking tracks. Experimental Neurology. 77 (3), 634-643 (1982).

22	Doi, K., Hattori, Y., Tan, S. H., Dhawan, V. Basic science behind functioning free muscle transplantation. Clinics in Plastic Surgery. 29 (4) (2002).

23	Vathana, T. et al. An Anatomic study of the spinal accessory nerve: Extended harvest permits direct nerve transfer to distal plexus targets. Clinical Anatomy. 20 (8), 899-904 (2007).

24	Chaiyasate, K., Schaffner, A., Jackson, I. T., Mittal, V. Comparing FK-506 with basic fibroblast growth factor (b-FGF) on the repair of a peripheral nerve defect using an autogenous vein bridge model. Journal of Investigative Surgery. 22 (6), 401-405 (2009).

25	Lee, B. K., Kim, C. J., Shin, M. S., Cho, Y. S. Diosgenin improves functional recovery from sciatic crushed nerve injury in rats. Journal of Exercise Rehabilitation. 14 (4), 566-572 (2018).

26	Lubiatowski, P., Unsal, F. M., Nair, D., Ozer, K., Siemionow, M. The epineural sleeve technique for nerve graft reconstruction enhances nerve recovery. Microsurgery. 28 (3), 160-167 (2008).

27	Luis, A. L. et al. Use of PLGA 90:10 scaffolds enriched with in vitro-differentiated neural cells for repairing rat sciatic nerve defects. Tissue Engineering, Part A. 14 (6), 979-993 (2008).

28	Shabeeb, D. et al. Histopathological and Functional Evaluation of Radiation-Induced Sciatic Nerve Damage: Melatonin as Radioprotector. Medicina (Kaunas). 55 (8) (2019).

29	Bain, J. R., Mackinnon, S. E., Hunter, D. A. Functional evaluation of complete sciatic, peroneal, and posterior tibial nerve lesions in the rat. Plastic and Reconstructive Surgery. 83 (1), 129-138 (1989).

30	Monte-Raso, V. V., Barbieri, C. H., Mazzer, N., Yamasita, A. C., Barbieri, G. Is the Sciatic Functional Index always reliable and reproducible? Journal of Neuroscience Methods. 170 (2), 255-261 (2008).

31	Lee, J. Y. et al. Functional evaluation in the rat sciatic nerve defect model: a comparison of the sciatic functional index, ankle angles, and isometric tetanic force. Plastic and Reconstructive Surgery. 132 (5), 1173-1180 (2013).

32	Shin, R. H. et al. Isometric tetanic force measurement method of the tibialis anterior in the rat. Microsurgery. 28 (6), 452-457 (2008).

33	Giusti, G. et al. Description and validation of isometric tetanic muscle force test in rabbits. Microsurgery. 32 (1), 35-42 (2012).

34	Bulstra, L. F. et al. Functional Outcome after Reconstruction of a Long Nerve Gap in Rabbits Using Optimized Decellularized Nerve Allografts. Plastic and Reconstructive Surgery. 145 (6), 1442-1450 (2020).

35	Giusti, G. et al. The influence of vascularization of transplanted processed allograft nerve on return of motor function in rats. Microsurgery. 36 (2), 134-143 (2016).

36	Giusti, G. et al. The influence of nerve conduits diameter in motor nerve recovery after segmental nerve repair. Microsurgery. 34 (8), 646-652 (2014).

37	Hundepool, C. A. et al. Comparable functional motor outcomes after repair of peripheral nerve injury with an elastase-processed allograft in a rat sciatic nerve model. Microsurgery. 38 (7), 772-779 (2018).

38	Lee, J. Y. et al. The effect of collagen nerve conduits filled with collagen-glycosaminoglycan matrix on peripheral motor nerve regeneration in a rat model. Journal of Bone and Joint Surgery. 94 (22), 2084-2091 (2012).

39	Shin, R. H., Friedrich, P. F., Crum, B. A., Bishop, A. T., Shin, A. Y. Treatment of a segmental nerve defect in the rat with use of bioabsorbable synthetic nerve conduits: a comparison of commercially available conduits. Journal of Bone and Joint Surgery. 91 (9), 2194-2204 (2009).

40	Coombes, J. S. et al. Effects of vitamin E deficiency on fatigue and muscle contractile properties. Eur J Appl Physiol. 87 (3), 272-277 (2002).

41	Kauvar, D. S., Baer, D. G., Dubick, M. A., Walters, T. J. Effect of fluid resuscitation on acute skeletal muscle ischemia-reperfusion injury after hemorrhagic shock in rats. Journal of the American College of Surgeons. 202 (6), 888-896 (2006).

42	Murlasits, Z. et al. Resistance training increases heat shock protein levels in skeletal muscle of young and old rats. Experimental Gerontology. 41 (4), 398-406 (2006).

43	Zhou, Z., Cornelius, C. P., Eichner, M., Bornemann, A. Reinnervation-induced alterations in rat skeletal muscle. Neurobiology of Disease. 23 (3), 595-602 (2006).

44	Schmoll, M. et al. In-situ measurements of tensile forces in the tibialis anterior tendon of the rat in concentric, isometric, and resisted co-contractions. Physiological Reports. 5 (8) (2017).

45	Heinzel, J. C., Hercher, D., Redl, H. The course of recovery of locomotor function over a 10-week observation period in a rat model of femoral nerve resection and autograft repair. Brain and Behavior. 10 (4), e01580 (2020).

46	Kingery, W. S., Vallin, J. A. The development of chronic mechanical hyperalgesia, autotomy and collateral sprouting following sciatic nerve section in rat. Pain. 38 (3), 321-332 (1989).

47	Weber, R. A., Proctor, W. H., Warner, M. R., Verheyden, C. N. Autotomy and the sciatic functional index. Microsurgery. 14 (5), 323-327 (1993).

48	Kunst, G., Graf, B. M., Schreiner, R., Martin, E., Fink, R. H. Differential effects of sevoflurane, isoflurane, and halothane on Ca2+ release from the sarcoplasmic reticulum of skeletal muscle. Anesthesiology. 91 (1), 179-186 (1999).

49	Schmoll, M. et al. A novel miniature in-line load-cell to measure in-situ tensile forces in the tibialis anterior tendon of rats. PLoS One. 12 (9), e0185209 (2017).

50	Paul, R. J. in Cell Physiology Source Book (Fourth Edition) (ed Nicholas Sperelakis) 801-821 (Academic Press, 2012).