# Submission ID #: 61989

# Scriptwriter Name: Bridget Colvin

Project Page Link: https://www.jove.com/account/file-uploader?src=18906693

Title: Studying Cavitation Enhanced Therapy

Authors and Affiliations: Michael Gray1*, Alexandra V. Vasilyeva1*, Veerle Brans1*, and Eleanor Stride1

*These authors contributed equally

1Institute of Biomedical Engineering, University of Oxford

# Corresponding Authors:

Michael Gray

michael.gray@eng.ox.ac.uk

# Eleanor Stride

# eleanor.stride@eng.ox.ac.uk

# Co-Authors:

alexandra.vasilyeva@eng.ox.ac.uk

veerle.brans@dtc.ox.ac.uk

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or similar?  N

2. Software: Does the part of your protocol being filmed demonstrate software usage?  Y

Videographer: screen capture files to be provided, film for reference only

3. Interview statements: Considering the Covid-19-imposed mask-wearing and social distancing recommendations, which interview statement filming option is the most appropriate for your group? Please select one.

☒ 	Interviewees self-record interview statements outside of the filming date. JoVE can provide support for this option.

4. Filming location: Will the filming need to take place in multiple locations (greater than walking distance)?   N

Protocol Length

Number of Shots: 31

# Introduction

- Introductory Interview Statements

REQUIRED:

  - Michael Gray: These benchtop system design, data collection, and analysis principles can be used as a rational foundation for the in vitro study of cavitation enhanced therapies [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

REQUIRED:

  - Michael Gray: Our techniques enable high throughput testing while preserving the critical experimental attributes of integral cavitation monitoring, repeatable sample alignment, and compatibility with common cellular analysis methods [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL:

  - Eleanor Stride: Cavitation enhanced therapy has multiple potential applications, including the treatment of diseases such as cancer and stroke. By better understanding the underpinning mechanisms, we can develop better therapies [1].

    - LAB MEDIA: WIN_20201213_18_49_36_Pro.mp4. Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL:

  - Veerle Brans: This easily reproducible system design and implementation framework allows the study of a wide variety of cavitation-induced cellular bioeffects, including drug-delivery, sonoprinting, and sonoporation [1].

    - LAB MEDIA: Named talent says the statement above in an interview-style shot, looking slightly off-camera NOTE: Uploaded to AWS

OPTIONAL:

  - Alexandra Vasilyeva: Obtaining a meaningful result requires ensuring reproducibility and accounting for all of the variables in the experiment. It is also imperative to perform all of the relevant controls and to measure the electrical noise [1].

    - LAB MEDIA: Named talent says the statement above in an interview-style shot, looking slightly off-camera NOTE: Uploaded to AWS

# Protocol

- Systems for Acoustic Transfection (SAT) Preparation

  - For preparation of the first system for acoustic transfection, degas the fill liquid under a pressure of minus 105 pascals for at least two hours to minimize the likelihood of cavitation in the propagation path [1]. Confirmation with a dissolved oxygen probe that the partial pressure of oxygen is below 10 kilopascals is recommended [2].

    - WIDE: Talent degassing fill liquid

    - Talent checking pressure with probe NOTE: don’t use take 1

  - Fill the test chamber slowly to minimize the re-introduction of air into the degassed liquid [1-TXT] and immediately clear any residual bubbles from the transducer and medium container surfaces [2-TXT].

    - Talent filling chamber TEXT: Continue degassing filled chamber as necessary

    - Talent clearing bubbles TEXT: Repeat bubble clearing just before exposure experiment initiation

    - Make sure the temperature of the chamber and its contents have stabilized before commencing exposure experiments.

  - Allow the ultrasound source power amplifier to warm up per the manufacturer’s recommendation so that the gain and output are stable with respect to time [1] and dilute the cavitation agent while gently and continuously stirring to make a uniform suspension without entrapping macrobubbles or destroying the agent [2].

    - Talent setting power source to warm

    - Talent diluting/stirring cavitation agent

  - When working with microbubbles, make sure to handle them carefully [1-TXT].

    - Agent being diluted/mixed TEXT: e.g. 18 G blunt fill needle NOTE: Audio slate

- SAT2 Preparation

  - For SAT2 preparation, before beginning the experiment, use ethanol to sterilize the PDMS (P-D-M-S) lid [1-TXT].

    - WIDE: Talent sterilizing lid TEXT: PDMS: polydimethlysiloxane

  - Press fit the sterilized lid to the culture dish to form the cell compartment [1] and load a 10-milliliter syringe equipped with an 18-gauge blunt needle with 10 milliliters of the fill liquid [2].

    - Talent press fitting lid onto culture dish

    - Talent loading syringe with solution

  - Insert the needle through one of the PDMS fill holes [1] and slowly fill the chamber while tilting so that macrobubbles can escape through the open fill hole [2].

    - Talent inserting needle into hole NOTE: This and next shot together Videographer: Important step

    - Chamber being filled/chamber being tilted Videographer: Important step TEXT: Tilt so open hole is vertically above fill hole

  - When the chamber has been filled, insert a 4-5-millimeter polymer rod into the open hole and position the assembly so that both holes are horizontal [1].

    - Rod being inserted into hole NOTE: Audio slate Videographer: Important step

    - Talent positioning assembly Videographer: Important step

  - Remove the needle while injecting extra fluid so that air is not drawn into the chamber [1] and close the fill hole with another polymer rod [2].

    - Needle being retracted/fluid being injected NOTE: 3.4.1, 3.5.1, and 3.5.2 shot together Videographer: Important/difficult step

    - Hole being closed Videographer: Important/difficult step

  - Visually check the compartment for evidence of entrapped macrobubbles [1-TXT] and press fit the cell exposure compartment into the compartment holder [2]. Consider the buoyancy of the particles in suspension and how their buoyancy will affect their contact with cells when deciding the orientation of the cell exposure compartment [3-TXT].

    - Compartment being checked for bubbles TEXT: Repeat filling if bubbles observed

    - Compartment being press fit into holder

    - Compartment being positioned NOTE: This shot was removed from 3.6, use this action for 3.7.1. TEXT: Use as little force as possible to minimize cell growth surface and detachment flexure

  - Then, lowering the chamber lid at a horizontal angle to discourage macrobubbles from resting on the submerged parts of the device, install the lid onto the top of the chamber [1].

    - Lid being lowered/placed Videographer: Important step

- Data Collection

  - Before performing the experimental measurements, allow the suspension to thermally equilibrate with the chamber temperature [1], using a fine needle thermocouple to confirm that the temperature in the chamber has stabilized [2].

    - WIDE: Talent checking compartment in holder TEXT: Establish control liquid background response and background electronic noise levels before measuring experimental agent data

    - Talent using thermocouple

  - To monitor the experiments in real time in both the time and frequency domains, start the data collection process [1] and turn on the ultrasound source drive signal [2].

    - Talent starting data collection process

    - Talent turning on ultrasound source drive signal

  - Use a high voltage probe to monitor the amplifier output signal that drives the ultrasound source throughout the experiment to ensure that the exposure is proceeding as expected [1] and make sure that the oscilloscope is set to compensate for the probe attenuation [2].

    - Talent using high voltage probe to monitor amplifier output signal

    - Talent checking oscilloscope setting(s)

  - Time domain monitoring of the passive cavitation detector reveals whether signals are sized appropriately for the current instrumentation settings and whether the cavitation signals are seen earlier than expected [1-TXT].

    - SCREEN: JoVE 4.4.1 screen capture.mp4. Shot of time domain TEXT: Caution: Avoid signal clipping and chamber leakage

  - Frequency domain monitoring allows analysis of the type of bubble behavior and can be used to adjust drive levels as necessary to achieve the desired cell stimulus [1].

    - SCREEN: JoVE 4.5.1 screen capture.mp4. Shot of frequency domain, then drive levels being adjusted

## Protocol Script Questions

A. Which steps from the protocol are the most important for viewers to see? Please list 4 to 6 individual steps.

3.3.-3.5., 3.7.

B. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1 or 2 individual steps from the script above.

3.5. fill and close without entrapping macrobubbles.

# Results

- Results: Representative Passive Cavitation Detector Responses

  - In this analysis, at the lowest incident pressure, the passive cavitation detector response consisted entirely of integer harmonics of the 0.5-megahertz fundamental ultrasound frequency [1].

    - LAB MEDIA: Figure 4 Video Editor: please emphasize red data line in Figure 4A

  - Increasing from 0.2 to 0.3 megapascals resulted in pronounced ultraharmonics in the spectrum in addition to further elevated integer harmonics [1]. The time domain waveforms at these two pressures looked similar, although the 0.3-megapascal results demonstrated more variability over the pulse duration [2].

    - LAB MEDIA: Figure 4 Video Editor: please emphasize red data line in Figure 4C

    - LAB MEDIA: Figure 4 Video Editor: please emphasize red data line in inset in Figure 4C

  - At the highest pressure, the time domain waveform amplitude grew nonlinearly relative to the lower pressures [1] as a result of clearly elevated broadband noise due to inertial cavitation caused by microbubble destruction [2].

    - LAB MEDIA: Figure 4 Video Editor: please emphasize red data line in Figure 4E

    - LAB MEDIA: Figure 4 Video Editor: please emphasize green data line in Figure 4F

  - Here full spectra are shown over a 50-second exposure time [1], during which the source emitted 2-millisecond pulses every 0.2 seconds [2].

    - LAB MEDIA: Figure 5

    - LAB MEDIA: Figure 5 Video Editor: please emphasize Figure 5A

  - As observed in this graph illustrating the corresponding total, harmonic, and broadband powers, large-amplitude broadband responses were generated [1], with the initial spike considered to correlate with the destruction of the largest bubbles [2].

    - LAB MEDIA: Figure 5 Video Editor: please sequentially add/emphasize red, blue, and green data lines

    - LAB MEDIA: Figure 5 Video Editor: please emphasize peak in all three data lines between 0-10 s

  - After a few seconds, the broadband response rapidly diminished, apparently due to bubble destruction [1].

    - LAB MEDIA: Figure 5 Video Editor: please emphasize data lines from 10-50 s

  - In this analysis using a 20:1 dilution of microbubbles in normal PBS [1], the time and sample-averaged spectra [2] showed that the unfocused passive cavitation detector contained a stronger broadband response, accompanied by a reduced sample-to-sample variability in both the harmonic [3] and ultraharmonic powers [4].

    - LAB MEDIA: Figure 6

    - LAB MEDIA: Figure 6 Video Editor: please emphasize Figure 6A

    - LAB MEDIA: Figure 6 Video Editor: please emphasize Figure 6B

    - LAB MEDIA: Figure 6 Video Editor: please emphasize 6C

# Conclusion

- Conclusion Interview Statements

  - Veerle Brans: Cavitation-induced bioeffects can be assessed using techniques such as fluorescence microscopy, flowcytometry, or biological assays to enable the establishment of a robust relationship between cavitation activity and biological responses [1].

    - LAB MEDIA: Named talent says the statement above in an interview-style shot, looking slightly off-camera  NOTE: Uploaded to AWS as 6.2.1.

  - Eleanor Stride: This technique has allowed us to better identify the relationships between bubble behavior and biological effects, which have revealed some potential new mechanisms underpinning cavitation-mediated drug delivery [1].

    - LAB MEDIA: WIN_20201213_18_50_16_Pro.mp4. Named talent says the statement above in an interview-style shot, looking slightly off-camera