# Submission ID #: 62015

# Scriptwriter Name: Bridget Colvin

Project Page Link: https://www.jove.com/account/file-uploader?src=18914623

# Title: Quantitative Atomic-Site Analysis of Functional Dopants/Point Defects in Crystalline Materials by Electron-Channeling-Enhanced Microanalysis

Authors and Affiliations: Masahiro Ohtsuka1 and Shunsuke Muto1

1Electron Nanoscopy Division, Advanced Measurement Technology Center, Institute of Materials & Systems for Sustainability, Nagoya University

☐   All author names are spelled correctly, and the affiliations are correct (city/state/country information not included in video title page).

# Corresponding Author:

Shunsuke Muto

smuto@imass.nagoya-u.ac.jp

# Co-authors:

m-ohtsuka@nagoya-u.jp

## Author Questionnaire

1. Microscopy: Does your protocol demonstrate the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or similar?  N

Enter make and model of microscope.

2. Software: Does the part of your protocol being filmed demonstrate software usage?  Y

If Yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps. Please upload all screen captured video files to your project page as soon as possible.

3. Interview statements: Considering the Covid-19-imposed mask-wearing and social distancing recommendations, which interview statement filming option is the most appropriate for your group? Please select one.

☐ 	Interviewees wear masks until the videographer steps away (≥6 ft/2 m) and begins filming. The interviewee then removes the mask for line delivery only. When the shot is acquired, the interviewee puts the mask back on. Statements can be filmed outside if weather permits.

☐ 	Interviewees self-record interview statements outside of the filming date. JoVE can provide support for this option.

☐ 	Interview Statements are read by JoVE’s voiceover talent.

☐ 	Author interview statement opt out (interview statements removed completely from paper).

3. Filming location: Will the filming need to take place in multiple locations (greater than walking distance)?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

- To ensure that your protocol can be filmed in one day, the protocol is restricted to 25 steps and/or 55 shots.

Protocol Length

Number of Shots: 39

# Introduction

- Introductory Interview Statements

Your answers to these questions will become author interview statements, which will be incorporated into the script. Authors will memorize the statements and then deliver them on camera.

- Enter the full name of the author who will deliver the statement on camera.

- Each author should deliver no more than two statements.

- Fill out both required statements. One optional statement may also be selected.

- Please answer in full sentences in a style suitable for being spoken aloud.

- Limit the length of each statement to 30 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: Why is your protocol significant? OR What key questions can this method help answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

REQUIRED: What is the main advantage of this technique?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Do the implications of this technique extend toward the therapy (or diagnosis) of a particular disease, disability, or challenge? How so?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Are there any specific areas of research that this method could provide insight into? OR Can this method be applied to any other systems?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: How would you expect an individual who has never performed this technique to struggle? Do you have any advice to offer to somebody who is trying this technique for the first time?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Why is visual demonstration of this method critical?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Has any author been assigned more than two statements? ☐ No

# Introduction of Demonstrator on Camera

Complete this statement ONLY if any of the individuals who will be demonstrating the procedure on camera will not be delivering an Introductory Interview Statement.

  - Enter name of author who will introduce demonstrator: Demonstrating the procedure will be Click here to enter name of demonstrator(s)., a Click here to enter demonstrator job title. from my laboratory. Include additional demonstrators as needed.[1][2].

    - INTERVIEW: Author saying the above

    - The named demonstrator(s) looks up from workbench or desk or microscope and acknowledges the camera

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The one-digit numbers represent sections of the video. The text will appear onscreen.

- The two-digit numbers (e.g. 2.1., 2.2.) represent steps of your protocol. The text will be recorded by a professional voiceover talent.

- The three-digit numbers (e.g. 2.1.1., 2.2.2.) represent the shots that our videographer will capture at your lab.

Please use this draft script to help you prepare for filming day.

- Filming should take no more than 10 minutes per step. If a step will take more than 10 minutes, prepare the product from that step in advance.

- Transmission Electron Microscopy (TEM) Alignment for Beam-Rocking

  - To mount a thin film sample for transmission electron microscopy, load the sample onto a double-tilting transmission electron microscope sample holder [1] and insert the holder into a transmission electron microscope equipped with a scanning mode and an energy-dispersive X-ray detector for imaging [2].

    - WIDE: Talent mounting sample onto sample holder

    - Talent inserting holder into microscope

  - After a routine transmission electron microscope beam alignment procedure has been performed [1], click Attachment Scanning Image Display to navigate to STEM mode [2].

    - SCREEN: To be provided by Authors: Beam alignment being performed

    - SCREEN: To be provided by Authors: ASIC being clicked

  - To perform an optical axis alignment, click Rocking and Stop to start and stop the beam-rocking, respectively [1].

    - Talent at microscope, initiating imaging, with monitor visible in frame

    - SCREEN: To be provided by Authors: Rocking and Stop being clicked

  - When the rocking has stopped, remove the sample from the field of view [1] and use the Magnification increment-decrement buttons to set the beam-rocking range to less than plus or minus 2 degrees [2].

    - Talent removing/moving sample

    - SCREEN: To be provided by Authors: Range being adjusted

  - Turn the Brightness knob clockwise to the limit [1] and adjust the Object Focus Coarse knob counterclockwise to an underfocused condition [2]. A caustic spot will appear on the fluorescent viewing screen [3].

    - Talent turning brightness knob

    - Talent turning Obj focus coarse knob

    - SCREEN: To be provided by Authors: Caustic spot appearing

  - Press BRIGHT TILT [1] and use the DEF-STIG X-Y knobs to move the caustic spot to the center of the fluorescent screen [2].

    - Talent pressing bright tilt

    - Talent moving spot, with monitor visible in frame

  - Press the Standard Focus button [1] and turn the BRIGHTNESS knob counterclockwise until an alternative caustic spot appears on the fluorescent screen [2].

    - Talent pressing button

    - SCREEN: To be provided by Authors: Knob being turned/alternative spot appearing on screen

  - Press F3 [1] and use the DEF-STIG X-Y knobs to move the beam spot to the center of the screen [2].

    - Talent pressing F3 TEXT: Alternative: Click Spot button in TEM Alignment Panel for Maintenance window

    - Talent moving knobs, spot moving to center of screen, with monitor visible in frame

  - Then repeat the just demonstrated optical alignment steps until the beam position remains in the center even if the lens condition is switched [1].

    - SCREEN: To be provided by Authors: Rocking and Spot being clicked

- Incident Beam Collimation and Pivot Point Setup

  - To collimate the incident beam, first use the two attached screws of the aperture knob to turn the knob clockwise to introduce the third largest condenser aperture at the center of the optical axis [1] and use the BRIGHTNESS knob in conjunction with the DEF-STIG knobs and Condenser Stigmator to adjust the condenser lens stigmator until the beam shape is coaxially defocused [2].

    - WIDE: Talent turning knob/aperture being introduced

    - Talent turning knob(s)/stigmator being adjusted

  - Press HT WOBB and adjust the BRIGHT TILT knob to minimize the beam size fluctuation with the change in acceleration voltage to adjust the beam convergence angle to a minimum [1]. Press HT WOBB again to stop the HT wobbler [2].

    - Talent pressing button and adjusting knob

    - Talent pressing button

  - To set the pivot point, activate the maintenance mode according to the manufacturer’s instructions [1] and select JEOL, Scan-Focus window, and Scan Control [2].

    - Talent activating maintenance mode, with monitor visible in frame

    - SCREEN: To be provided by Authors: JEOL, Scan-Focus window, and Scan Control being clicked

  - Click Cor and Scan and use the DEF-STIG and Object Focus Fine knobs to minimize the beam shift with the beam-rocking [1-TXT].

    - Talent turning knobs, with monitor visible in frame TEXT: Turn OBJ FOCUS FINE slightly counterclockwise

  - Then use the Z control keys to match the sample and pivot point height [1] so that the sample is focused on the fluorescent screen [2].

    - Talent using Z control keys

    - SCREEN: To be provided by Authors: Sample coming into focus on screen

- Electron-Channeling Pattern (ECP) Acquisition

  - To perform a final beam alignment to obtain an electron-channeling pattern for the sample, move the sample area of interest back to the center [1] and click Scan to start the beam-rocking [2].

    - WIDE: Talent moving sample into center

    - Talent clicking Scan, with monitor visible in frame

  - Manually turn the annular dark field detector cylinder clockwise [1] and insert the detector [2].

    - Talent turning cylinder

    - Talent inserting detector

  - Adjust the DEF-STIG knobs while holding the PLA key to set the detector position to the center of the beam position [1] and check STEI-DF and STEM. An electron-channeling pattern will appear [2].

    - Talent adjusting knob(s)/holding button/position being centered

    - SCREEN: To be provided by Authors: STEI-DF and STEM being checked, then ECP appearing

  - Adjust the Brightness and Contrast to best view the pattern, slightly turning the BRIGHTNESS knob obtain the sharpest contrast as necessary [1].

    - SCREEN: To be provided by Authors: Brightness and contrast being adjusted/sharpened

- Energy-Dispersive X-Ray Analysis

  - To collect the energy-dispersive x-ray spectra, in the beam-rocking mode, use the spectral imaging method [1] as a function of the beam tilting angles in the x and y directions to display the elemental intensity distribution for specified elements [2].

    - WIDE: Talent selecting spectral imaging method, with monitor visible in frame

    - SCREEN: To be provided by Authors: Intensity distribution being displayed for at least one element

  - To obtain an ionizing-channeling pattern, use the Line Scan function to perform a 1D tilting measurement of a systematic row of reflections. Yellow arrows will appear in the electron-channeling pattern preview to specify the measuring range [1].

    - SCREEN: To be provided by Authors: Line Scan being selected, then measurement being performed and yellow arrow(s) appearing Video Editor: please indicate yellow arrow(s) when they appear

  - Stop the measurements when sufficient data statistics are obtained [1].

    - SCREEN: To be provided by Authors: Shot of sufficient data, then measurement being stopped

## Protocol Script Questions

# Authors: Please use the step numbers from the script above (not step numbers from the manuscript) when answering the questions below. Please do not include steps that will be screen-captured and do not list entire sections.

A. Which steps from the protocol are the most important for viewers to see? Please list 4 to 6 individual steps.

Click here to list 4 to 6 individual steps, using the step numbers from the protocol section of the video script.

B. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1 or 2 individual steps from the script above.

Click here to list 1 or 2 individual steps, using the step numbers from the protocol section of the video script.

# Results

Please review this section to make sure that it accurately reflects your findings.

- Use Track Changes when making edits or revisions.

- If you would like the video to include different results, please revise this section.

- When revising, please keep the length of the voiceover below 200 words. Current word count: 166. (Voiceover is the text that follows the two-digit numbers.)

- Please note that the video cannot include voiceover without an accompanying visual.

- Notes in blue italics are for JoVE’s video editor.

- Results: Representative ECP and ICP Emission Imaging

  - In these representative images [1], experimental electron- and ionization-channeling patterns for barium titanate [2] and barium-lanthanum, titanium-potassium-alpha, and oxygen-potassium-alpha [3] near the 100 [4] and 110 zone axes, respectively, are shown [5].

    - LAB MEDIA: Figure 6

    - LAB MEDIA: Figure 6 Video Editor: please emphasize ECP image column

    - LAB MEDIA: Figure 6 Video Editor: please sequentially emphasize Ba-L ICP, Ti-K ICP, and O-K ICP image columns

    - LAB MEDIA: Figure 6 Video Editor: please emphasize [100] image row

    - LAB MEDIA: Figure 6 Video Editor: please emphasize [110] image row

  - Here the electron- [1] and ionization-channeling patterns of calcium-potassium [2], tin-lanthanum [3], oxygen-potassium [4], europium-lanthanum [5], and yttrium-lanthanum for the Ca1.8Eu0.2Y0.2Sn0.8O4 sample near the 100 zone [6].

    - LAB MEDIA: Figure 7 Video Editor: please emphasize ECP image

    - LAB MEDIA: Figure 7 Video Editor: please emphasize Ca-K ICP image

    - LAB MEDIA: Figure 7 Video Editor: please emphasize Sn-L image

    - LAB MEDIA: Figure 7 Video Editor: please emphasize O-K image

    - LAB MEDIA: Figure 7 Video Editor: please emphasize Eu-L image

    - LAB MEDIA: Figure 7 Video Editor: please emphasize Y-L image

  - The europium-lanthanum ionizing-channel pattern in this analysis [1] was closer to the calcium-potassium pattern [2], while the yttrium-lanthanum pattern [3] was closer to that observed for tin-lanthanum [4]. These data suggest that the europium and yttrium occupation sites could be biased, as expected [5].

    - LAB MEDIA: Figure 7 Video Editor: please emphasize Eu-L pattern

    - LAB MEDIA: Figure 7 Video Editor: please emphasize Ca-K pattern

    - LAB MEDIA: Figure 7 Video Editor: please emphasize Y-L pattern

    - LAB MEDIA: Figure 7 Video Editor: please emphasize Sn-L pattern

    - LAB MEDIA: Figure 7

  - The site occupancies of the impurities [1] and the impurity concentrations of all of the samples are indicated in the Table [2].

    - LAB MEDIA: Table 1 Video Editor: please emphasize fCa and fSn columns

    - LAB MEDIA: Table 1 Video Editor: please emphasize cx column

  - As observed, for Ca1.9Eu0.2Sn0.9O4 [1], europium occupied the calcium and tin sites equally, consistent with the results of the x-ray diffraction-Rietveld analysis [2].

    - LAB MEDIA: Table 1

    - LAB MEDIA: Table 1 Video Editor: please emphasize the top row alphaCa and alphaSn values

  - In contrast, europium and yttrium occupied the calcium [1] and tin sites in the co-doped samples at ratios of approximately 7:3 and 4:6, respectively [2], significantly biased as expected while maintaining the charge neutrality condition within the present experimental accuracies [3].

    - LAB MEDIA: Table 1 Video Editor: please emphasize bottom row alphaCa values

    - LAB MEDIA: Table 1 Video Editor: please emphasize bottom row alphaSn values

    - LAB MEDIA: Table 1

# Conclusion

- Conclusion Interview Statements

Below are prompts for interview statements that can be used to further emphasize the significance of your protocol.

- Please answer one or two of the statements.

- Each statement is limited to 30 words.

- Answer the questions in full sentences; you will need to memorize and deliver the sentences as spoken interview statements during filming.

- Indicate the full name of the author who will give each statement.

# What is the most important thing to remember when attempting this procedure? Please indicate the steps (e.g., 2.4., 2.5.) in the Protocol section of the script that this advice applies to.

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera (Enter step numbers referred to.)

# Following this procedure, what other methods can be performed? What questions would these additional methods answer?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

# After its development, did this technique pave the way for researchers to explore new questions within a specific scientific field? If so, how?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Thank you for addressing our questions. We will incorporate your answers and suggestions and send you the final script before your filming day. You will also receive detailed preparation instructions in the email accompanying the final script.