JoVE62053

„A simplified stepwise approach to echo guidance during the MitraClip procedure”

Dear Vineeta Bajaj,

We would like to thank the editorial team of JoVE for the opportunity to revise our manuscript. We would also like to thank the reviewers for their comments and ideas for further improvement of the manuscript. Below you will read our response to the reviewers´ comments along with the revised passages highlighted in blue.

Editorial and production comments:

Changes to be made by the Author(s):

1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues.

The manuscript has been proofread from a native speaker.

2. Please remove the commercial term MitraClip from the title and use generic term instead. Please ensure that the title directly reflects the protocol being presented.

The title has been changed in:

A simplified stepwise approach to echo guidance during percutaneous mitral valve repair.

3. JoVE policy states that the video narrative is objective and not biased towards a particular product featured in the video. The goal of this policy is to focus on the science rather than to present a technique as an advertisement for a specific item. To this end, we ask that you please reduce the number of instances of "MitraClip" within your text. The term may be introduced once somewhere in the introduction but please use it when directly relevant. Otherwise, please refer to the term using generic language.

We reduced the instances of the term “MitraClip” within the text as much as possible. To present the study results concerning the MitraClip or the Pascal system, we preferred to not use the generic language for better understanding.

4. JoVE cannot publish manuscripts containing commercial language. This includes trademark symbols (™), registered symbols (®), and company names before an instrument or reagent. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials and Reagents.
For example: MitraClips, (Abbott Laboratories, Melo Park, California, USA), PASCAL transcatheter, Edwards Lifesciences, Irvine, California, etc.

As mentioned before we reduced the term MitraClip or Pascal system within the text as much as possible and removed all trademark symbols (™) and registered symbols (®) within the text.

5. Please rephrase the Summary to clearly describe the protocol and its applications in complete sentences between 10-50 words: “This protocol presents…”

We rephrased the Summary as recommended.

This protocol presents in detail how to perform real-time echocardiographic guidance during transcatheter mitral valve repair. The fundamental views and the necessary measurements are described to each stage of the procedure.

6. Please expand the Introduction include all of the following with citation:

Echocardiography during a percutaneous mitral valve reconstruction procedure is mandatory. We want to explain the method of echocardiographic visualisation during this procedure. To the best of our knowledge, there are no studies of a percutaneous valve repair without echocardiography as there a no studies to prove that general anesthesia is necessary for open-heart surgery. Although there are various options how to perform general anesthesia, up to this point there is no alternative to echocardiographc visulisation during a percutaneous MV repair. Measurements such as 2D and 3D need to be performed during the procedure. There are different manufacturers of echocardiography machines, but these measurements are the same. All in all, performing percutaneous MV reconstruction without real-time echocardiography is not possible, so there are no studies comparing the procedure with and without echo. Thus, it is also not possible to describe whether echocardiography is the adequate method for performing MC.

a) A clear statement of the overall goal of this method

We added a clear statement of the overall goal of this method to the protocol.

Mitral regurgitation is the second-most frequent indication for valve surgery in Europe. Untreated, it may lead to severe heart failure and reduced life quality . The percutaneous transcatheter mitral valve repair is a catheter-based technique, that mimics the Alfieri stich surgical method to mitral repair by connecting together the A2 and P2 scallops.

b) The rationale behind the development and/or use of this technique

The percutaneous mitral valve repair (PMVR) is a catheter-based technique, which mimics the Alfieri stitch surgical method to mitral repair by connecting the A2 and P2 scallops.

c) The advantages over alternative techniques with applicable references to previous studies

The advantage of a percutaneous approach over an open-heart surgery is explained in the introduction.

For patients with high surgical risk, this technique offers a minimal invasive approach for the treatment of severe mitral regurgitation.

d) A description of the context of the technique in the wider body of literature

A description of the context of the technique in the wider body of the literature is in the introduction.

MitraClip has been the world’s first transcatheter mitral valve repair therapy. Data of several registries and trials have shown, that the MitraClip procedure is an effective and safe method. In 2019 a similar device, the PASCAL transcatheter valve repair system was introduced to the market. It has shown feasibility and acceptable safety in the treatment of patients with severe mitral regurgitation.

e) Information to help readers to determine whether the method is appropriate for their application

This information is included in the paper in the introduction.

In contrast to other percutaneous techniques, like the percutaneous transvalvular replacement (TAVR), that can be done with fluoroscopy only, the transcatheter edge-to-edge repair of the mitral valve requires echocardiographic guidance.  This review describes step by step the echocardiographic approach during transcatheter mitral valve repair, including measurements, suggestions for intraprocedural quantification of the mitral regurgitation and important views to prevent periprocedural complications.

7. Please include an ethics statement before the numbered protocol steps, indicating that the protocol follows the guidelines of your institution’s human research ethics committee.

We added the phrase:

The protocol follows the guidelines of our’s institution’s human research ethics committee.

8. Please adjust the numbering of the Protocol to follow the JoVE Instructions for Authors. For example, 1 should be followed by 1.1 and then 1.1.1 and 1.1.2 if necessary.

We adjusted the numbering of the protocol to follow the Jove Instructions for authors.

9. Please ensure that all text in the protocol section is written in the imperative tense as if telling someone how to do the technique (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “Note.”

We changed the text and rewrote the sentences in the imperative tense.

10. The Protocol should contain only action items that direct the reader to do something.

We removed the information about the different clip sizes, as these are only available for the MitraClip System. This information does also not direct the reader to do something.

- Strategy

  - Discuss the strategy with the operator before inserting the steerable guide catheter (SGC) and the clip delivery system (CDS) into the left atrium.

    - Evaluate a one device strategy if the orifice is < 1 cm wide and position the clip directly above the regurgitation jet if the orifice is circular.

    - Evaluate the implantation of ≥ 2 clips in case of large elliptic or multiple jets. Implant the device starting medially of the regurgitant orifice, as the positioning of a second device is often easier when the first has been implanted in this way, rather than after starting laterally.

11. Please add more details to your protocol steps. Please ensure you answer the “how” question, i.e., how is the step performed? This can be done by including mechanical actions, button clicks, knob turns etc.

As recommended we have added more details to our protocol steps. We described mechanical actions e.g. “insert the probe deeper”, “retroflex the probe”, and “rotate the probe anti-clockwise”. In addition to that, we added information on button clicks e.g. “take an X-plane”, “asses xx with CW”, “take a 3D-colour data-set”, “take a wide sector zoom image”, “record..”

We described knob turns e.g. “sweep to degree..”, “Perform a 180° clockwise rotation into the 3D en-face surgical view..”

12. Please revise the protocol text to avoid the use of any personal pronouns in the protocol (e.g., "we", "you", "our" etc.).

We revised the protocol and deleted the personal pronouns.

Please include a single line space between each protocol step.

We included a single line space between each protocol step.

14. Please include a figure or a table in the Representative Results showing the effectiveness of your technique backed up with data. Example: Outcomes, before and after surgery tables.

Our protocol explains in detail the general method of echocardiographic visualization during a percutaneous mitral valve repair. Our goal of this paper was not the representation of any data. We did not produce any results, because, as already mentioned above, no comparing studies on echocardiography during percutaneous mitral valve reconstruction have been carried out. Nevertheless, we included slides with the results of previous studies of the percutaneous mitral valve repair as in the video.

15. Please include details about the patients (number, clinical symptoms, age, sex, etc) analyzed with this protocol.

Not applicable, general recommendation. Our protocol explains in detail the general method of echocardiographic visualization during a percutaneous mitral valve repair.

16. Please ensure that the results are described with respect to your experiment, you performed an experiment, how did it help you to conclude what you wanted to and how is it in line with the title.

Our protocol explains in detail echocardiographic visualisation during percutaneous mitral valve reconstruction. This is not an experiment, but a method that makes the intervention possible. In our case, the result of the method may be the change of the mitral regurgitation (reduced, increased, or absent). Under “Results” we described in detail the results of previous studies which examined the procedure and the quantification of the mitral valve regurgitation (result of the method) with the corresponding literature.

17. Each Figure Legend should include a title and a short description of the data presented in the Figure and relevant symbols.

We added a title and a short description to each Figure

18. Please obtain explicit copyright permission to reuse any figures from a previous publication. Explicit permission can be expressed in the form of a letter from the editor or a link to the editorial policy that allows re-prints. Please upload this information as a .doc or .docx file to your Editorial Manager account. The Figure must be cited appropriately in the Figure Legend, i.e. “This figure has been modified from [citation].”

We changed the previous figures with the original figures from the papers and complemented the citation of the original paper under the figures.

19. As we are a methods journal, please ensure that the Discussion explicitly cover the following in detail in 3-6 paragraphs with citations:
a) Critical steps within the protocol

Critical steps within the protocol were added to the paper.

Complications due to echocardiography may occur but rarely lead to significant damage. Nevertheless, esophageal lesions are possible after performing transesophageal echocardiography. This incidence is reduced with a shorter duration of the intervention. On the contrary, several complications associated with the transcatheter edge-to-edge mitral valve repair are described. Major bleeding is the most frequent major complication and blood transfusion after transcatheter valve procedures have been strongly associated with mortality. Pericardial tamponade is rare and may occur during several steps of the procedure. Caution during transseptal puncture, insertion of the SGC into the left atrium and retrieving of the SGC-tip after clip deployment reduces the probability of this life-threatening complication. Rates of clip-specific complications, as embolization and partial clip detachment, are low and can be reduced by an accurate echocardiographic assessment during and after grasping of the leaflets.

b) Any modifications and troubleshooting of the technique

Despite the carefulness during echocardiography, aggravation of the MR, may occur. Additional clip deployment in the medial and lateral position may increase the risk of entangling and rupturing of the chordae. Furthermore, repeated grasping of the leaflets can lead to leaflet laceration and perforation. Residual MR is not only a determinant of procedural success but is also associated with poorer prognosis, especially when there is a concomitant transmitral pressure gradient >5 mm Hg. Clear communication with the interventionist and a meticulous echocardiographic assessment are necessary to reduce peri-procedural clip-related complications.

c) Any limitations of the technique

The percutaneous edge-to-edge repair is not eligible for every pathomechanism of MR. For some determinants of MR, such as annulus dilatation, rupture of the chordae tendinae or severe calcifications of the mitral leaflets, the device reaches its limits.

d) The significance with respect to existing methods

In this case, mitral ring dilation can be sutured surgically or percutaneously in place to provide annular stabilization.  Polytetrafluoroethylene (PTFE) neochordae can be surgically sewed as new chordae tendinee. A MV with severe calcified leaflets may be replaced with a prosthetic valve.

e) Any future applications of the technique

New percutaneous devices designed to address these aspects are under development.

20. Please revise the table of the essential supplies, reagents, and equipment. The table should include the name, company, and catalog number of all relevant materials in separate columns in an xls/xlsx file.

We complemented the table of essential supplies.

Reviewer #1:

The manuscript by Gatti et al addresses a real clinical need about transcatheteter edge to edge MV repair: echocardiographic guidance is really mandatory and essential for procedural safety and success. As written in the manuscript, it is essential to build a shared planning between the imager and the interventionalist and the echo protocol should consider peculiarities related to the very single interventionalist and center.

We would like to thank reviewer # 1 for the recommended changes and the appreciation of our manuscript.

Title:
Avoid MitraClip in the title => percutaneous edge to edge repair; you speak also about Pascal system in the text. It could be very useful to add peculiarities of echo guidance also for this novel device

The term MitraClip has been replaced with “mitral valve repair”.

A simplified stepwise approach to echo guidance during the percutaneous mitral valve repair.

Abstract:
check for typos; add also a sentence on shared standardization between imager and interventionalist=> useful for safety and efficacy.

We checked for typos and corrected possible mistakes. In the conclusion of the abstract and the discussion we mentioned that shared standardization between imager and interventionalist is indispensable for a safe and effective procedure.

A structured approach improves the collaboration between interventionist and imager and is indispensable for a safe and effective procedure.

Both, the imager and the interventionist should have a fundamental understanding of echocardiographic assessment to achieve a successful and safe procedure.

CW is not colour...but continuous

Thank you very much! We corrected the abbreviation CW.

There are official recommendations: Guidelines for the Evaluation of Valvular Regurgitation After Percutaneous Valve Repair or Replacement ASE https://doi.org/10.1016/j.echo.2019.01.003

As far as we know, there are no official recommendation for the evaluation of valvular quantification during the procedure of percutaneous valve repair. The official recommendations from ASE are allowed for the evaluation of valvular quantification after percutaneous valve repair. We quoted the Guidelines for the Evaluation of Valvular Regurgitation After Percutaneous Valve Repair or Replacement ASE in the text, when appropriate.

Protocol
Add more frequently/underline the role of X plane (es. for check perpendicularity)

We mentioned many times the role of X-plane without writing “X-plane” e.g when writing “choose a intercommissural view and a long axis view to ensure..” or “an intercommissural view combined with a long axis view to show perpendicular positioning....”.

We’ve added the role of X-plane in point 1.9.

Finally, take a bicaval view (90-110°) with X-plane, to show the aortic valve, for the transeptal puncture.

1. Measurement before intervention => better "evaluation";

We changed point 1. in

Evaluation before intervention

1.1 / 1.2 Also evaluation of pericardial effusion and PV flow pattern (not necessarily measure)

We think that it might be too many “evaluate” in series. We left “exclusion of the pericardial effusion” but changed the measurement in pulmonary vein flow in

Evaluate the pulmonary venous flow pattern in the left upper pulmonary vein…

1.4 hemodynamic state should not necessarily be kept during the procedure but at least at basal and final evaluation

We moved the changed phrase from 1.4 to point 1.3

Ensure hemodynamic state is the same at pre- and post-procedural evaluation.

Be consistent during protocol: always use for example keep or sentences without the verb

As recommended by the editor in chief of JOVE We changed the protocol into an imperative tone, which limited the number of verbs.

1.5 contraindication is MV area!!not gradient that coudl be influenced by MR;

As a transmitral pressure gradient > 5 mmHg is associated with poorer prognosis after percutaneous mitral valve repair, literature define this condition as a relative contraindication. We added the adjective “relative”.

A mean pressure gradient (MPG) > 5 mmHg is a relative contraindication for percutaneous mitral valve repair.

1.5 mean gradient should be measured by CW not PW!

PW can be alternatively used to CW. We changed the sentence into…

Assess the transmitral pressure gradient with CW in the long axis view (120-140°)

1.6 it is also possible (however less elegant) to directly measure it on the 3d volume rendering

We complemented the 3D-zoom-mode.

Take a 3D-colour doppler dataset or a wide-sector zoom image with colour and measure the 3D-vena contracta (3D-VCA).

1.7 add absolute and relative contraindication according to MV area

We changed the sentence into

An area < 4cm² is a relative contraindication, an area < 3 cm² an absolute contraindication to the procedure.

1.8 however it is not the some for everybody! always state where aorta is => allign the view in accordance to your interventionist and use the same perspective as much as you can during the procedure to avoid confusion

We mentioned both views; the en-face surgical view and the “modified” en-face view, specifying the localisation of the aortic valve in both views.

Show the 3D en-face surgical atrial view (aortic valve at 12 o’clock) of the mitral valve. NOTE: The segments of the valve are named “lateral” for the segments 1 and “medial” for the segments 3. The sequence of the segments in the en-face surgical view, is inverse to the sequence in the commissural view. Perform a 180° clockwise rotation into the 3D en-face surgical view (aortic valve at 6 o’clock), that will result in equal sequence of segments in both views.

9. state that is for transeptal puncture

We changed point 1.9 in

Finally, take a bicaval view (90-110°) with X-plane for the transeptal puncture.

Clip strategy

To reduce the word “clip”, we changed point 2. In

2. Strategy

2. in front => above!

We changed the expression “in front” into “above”

“2.1.1…position the clip directly above the regurgitation jet..”

4. MitraClip Generation 4 introduced two additional clip sizes NTW and XTW (wider versions of NT and XT) => so 4 sizes are now available; add also a sentence on specificities about clip width

To make the article as neutral as possible, we decided to describe just the strategy and not the different types of the MitraClip devices.

2.1. Discuss the strategy with the operator before inserting the steerable guide catheter (SGC) and the clip delivery system (CDS) into the left atrium.

2.1.1. Evaluate a one device strategy if the orifice is < 1 cm wide and position the clip directly above the regurgitation jet, if the orifice is circular.

2.1.2. Evaluate the implantation of ≥ 2 clips in case of large elliptic or multiple jets. In this case, implant the device starting medially of the regurgitant orifice, as the positioning a second device is often easier than after starting laterally.

Transeptal puncture

State that SAX view at the base is useful for live monitoring to avoid aortic injury

We added information on the position of the aorta during transseptal puncture to avoid it´s injury.

3.1. Ensure the aortic valve is visible, to avoid aortic injury.

Advancement of CDS in the left atrium

one important point to be stressed is perpendicularity and trajectory

We changed the phrase into …

5.2. Check that the CDS is positioned perpendicularly to the coaptation line to guarantee a correct trajectory.

Orientation of the clip

1.1 an INTER-commissural view combined with a long axis view is required: role of X plane; specify how to check orientation in 2D (arms visible only in LAX)

Specification of orientation of the device in 2D was added to the paper:

6.1.1. In the event of poor image quality, show an intercommissural view combined with a long axis view.

6.1.2. Adjust the intercommissural view angle in medially and laterally positioned devices to visualise the complete length of both arms. Sweep at ca. 30-45° for medially positioned devices and ca. 70-90° for laterally positioned devices.

Point 1.1 has been changes in point 6.1.1. We completed point 6.1.1 with:

NOTE: the clip arms are visible only in the LAX.

3. also fluoroscopy could be useful for this purpose

That is right. But we want to describe the steps that are relevant for the echocardiographer. Nevertheless, we want to emphasise the role of fluoroscopy when echocardiography imaging is insufficient, as in point 9.2.

3.1 after gain reduction

Point 3.1. has been changed in 6.4. This depends on the adjustment of the echomachine at the beginning of the procedure and differs between the examiners.

3.2 if gross re-orientation is necessary, clip must be retrieved into LA

We totally agree with this statement and added this information to the protocol. Point 3.2. has been changed in 6.5.

If gross reorientation of the device is necessary, show the intercommissural view with X-plane to visualize the inversion of the clip that will be retrieved in the LA.

Grasping
3. detachment => attachment

What we ment, is not the single leaflet clip attachment, but a grasping of both leaflets resulting/ leading to a partial clip detachment.

4.2 not necessarily measure

We changed point 4.2. in 7.4.2.

Evaluate the PW flow in the pulmonary veins.

4.3 mean G > 5 is not necessarily a contraindication, as gradients are flow dependent => stress role of MV area

Point 4.3. has been changed in 7.4.3. We changed the sentence into…

A gradient > 5 mmHg is a relative contraindication for clip deployment.

4.4 or TG SAX of MV

Point 4.4. has been changed in 7.4.4.

Use the 3D en -face view of the mitral valve or a transgastric short axis view of the mitral valve to show the double orifice.

4.5 scanning the valve with X plane; rule out intraclip jet
Check should be done after clip release! possible role of tension of the system

We described how to rule out intraclip jets in detail in point 7.4.1. and in the “Final MR Assessment” in point 9. The Check should be done before and after clips release. We mentioned the Check after clip release in point 7.5.

Rotate the TEE probe medially and laterally to the clip with colour doppler or use X-plane to find eccentrics jets close to the clip. NOTE: Due to shadowing artifacts caused by the CDS underestimation of the MR may occur. To avoid this, Insert the probe deeper into the esophagus or show the transgastric view to visualize the residual insufficiency jets without shadowing artifacts.

(In “final MR assessment”): Show the commissural view with colour doppler in combination with perpendicular X-planes in the residual insufficiency jets if present.

We added the role of the tension of the system in point 7.5.

NOTE: due to the tension of the system on the MV, the residual insufficiency jets after releasing the device may be aggravated.

Implantation of additional clips

deferred? what does it mean?

Point 8. Has been changed to Point 9.0. If a significant reduction of MR has not been achieved, implantation of additional clips is/may be required (if possible).

1.3 it is not clear => if necessary, reposition of the clip?

Point 1.3 has been changed to point 9.4. What we meant was that repositioning of the clip may be necessary in case of an aggravated MR. To avoid confusion, we changed the phrase in:

Repeat the five steps as explained in point 7.4 to evaluate MR after grasping the leaflets with the additional clip.

Final MR assessment

Stress role of technical settings for VCA
role of hemodynamic monitoring: LA pressure and systemic arterial pressure

Point 9.0 has been changed in point 8.0. We added a NOTE in point 8.3

NOTE: continuous LA pressure monitoring may be a useful tool during transcatheter mitral valve repair.

line 272: also clip G4 has the possibility of independent grasping

The information on the possibility of independent grasping of the 4th generation Clip was rephrased in the paper.

“It can independently grasp leaflets, as the fourth generation MitraClip, in the presence of a central spacer.”

line 279: A critical appraisal of echocardiographic assessment of mitral regurgitation, especially functional one, remains.

We changed the phrase as recommended.The nativ speaker changed the phrase in

A critical appraisal of echocardiographic assessment of MR, especially a functional one, is still to be done.

lines 290.. add role of hemodynamic data

We added:

Different studies demonstrate the value of real-time monitoring of LA pressure during percutaneous mitral valve repair. An increase in left atrial mean pressure is a predictive of worse clinical outcomes at short-term follow-up, independent from echocardiographic findings. Continuous hemodynamic monitoring may complement transesophageal echocardiography to assess and optimize percutaneous edge-to-edge mitral valve repair. 

Figure 3: x plane view starting from..

We changed all titles of the figures as recommended from the Editorial team.

Figure 3 2D biplanar view of the MV with colour doppler: medial insufficiency jet

Figure 6: multiplanar reformatting

Figure 6 Multiplanar reconstruction of the 3D dataset: mitral valve area in 3D

Figure 20: PW is not correct

We choosed another figure for this title

line 338 not only perforation => laceration

We added

Furthermore, repeated grasping of the leaflets can lead to leaflet laceration and perforation.

lines 343-348 add this part to the previous section when you speak about Pascal

We moved the paragraph to the result section where we speak about the Pascal system (line 250)

Since 2019, another device for transcatheter mitral valve repair is available. Compared to the MitraClip device, the PASCAL system has wider paddles. It can independently grasp leaflets, as the generation 4 MitraClip, in the presence of a central spacer. After a multicentre, prospective, observational, first-in-man study and the multicenter early-feasibility trial of the PASCAL transcatheter mitral repair device (CLASP study), the PASCAL repair system received CE mark approval for the treatment of both primary and secondary MR. The early feasibility study of the PASCAL repair system showed high survival, low complication rates as well as improvements in functional status and quality of life. Furthermore, the CLASP study showed a significant reduction of the MR with this device. As there are no head-to-head studies available yet, the preferred device related to the valve anatomy (e.g., larger flail segments, large prolapse, cleft, short posterior leaflet, mitral annular calcification) is still not defined. A head-to-head comparison trial of the MitraClip with the Pascal System (CLASP IID trial) started in 2018 and the estimated primary completion date is expected to be in 2023.

line 352: not only surgically but also percutaneously

We added

In this case, mitral ring dilation can be sutured surgically or percutaneously in place to provide annular stabilization.

line 362: cite guidelines

We added the citation of the current guidelines.

lines 374-376 => it is better to add this part as a conclusion

The structure of the article given by Jove does not contain a conclusion, but we added a conclusive sentence at the end of the discussion.

In conclusion percutaneous transcatheter mitral valve repair is fully dependent on transesophageal echocardiography. Both, the imager and the interventionist should have a fundamental understanding of echocardiographic assessment to achieve a successful and safe procedure.

Check for typos:

echo guidance,

oe/esophageal, manoeuvre..

british vs american english,

incorrect comas,

abbreviations should be consistent: es MR not always used; MC line 97 never explained

line 384: there is a bracket

The manuscript has been proofread from a native speaker.

Reviewer #2:

Manuscript Summary:

The below manuscript titled "A simplified stepwise approach to echoguidance during the MitraClip procedure", is a well written article. After careful review of the manuscript multiple times, I could not find any major flaws in the manuscript.

Major Concerns:

None upon multiple readings

Thank you very much. We are happy to hear this.

Minor Concerns:

page 2- line 69- please change the words to "..procedure are dependent on the individual operator's skill and experience. In c..."

We changed the phrase:

Duration and success of the percutaneous mitral valve repair are dependent on the individual operator’s skill and experience.

Reviewer #3:

The manuscript is well written and gives a brief and precise overview of the imaging strategy for MitraClipping. This topic is very important for interventional cardiologists. I have some minor comments:

Please acknowledge that many centers use the "cardiologist" view which shows the aortic valve at 6 o`clock.

We added this information in the discussion, line 353:

Although the 3D en-face view is defined showing the aortic valve at 12 o’clock, many centers prefer the “cardiologist” 3D en-face view, which shows the aortic valve at 6 o’clock. There is no literature comparing both views during the percutaneous mitral valve repair.

Please add to this section (2.4) accroding to MitraClip G4 device which is - to my knowledge - now available throughout Europe.

Jove’ Editor demanded that we remove any commercial term. To make the article as neutral as possible, we decided to describe just the strategy and not the different types of the MitraClip devices.

2.1. Discuss the strategy with the operator before inserting the steerable guide catheter (SGC) and the clip delivery system (CDS) into the left atrium.

2.1.1. Evaluate a one device strategy if the orifice is < 1 cm wide and position the clip directly above the regurgitation jet, if the orifice is circular.

2.1.2. Evaluate the implantation of ≥ 2 clips in case of large elliptic or multiple jets. In this case, implant the device starting medially of the regurgitant orifice, as the positioning a second device is often easier than after starting laterally.

3.1.2. puncture height of >3.5cm is neccesarry instead of ca. 3.5

Point 3.2.1. We changed the phrase in

Choose a puncture height of 4-5 cm in case of degenerative MR (e.g. prolapse) and of > 3,5 in functional MR. Avoid a patent foramen ovale, as the entry is far too anterior.

I do not agree with a strict superiorly transseptal puncture. I would agree with "slightly superior".

Point 3.2. We changed the phrase in

Ensure the aortic valve is visible, to avoid aortic injury. Make sure the puncture site is slightly superiorly and posteriorly.

The resolution of the figures seems very low and especially 3D enface images are not nicely seen. Sometimes there may be a quality loss due to the editorial system. Please check and improve prior to publication.

Thank you for mentioning this, but the lower quality maybe due to the editorial system; the editors have accepted the quality of the mentioned images. Nevertheless, we improved image quality of every image.

Changes to be made by the Author(s) regarding the video:

1. Please increase the homogeneity between the video and the written manuscript. Ideally, all figures in the video would appear in the written manuscript and vice versa. The video and the written manuscript should be reflections of each other. Currently the results figures are missing from the text.

We complemented the figures of the results in the manuscript.

2. Furthermore, please revise the narration to be more homogenous with the written manuscript. Ideally, the narration is a word for word reading of the written protocol.

As we rewrote the text in the imperative tense, we also changed the narration.

3. Please ensure that the section highlight is same as that in the text.

We complemented the sections from the text.

4. Please include an ethics statement before the numbered protocol steps, indicating that the protocol follows the guidelines of your institution’s human research ethics committee.

We added the ethic statement at the beginning of the protocol.

5. 2:00: Please zoom the monitor to not show the commercial name throughout the video.

The commercial name is not visible in the current version of the video.

6. Representative results: Please do not use commercial terms.

We minimized the commercial terms as much as possible.

7. Text

All the title looks more like a sentence. Please make each word have a capital letter to make it look more like a title: 00:00, 00:51, 03:06, 06:36
Please add a person caption with their name and affiliated lab to the left of the person: 00:11, 08:58, 09:12

The titles have been corrected and a person caption has been added as requested.

8. Pacing

00:09 - 00:09 The title doesn't need to be 10 seconds long. Please reduce this to about 5 seconds.

We reduced the length of the titles as requested.

00:34 - 00:34 There's a weird fade/dip to black that happens here that's very jarring. Please use a straight cut without any dissolve or flash/dip.

We made the requested changes.

00:55 - 00:55 The audio comes in a little late. Please have the narration happen right after the video fades up.

The narration starts right after the video fades up.

The narration ends and we hold on this shot for too long. Please move onto the next narration much sooner within these three spots: 01:26 - 01:31, 02:05 - 02:05, 02:45 - 02:53.

9. Audio

Interview audio should come out of both channels, instead of just the left. Please balance the audio so it can be heard equally left and right: 00:10 - 00:10, 08:56 - 08:56
00:57 - 00:57 Please remove any unnecessary sounds like breathing or paper rustling. I can here lip smacking here.

We made the requested changes.

I can hear background noise whenever certain clips are shown. Make sure the audio from any of the clips being shown are removed from the timeline in your editing program. We only want to hear the narration and nothing else please: 01:12 - 01:25, 01:50 - 02:04, 02:28 - 02:44, 03:52 - 03:59, 04:55 - 05:46, 05:55 - 06:10, 06:22 - 06:33

We deleted the background noise as requested.

10. Composition

00:35 - 00:35 The backdrop looks to be greenscreen. If so, I would recommend using a different backdrop to separate this interview from the previous. Maybe flipping the image so the person is on the left side of the screen would help as well.

The backdrop is not a greenscreen, but the official background requested from our hospital for portraits and interviews.