TITLE:

Use of a percutaneous ventricular assist device/left atrium to femoral artery bypass system for cardiogenic shock

AUTHORS AND AFFILIATIONS:

Swethika Sundaravel1, Poonam Velagapudi1, Mamas Mamas2, Sandeep Nathan3, Alexander Truesdell4

- University of Nebraska Medical Center, Omaha, NE

- Keele Cardiovascular Research Group, Keele University, UK

- University of Chicago Medicine, Chicago, IL, USA

- Virginia Heart/Inova Heart and Vascular Institute, Falls Church, VA, USA

s.sundaravel@unmc.edu

poonamchou@gmail.com

mamasmamas1@yahoo.co.uk

snathan@medicine.bsd.uchicago.edu

agtruesdell@gmail.com

KEYWORDS:

Mechanical circulatory support, cardiogenic shock, percutaneous ventricular assist device, left atrial to femoral artery bypass system, transseptal puncture, right atrium to pulmonary artery bypass, Protek Duo, acute myocardial infarction

SUMMARY:

The following article describes the stepwise procedure for placement of a device (e.g., Tandemheart) in cardiogenic shock (CS) that is a percutaneous left ventricular assist device (pLVAD) and a left atrial to femoral artery bypass (LAFAB) system that bypasses and supports the left ventricle (LV) in CS.

ABSTRACT:

The left atrial to femoral artery bypass (LAFAB) system is a mechanical circulatory support (MCS) device used in cardiogenic shock that bypasses the left ventricle by draining blood from the left atrium and returning it to the systemic arterial circulation via the femoral artery. It can provide flows ranging from 2.5-5 L/min depending on the size of the cannula. Here, we discuss the mechanism of action of LAFAB, available clinical data, indications for its use in cardiogenic shock, steps of implantation, post-procedural care, complications associated with the use of this device and their management, and finally the advantages of this device over other MCS devices.

We also provide a brief video of the procedural component of device therapy, including the pre-placement preparation, percutaneous placement of the device via trans-septal puncture under echocardiographic guidance and the post-operative management of device parameters and clinical course.

INTRODUCTION:

Cardiogenic shock (CS) is a state of tissue hypoperfusion with or without concomitant hypotension, in which the heart is unable to deliver sufficient blood and oxygen to meet the body’s demands, resulting in organ failure. It is classified into stages A to E by the Society of Cardiovascular Angiography and Interventions (SCAI): stage A - patients at risk for CS; stage B - patients at beginning stage of CS with hypotension or tachycardia without hypoperfusion; stage C - classic CS with cold and wet phenotype requiring inotropes/vasopressors or mechanical support to maintain perfusion; stage D - deteriorating on current medical or mechanical support requiring escalation to more advanced devices; and stage E - includes patients with circulatory collapse and refractory arrhythmias who are actively experiencing cardiac arrest with ongoing cardiopulmonary resuscitation1. The most common causes of CS are acute MI (AMI) representing 81% of cases in a recently reported analysis2, and acute decompensated heart failure (ADHF). Cardiogenic shock is classically characterized by congestion and impaired perfusion, manifested by elevated filling pressures (pulmonary capillary wedge pressure [PCWP], left ventricular end-diastolic pressure, central venous pressure, and right ventricular end-diastolic pressure [LVEDP]) and poor CO (cardiac index [CI], cardiac power output, and end-organ malfunction)3. In the past, the only available treatments for AMI complicated by CS were early revascularization and medical management with inotropes and/or vasopressors4. More recently, with the advent of MCS devices and the recognition that escalation of vasopressors is associated with increased mortality, there has been a paradigm shift in the treatment of both AMI and ADHF related CS5 6.

In the current era of percutaneous ventricular assist devices (pVAD), there are a number of MCS device platforms/configurations available, which provide univentricular or biventricular circulatory and ventricular support with and without oxygenation capability7. Despite steady increases in the use of pVADs to treat both AMI and ADHF CS, mortality rates have remained largely unchanged5. With emerging evidence for possible clinical benefits to early unloading of the left ventricle in acute MI8 and early use of MCS in AMI CS9, the use of MCS continues to increase.

The Left Atrial to Femoral Artery Bypass (LAFAB) MCS device bypasses the left ventricle by draining blood from the left atrium and returning it to the systemic arterial circulation via the femoral artery (Figure 1). It is supported by an external centrifugal pump that offers 2.5-5.0 liters per minute (L/m) flow (new generation pump, designated as LifeSPARC, capable of up to 8 L/m flow) depending on the size of the cannulas. Once the blood is extracted from the left atrium via the transseptal venous cannula, it passes through the external centrifugal pump which recirculates the blood back into the patient’s body via the arterial cannula placed in the femoral artery.

[FIGURE 1 HERE]

PROTOCOL:

This procedure and protocol have been approved by the institutional review board and FDA.

- Patient criteria

  - Include patients with cardiogenic shock stage B and above as defined by the SCAI consensus statement1.

  - Include as bridge to transplant or durable left ventricular assist system in stage D heart failure.

  - Include as bridge to recovery in acute myocardial infarction complicated or cardiogenic shock.

  - Exclude contra-indication to systemic anticoagulation.

  - Exclude life-expectancy <6months (active malignancy).

  - Exclude if there is the presence of left atrial thrombus.

  - Exclude if patients have peripheral vascular disease with small arteries that cannot house the large cannulas.

  - Exclude if patients have an irreversible neurological injury/coma.

  - Exclude if patients have a severe aortic insufficiency.

  - Exclude if patients have a ventricular septal defect.

NOTE: The placement of the TH device includes three separate processes: 1) setting up the controller and the pump; 2) placement of arterial and venous cannulas and transeptal access under transesophageal echocardiogram (TEE) or intracardiac echocardiogram (ICE); and 3) connecting the system to the circuit.

- Placement of the left atrium to femoral artery bypass device

  - Setting up the controller

NOTE: This step can be performed as the patient is being transported to the lab and the table is being set up for the procedure. Usually, the device representative and perfusionist team is present to assist with the process.

    - Open the box and initiate the controller setup with the following steps.

      - Turn on power in the back of the controller and allow the system to power up.

      - Select either View Instructions (for a step by step tutorial) or Skip Instructions on the screen.

    - Remove the intravenous tubing set from the sterile package. Priming the infusate is a critical step in the whole process.

    - Plug in the pressure transducer cable and wait for it to calibrate. Hang a bag of saline by the holder at the controller.

    - Turn the stopcock off to the saline spike and then connect the saline spike to the bag of saline. Fill the drip chamber two-thirds with saline.

    - Prime the lower housing of the circuit. Connect a 20 mL syringe to the stopcock and fill 20 mL from the saline bag.

    - Open the stopcock to the transducer to expel any air between the transducer and the filter.

  - Setting up the pump

NOTE: Pump priming requires two people — one person is scrubbed and stays in the sterile field. The other person handles the controller in the non-sterile field. Pump de-airing is a crucial step and needs to be performed very carefully.

    - Have the primary operator from the sterile field hand the ends of the pump’s external communicating line and infusate line to the secondary operator managing the controller (in the unsterile field).

    - Have the secondary operator connect the infusate line to the bacteriologic filter at the end of the infusion assembly and connect the power line to the controller.

    - Have the controller operator refill the 20 mL syringe from the bag and slowly inject 15 mL of heparinized saline into the pump’s infusion line to prime the lower housing of the pump. Then refill the syringe.

    - Plug in the power cable and connect it to the infusate line (from the scrub table). Inject 15 mL of heparinized saline into the pump’s lower chamber.

    - Start the pump and inject additional saline with the pump running. Then turn off the pump and inspect the lower chamber for any air bubbles (the pump will shut down automatically if it is on screen three of the “View Instructions Mode”). Repeat the above steps if air bubbles are detected.

    - From the scrub table, inject 60 mL of saline into the pump’s upper chamber. Refill the syringe and inject an additional 60 mL of saline into the pump tubing.

    - Tap the pump and the tubing to ensure that all the air bubbles are removed.

    - Clamp the tubings below the fluid line.

    - Place the infusate tubing into the infusion pump and the air detector on the side of the controller.

  - Transseptal access10,32

    - Prepare and drape the patient in a sterile manner.

    - Perform the procedure under general anesthesia with the anesthesia team.

    - Once the patient is intubated and sedated adequately, pass the TEE probe into the esophagus and obtain the basic images. If using intracardiac echocardiogram (ICE), then the images can be obtained after venous access.

    - Identify the ideal spot for septostomy on the inter-atrial septum using TEE (Figure 2) or ICE (Figure 3). Use the bicaval view on TEE showing the membranous part of the inter-atrial septum at the region of the fossa ovalis to expose the inter-atrial septum better.

    - Confirm the absence of any thrombus in the left atrium where the inflow cannula will be positioned using TEE or ICE.

[FIGURE 2/3 HERE]

    - Obtain femoral venous access via ultrasound guidance with modified Seldinger technique and insert an 0.035” guidewire.

    - Advance the guidewire to the inferior vena cava – right atrial junction and then direct it towards the inter-atrial septum under fluoroscopic and TEE or ICE guidance. Using multiple angiographic projections (Right or Left Anterior Oblique), identify the optimal site for transeptal puncture (Figure 4).

NOTE: Ideally, it should be done in the region of the foramen ovale10 to minimize complications. In patients with thick or aneurysmal septa or an inter-atrial septum which has been previously patched or instrumented surgically or closed percutaneously, one may consider use of an energized or radiofrequency transseptal needle to insure precise puncture without needle deflection.

[FIGURE 4 HERE]

    - Anticoagulate the patient (ACT more than 250 seconds). Perform transseptal puncture using a trans-septal needle, insert guidewire to the left atrium (LA).

    - Dilate the venous access and the interatrial septum with a 2-stage dilator. Insert the transseptal cannula and advance it into LA, remove introducer and guidewire, wait for back-bleed, and clamp. Secure cannula to the patient.

  - Arterial access

    - Obtain femoral arterial access via the modified Seldinger technique using ultrasound and angiographic guidance at the level of the femoral head. Insert an 0.035” guidewire.

    - Consider use of “pre-closure” technique (e.g., Perclose ProGlide Suture Mediated Closure system) prior to upsizing arterial access.

    - Serially dilate the arterial access site appropriate to the size of the arterial cannula selected. Insert the Protek Solo arterial cannula, and remove the introducer and guidewire. Wait for back-bleed, and then clamp. Secure the cannula to the patient using the holders.

  - Connecting the components

NOTE: It is crucial to make wet-to-wet connections to the cannulas to avoid introduction of any air bubbles in the circuit.

    - Use saline immersion or constant infusion of saline (“waterfall”) over the two ends of the cannulas as they are being connected.

    - Connect the transseptal cannula (venous) to the pump inlet which is marked in blue and the arterial cannula to the pump outlet which is marked in red.

    - First, remove the venous clamps and start the pump (from the controller box). Then release the other clamps sequentially, constantly checking for any air, releasing the arterial clamp last.

    - Adjust pump speed (by adjusting the RPMs) to optimize flow. Confirm the position of the cannula under fluoroscopy and TEE or ICE and secure the circuit to patient.

    - Maintain therapeutic anticoagulation (activated clotting time ACT at 180–220 s or activated partial thromboplastin time aPTT at 65–80 s) for as long as the pump is in place to prevent pump thrombosis and stroke.

- Right Atrium to Pulmonary Artery Bypass (RAPAB) system placement

  - Initiate the controller, prime the lower chamber of the pump and check for any air bubbles – same steps described as above.

  - Then prime the upper chamber of the pump, check for any air and clamp it – same steps described as above.

  - Patient procedure

    - Obtain venous access at the right internal jugular (RIJ) vein via modified Seldinger technique under ultrasound guidance.

    - Insert a Pulmonary Artery (PA) Catheter with a 0.035 in lumen and advance it to the main PA just before the bifurcation. Insert a stiff-bodied 0.035 in guidewire and remove the PA Catheter.

    - Anticoagulate the patient (ACT > 250 s).

    - Dilate the venous access site sequentially using the stepwise dilators provided in the package until the desired size is reached (29 French or 31 French).

    - Insert the venous cannula (e.g., ProtekDuo) over the guidewire.

    - Remove the guidewire, wait for back-bleed, and then clamp the distal port which is marked as “Distal.”

    - Remove the hemostasis cap, wait for back-bleed, and clamp the proximal port which is marked as “Proximal.” Secure the cannula to patient via sutures.

    - Verify the cannulas and make wet-to-wet connections from the pump to the cannula.

    - Connect the proximal cannula to the pump inlet which is marked in blue and the distal cannula to the pump outlet which is marked in red. Turn on the pump (from the controller).

    - Release clamps sequentially, constantly checking for any air bubbles. Adjust pump speed (by controlling the RPMs) to optimize flow.

    - Confirm cannula position under fluoroscopy (may use TEE guidance to confirm position in the main PA) and secure the circuit to patient. Maintain therapeutic anticoagulation (ACT at 180–220 s; aPTT at 65–80 s).

  - Device removal

NOTE: Once the patient’s end organ function has improved and hemodynamics have stayed stable with either left ventricular recovery or advanced therapies such as durable LVAD placement/transplant, the device can be removed.

    - Before removing the device, slowly turn down the speed by 0.5 L/min in a stepwise manner, carefully observing the hemodynamics to make sure there is adequate cardiac output and normal filling pressures with less support (also known as the ramp down or turn down study).

    - Turn off the pump once the turndown study is successful.

    - Make sure the ACT is < 150 seconds before removing the arterial cannula. Tighten the previously placed intravascular suture to occlude the arteriotomy site or manual pressure can be held for at least 40 minutes to ensure hemostasis.

    - Withdraw the trans-septal cannula into the IVC and remove it slowly from the femoral vein. Apply a figure of 8 suture to the venous site and additionally, hold manual pressure over the site to achieve hemostasis.

NOTE: The atrial septal defect is usually small and is not closed routinely.

    - Post procedure, perform continuous monitoring of patient’s hemodynamics and end organ function to ensure stability.

    - Removal of RAPAB

    - Similar to the LAFAB removal, when the patient’s end-organ function has stabilized with recovery or advanced therapies, slowly turn down the pump by 0.5 L/min, carefully observing the hemodynamics.

    - Turn off the pump when turn-down study is successful.

    - Once ACT is < 150 seconds, remove the venous cannula from the neck and place a figure of 8 suture to secure the puncture site. Hold manual pressure in addition to the suture to achieve complete hemostasis.

    - Post removal, closely monitor patient’s hemodynamics and end-organ function to ensure stability.

TABLE 1 HERE

REPRESENTATIVE RESULTS:

Clinical applications of LAFAB device

The technique and feasibility of a percutaneous trans-atrial left ventricular bypass system were first described in the 1960s by Dennis et al.11,12. However trans-septal puncture was not initially widely adopted due to complications with the septostomy technique. Over the last decade, with advancements in the field of percutaneous interventions, operators have accumulated experience with atrial septostomy (most commonly also used for percutaneous mitral valve interventions), which has led to a resurgence of the trans-septal ventricular assist device or LAFAB device.

From the initial clinical studies conducted in the 1990s, the LAFAB strategy has demonstrated a high level of myocardial preservation with larger decrease in infarct size from, and greater LV unloading in AMI CS, as compared to the intra-aortic balloon pump (IABP)13.

LAFAB device for cardiogenic shock

The LAFAB circuit is most widely used in the setting of CS, where it has been shown to be safe and effective in augmenting CO14. A 2006 randomized trial comparing the LAFAB device to IABP in CS15 in 42 patients demonstrated that the LAFAB device was more effective in lowering PCWP and improving CI (1.2 ± 0.8 (P < .05 vs baseline), although mortality did not differ between the two groups. The incidence of severe adverse events was likewise not significantly different between the two groups in this small cohort. Another trial comparing IABP and LAFAB showed significant improvement in cardiac power index with the LAFAB device (0.22 to 0.37 W/m2, p<0.001), but with a trend towards more complications and no mortality benefit in the LAFAB group16. More patients in the LAFAB group experienced significantly higher acute limb ischemia necessitating revascularization (n=7 vs. n=0, P=0.009) and bleeding complications (n=19 vs. n=8, P=0.002). There was also a trend towards increased incidence of disseminated intravascular coagulopathy in the LAFAB group.

Table 2 here

A 10-patient report also demonstrated successful LAFAB device use in patients with severe aortic stenosis and CS. In this case series, the LAFAB device was used as a temporizing measure to stabilize hemodynamics and end-organ function prior to definitive valve replacement (8 patients who received it in the cath lab before surgical valve replacement had improved renal function prior to surgery, 2 patients received the device support in the operating room after valve replacement and both of them died. Overall, 3/10 patients died)17.

More recently, in 2011, Kar et al. published a study of 117 patients who underwent LAFAB implantation for severe CS refractory to IABP and vasopressor therapy18. 80 patients had ischemic cardiomyopathy, and 37 had non-ischemic cardiomyopathy. There was an immediate improvement in hemodynamics (as evidenced by improved CI – 0.53 l/ (min·m2) to 3.0 l/ (min·m2), p<0.001 and PCWP-31.53 ± 10.2 mm Hg to 17.29 ± 10.82 mm Hg (p < 0.001)) as well as end-organ function in both groups. Mortality rates at 1 and 6 months were 40.2% and 45.3% respectively.

LAFAB device as a bridge to transplant/durable VAD

Another important application of pVADs is in end-stage heart failure, where they are used as a bridge to definitive therapy – durable left ventricular assist device (LVAD) or cardiac transplantation19.

The LAFAB device has been shown to be effective in this scenario as well. In a series of 25 patients with CS20, 44% of patients underwent placement of durable left ventricular assist device (LVAD), 30% recovered, and 36% died on support. The mean duration of LAFAB support was 4.8 ± 2.1 days. 56% of patients experienced device-related complications, of which 90% were secondary to vascular access complications.

Gregoric et al. reported a series of 9 patients with end-stage cardiomyopathy and refractory shock who were bridged with the LAFAB device to a durable LVAD. Eight of these nine patients were supported with IABP prior to being upgraded to LAFAB. Mean duration to durable LVAD was 5.9 days. All 9 patients showed improvement in their hemodynamics and end-organ function on LAFAB support before undergoing LVAD implantation21.

In another series of 5 critically ill patients with CS22, the LAFAB device was placed emergently for left ventricular support as a bridge to transplantation. The mean duration of support was 7.6 ± 3.2 days, and all five patients underwent cardiac transplantation successfully. No device-related complications were reported in this cohort.

Agarwal et al. in 2015 also described a rare, interesting case of durable LVAD thrombosis with CS treated successfully with fibrinolytic therapy and temporary support with LAFAB23.

LAFAB device for high-risk percutaneous interventions

Another clinical application for LAFAB use is in the setting of high-risk percutaneous coronary intervention (PCI)24. The LAFAB device has been studied and utilized extensively in this setting25,26,27,28 with device-related vascular complications within acceptable ranges compared to other comparable devices. In a recent series of 37 patients who underwent LAFAB device placement for CS, 28 patients also underwent high-risk PCI successfully while being supported with the LAFAB device. 71% of patients in the cohort were successfully discharged home. The mean EuroSCORE was 11 ± 3.4, indicating a high complexity of this critically ill cohort. The authors concluded that the LAFAB device could be used safely and effectively for high-risk PCI29. In addition to high-risk PCI, the LAFAB device has also been used for high risk patients undergoing percutaneous aortic valve replacement, as demonstrated in one small series30. Finally, there is also a small 2-patient report of successful LAFAB device use for post-cardiotomy shock31.

FIGURE AND TABLE LEGENDS:

Figure 1: LAFAB setup. Image courtesy of TandemLife, a wholly owned subsidiary of LivaNova US Inc.

Figure 2: TEE with biplane in the bicaval view showing the inter-atrial septum. LA – Left Atrium, RA – Right Atrium, IVC – Inferior Vena Cava, SVC – Superior Vena Cava, Ao – Aorta, PV – Pulmonary Vein, RV – Right Ventricle. Picture courtesy Alkhouli M, Rihal CS, Holmes DR. Transseptal Techniques for Emerging Structural Heart Interventions. JACC: Cardiovascular Interventions. 2016;9 (24):2465-8010

Figure 3: ICE for transeptal access (A) The fossa ovalis (FO) as seen in the “septal” view. (B) Tenting of the FO with the transseptal needle. (C) The transseptal sheath across the FO in the Left Atrium (LA). RA – Right Atrium, IVC – Inferior Vena Cava, SVC – Superior Vena Cava. Picture courtesy Alkhouli M, Rihal CS, Holmes DR. Transseptal Techniques for Emerging Structural Heart Interventions. JACC: Cardiovascular Interventions. 2016;9 (24):2465-8010

Figure 4: Multiple angiographic projections to optimize the site for transseptal access. RAO – Right Anterior Oblique. LAO – Left Anterior Oblique. RA – Right Atrium, LA- Left Atrium, RV – Right Ventricle, LV – Left Ventricle, FO – Foramen Ovale, CS – Coronary Sinus, TV - Tricuspid Valve, AV – Aortic Valve, AR – Aortic Root, PA – Pulmonary Artery and Ao – Aorta. Picture courtesy Alkhouli M, Rihal CS, Holmes DR. Transseptal Techniques for Emerging Structural Heart Interventions. JACC: Cardiovascular Interventions. 2016;9 (24):2465-8010

Figure 5: Schematic of ProTek Duo system. Image courtesy of TandemLife, a wholly owned subsidiary of LivaNova US Inc.

Table 1: Complications of LAFAB device33.

Table 2: Table of results for hemodynamic improvement post LAFAB compared with intra-aortic balloon pump. Adapted from Thiele et al.17.

DISCUSSION:

Hemodynamics of LAFAB device:

The hemodynamic profile of the LAFAB device is distinct from other pVADs. By draining blood directly from the left atrium and returning it to the femoral artery, the device bypasses the LV completely. In doing so, it reduces LV end diastolic volume and pressure, contributing to improved LV geometry thereby effecting a decrease in LV stoke work. However, by returning the blood back into the iliac artery/descending aorta, afterload increases. This results in loading of the LV via elevated left ventricular end-systolic pressure (LVESP). Overall, the pressure-volume loop narrows and shifts modestly to the left34. The result is a reduction in myocardial workload and decreased oxygen consumption35. The LVEDP and LVESP may be optimized by increasing pump speed, further unloading the LV36.

The advantages of the LAFAB device include decompression of the LV and decreased stroke work, leading to myocardial recovery. LAFAB unloads the LV indirectly by venting the left atrium (and does not require a separate device for LV unloading). It can provide cardiac support up to 4.5 L/min at the bedside, and its durability has been well established up to 2 weeks in vivo37.

Disadvantages of LAFAB lie in the advanced procedural skills required for device placement, as trans-septal puncture is not a ubiiquitous cardiac catheterization lab skill. The need for TEE or ICE as well as fluoroscopy limits the use of this device at bedside in emergent situations. Femoral vascular access prevents patients from ambulating adversely affecting patient rehabilitation. As with other MCS, both hemolysis and access site complications may also be significant. Residual atrial septal defect may also occur and adversely affect outcomes in patients bridged to recovery.

Contraindications to LAFAB device:

Absolute contraindications to LAFAB system use include pre-existing ventricular septal defects, severe aortic regurgitation, left atrial thrombus, and severe peripheral vascular disease precluding device insertion38.

Right Atrium to Pulmonary Artery Bypass (RAPAB) for right ventricular failure:

Right ventricular (RV) dominant or biventricular CS is as fatal as left ventricular dominant CS, with in-hospital mortality rates ranging up to 75%39. Various MCS devices have been successfully used in the management of RV CS. Similar to the LAFAB system, there is also a novel right ventricular support device draining blood from the right atrium to the pulmonary artery. It is called the ProtekDuo® system (Figure 5), which can provide up to 4 lpm of flow for the RV40. It consists of a long 31 F (also available in smaller sizes) dual lumen cannula that traverses from the internal jugular vein through the right atrium and right ventricle, into the main pulmonary artery41. The system is supported by a centrifugal pump similar to the LAFAB system. It empties blood from the right atrium and returns it to the main pulmonary artery, bypassing the RV. It may also be connected to an oxygenator for pulmonary support. The RAPAB system helps unload the RV and reduce the RV stroke work42. It requires TEE guidance for confirmation of position and placement in the main pulmonary artery. This device is being used increasingly in RV cardiogenic shock43, massive pulmonary embolism44, and postcardiotomy RV failure after LVAD or transplant45. The bulk of the data for RV support comes from the TandemHeart Experiences and Methods (THEME) registry46, where the RAPAB system was used in 30 patients with RV failure of various etiologies, with a 30-day survival of 72.4%, and most patients bridged to recovery.

[FIGURE 5 HERE]

Use of the LAFAB circuit may provide differential benefit over a transvalvular axial flow pump such as Impella in a number of specific clinical scenarios. As stated previously, the LifeSPARC pump is capable of delivering supraphysiologic levels of flow (up to 8 L/min) with current cannula configurations and thus larger patients and those not adequately supported or fully unloaded on transvalvular axial flow pumps (Impella 2.5 L or CP 4 L) devices may benefit from LAFAB. Secondly, patients with prosthetic aortic valves or unfavorable aortic valve/root anatomy such as those with critical aortic stenosis or aortic valve endocarditis and patients with left ventricular thrombus may be preferentially considered for LAFAB. Finally, LAFAB may be part of a staged support strategy that began with TandemLife VA-ECMO for acute cardiopulmonary rescue followed by subsequent left atrial drainage into the circuit, thus effectively providing near-complete biventricular support plus oxygenation.

The LAFAB device is an effective and safe pVAD that can be used in acute or chronic LV failure providing up to 4.5 to 5 lpm of circulatory support as well as indirect LV unloading via left atrial decompression. The need for trans-septal access may limit the widespread use of this device for emergent LV support in CS. Although observational data suggests both safety and efficacy of the device in both CS and high-risk PCI, well-conducted large, randomized control trials are needed. In the meantime, results of the ongoing THEME registry may shed more light on the real-world outcomes in patients managed with the LAFAB device.

ACKNOWLEDGMENTS:

To the TandemHeart team at LifeSparc.

DISCLOSURES:

Sandeep Nathan – Disclosures: Consultant, Abiomed, Getinge, CSI, Inc.

Alexander Truesdell - Disclosures: Consultant, Abiomed Inc.

REFERENCES:

1.	Baran, D.A. et al. SCAI clinical expert consensus statement on the classification of cardiogenic shock. Catheterization and Cardiovascular Interventions. 94 (1), 29-37 (2019).

2.	Harjola, V.-P. et al. Clinical picture and risk prediction of short-term mortality in cardiogenic shock. European Journal of Heart Failure. 17 (5), 501-509 (2015).

3.	Furer, A., J. Wessler, Burkhoff, D. Hemodynamics of Cardiogenic Shock. Interventional Cardiology Clinics  6 (3), 359-371 (2017).

4.	Hochman, J.S. et al. Cardiogenic shock complicating acute myocardial infarction--etiologies, management and outcome: a report from the SHOCK Trial Registry. SHould we emergently revascularize Occluded Coronaries for cardiogenic shocK? Journal of the American College of Cardiology. 36 (3 Suppl A), 1063-70 (2000).

5.	Shah, M. et al. Trends in mechanical circulatory support use and hospital mortality among patients with acute myocardial infarction and non-infarction related cardiogenic shock in the United States. Clinical Research in Cardiology. 107 (4), 287-303 (2018).

6.	van Diepen, S. et al. Contemporary Management of Cardiogenic Shock: A Scientific Statement From the American Heart Association. Circulation. 136 (16), e232-e268 (2017).

7.	Alkhouli, M. et al. Mechanical Circulatory Support in Patients with Cardiogenic Shock. Current Treatment Options in Cardiovascular Medicine. 22 (2), 4 (2020).

8.	Basir, M.B. et al. Feasibility of early mechanical circulatory support in acute myocardial infarction complicated by cardiogenic shock: The Detroit cardiogenic shock initiative. Catheterization and Cardiovascular Interventions. 91 (3), 454-461 (2018).

9.	Basir, M.B. et al. Improved Outcomes Associated with the use of Shock Protocols: Updates from the National Cardiogenic Shock Initiative. Catheterization and Cardiovascular Interventions. 93 (7), 1173-1183 (2019).

10.	Alkhouli, M., C.S. Rihal, Holmes, D.R. Transseptal Techniques for Emerging Structural Heart Interventions. JACC: Cardiovascular Interventions. 9 (24), 2465-2480 (2016).

11.	Dennis, C. et al. Clinical use of a cannula for left heart bypass without thoracotomy: experimental protection against fibrillation by left heart bypass. Annals of Surgery. 156 (4), 623-637 (1962).

12.	Dennis, C. et al. Left atrial cannulation without thoracotomy for total left heart bypass. Acta Chirurgica Scandinavica. 123, 267-79 (1962).

13.	Fonger, J.D. et al. Enhanced preservation of acutely ischemic myocardium with transseptal left ventricular assist. Annals of Thoracic Surgery. 57 (3), 570-5 (1994).

14.	Thiele, H. et al. Reversal of cardiogenic shock by percutaneous left atrial-to-femoral arterial bypass assistance. Circulation. 104 (24), 2917-22 (2001).

15.	Burkhoff, D. et al. A randomized multicenter clinical study to evaluate the safety and efficacy of the TandemHeart percutaneous ventricular assist device versus conventional therapy with intraaortic balloon pumping for treatment of cardiogenic shock. American Heart Journal. 152 (3), 469.e1-469.e8 (2006).

16.	Thiele, H. et al. Randomized comparison of intra-aortic balloon support with a percutaneous left ventricular assist device in patients with revascularized acute myocardial infarction complicated by cardiogenic shock. European Heart Journal. 26 (13), 1276-1283 (2005).

17.	Gregoric, I.D. et al. TandemHeart as a rescue therapy for patients with critical aortic valve stenosis. Annals of Thoracic Surgery. 88 (6), 1822-6 (2009).

18.	Kar, B. et al. The percutaneous ventricular assist device in severe refractory cardiogenic shock. Journal of the American College of Cardiology. 57 (6), 688-96 (2011).

19.	Patel, C.B., K.M. Alexander, Rogers, J.G. Mechanical Circulatory Support for Advanced Heart Failure. Current Treatment Options in Cardiovascular Medicine. 12 (6), 549-565 (2010).

20.	Tempelhof, M.W. et al. Clinical experience and patient outcomes associated with the TandemHeart percutaneous transseptal assist device among a heterogeneous patient population. Asaio Journal. 57 (4), 254-61 (2011).

21.	Gregoric, I.D. et al. The TandemHeart as a bridge to a long-term axial-flow left ventricular assist device (bridge to bridge). Texas Heart Institute Journal. 35 (2), 125-9 (2008).

22.	Bruckner, B.A. et al. Clinical experience with the TandemHeart percutaneous ventricular assist device as a bridge to cardiac transplantation. Texas Heart Institute Journal. 35 (4), 447-450 (2008).

23.	Agarwal, R. et al. Successful treatment of acute left ventricular assist device thrombosis and cardiogenic shock with intraventricular thrombolysis and a tandem heart. Asaio Journal. 61 (1), 98-101 (2015).

24.	Vetrovec, G.W. Hemodynamic Support Devices for Shock and High-Risk PCI: When and Which One. Current Cardiology Reports. 19 (10), 100 (2017).

25.	Al-Husami, W. et al. Single-center experience with the TandemHeart percutaneous ventricular assist device to support patients undergoing high-risk percutaneous coronary intervention. Journal of Invasive Cardiology. 20 (6), 319-22 (2008).

26.	Vranckx, P. et al. Clinical introduction of the Tandemheart, a percutaneous left ventricular assist device, for circulatory support during high-risk percutaneous coronary intervention. International Journal of Cardiovascular Interventions. 5 (1), 35-9 (2003).

27.	Vranckx, P. et al. The TandemHeart, percutaneous transseptal left ventricular assist device: a safeguard in high-risk percutaneous coronary interventions. The six-year Rotterdam experience. EuroIntervention. 4 (3), 331-7 (2008).

28.	Vranckx, P. et al. Assisted circulation using the TandemHeart during very high-risk PCI of the unprotected left main coronary artery in patients declined for CABG. Catheterization and Cardiovascular Interventions. 74 (2), 302-10 (2009).

29.	Thomas, J.L. et al. Use of a percutaneous left ventricular assist device for high-risk cardiac interventions and cardiogenic shock. Journal of Invasive Cardiology. 22 (8), 360 (2010).

30.	Vranckx, P. et al. Assisted circulation using the Tandemhear , percutaneous transseptal left ventricular assist device, during percutaneous aortic valve implantation: the Rotterdam experience. EuroIntervention. 5 (4), 465-9 (2009).

31.	Pitsis, A.A. et al. Temporary assist device for postcardiotomy cardiac failure. The Annals of Thoracic Surgery. 77 (4), 1431-1433 (2004).

32.	Singh, G.D., Smith, T.W. Rogers, J.H.Targeted Transseptal Access for MitraClip Percutaneous Mitral Valve Repair. Interventional Cardiology Clinics. 5 (1), 55-69 (2016).

33.	Subramaniam, A.V. et al. Complications of Temporary Percutaneous Mechanical Circulatory Support for Cardiogenic Shock: An Appraisal of Contemporary Literature. Cardiology and Therapy. 8 (2), 211-228 (2019).

34.	Morley, D. et al. Hemodynamic effects of partial ventricular support in chronic heart failure: Results of simulation validated with in vivo data. The Journal of Thoracic and Cardiovascular Surgery. 133 (1), 21-28.e4 (2007).

35.	Naidu, S.S., Novel Percutaneous Cardiac Assist Devices. Circulation. 123 (5), 533-543 (2011).

36.	Kapur, N.K. et al. Hemodynamic Effects of Left Atrial or Left Ventricular Cannulation for Acute Circulatory Support in a Bovine Model of Left Heart Injury. ASAIO Journal. 61 (3), 301-306 (2015).

37.	Smith, L. et al. Outcomes of patients with cardiogenic shock treated with TandemHeart percutaneous ventricular assist device: Importance of support indication and definitive therapies as determinants of prognosis. Catheterization and Cardiovascular Interventions. 92 (6), 1173-1181 (2018).

38.	Ergle, K., Parto, P., Krim, S.R. Percutaneous Ventricular Assist Devices: A Novel Approach in the Management of Patients With Acute Cardiogenic Shock. The Ochsner Journal. 16 (3), 243-249 (2016).

39.	Sultan, I., Kilic, A., Kilic, A.Short-Term Circulatory and Right Ventricle Support in Cardiogenic Shock: Extracorporeal Membrane Oxygenation, Tandem Heart, CentriMag, and Impella. Heart Failure Clinics. 14 (4), 579-583 (2018).

40.	Bermudez, C. et al. Percutaneous right ventricular support: Initial experience from the tandemheart experiences and methods (THEME) registry (2018).

41.	Aggarwal, V., Einhorn, B.N., Cohen, H.A. Current status of percutaneous right ventricular assist devices: First‐in‐man use of a novel dual lumen cannula. Catheterization and Cardiovascular Interventions. 88 (3), 390-396 (2016).

42.	Kapur, N.K. et al. Mechanical circulatory support devices for acute right ventricular failure. Circulation. 136 (3), 314-326 (2017).

43.	Kapur, N.K. et al. Mechanical Circulatory Support for Right Ventricular Failure. JACC: Heart Failure. 1 (2), 127-134 (2013).

44.	Geller, B.J., Morrow, D.A., Sobieszczyk, P. Percutaneous Right Ventricular Assist Device for Massive Pulmonary Embolism. Circulation: Cardiovascular Interventions. 5 (6), e74-e75 (2013).

45. Bhama, J. et al. Initial Experience with a Percutaneous Dual Lumen Single Cannula Strategy for Temporary Right Ventricular Assist Device Support Following Durable LVAD Therapy. The Journal of Heart and Lung Transplantation. 35 (4), S323 (2013).

46.	O'Neill, B. et al. Right ventricular hemodynamic support with the PROTEKDuo Cannula. Initial experience from the tandemheart experiences and methods (THEME) registry category: Miscellaneous (2018).