# Submission ID #: 62190

# Scriptwriter Name: Bridget Colvin

Project Page Link: https://www.jove.com/account/file-uploader?src=18967998

Title: A 3D Cartographic Description of the Cell by Cryo Soft X-Ray Tomography

Authors and Affiliations: Johannes Groen1, Andrea Sorrentino1, Lucía Aballe1, Robert Oliete1, Ricardo Valcárcel1, Chidinma Okolo2, Ilias Kounatidis2, Maria Harkiolaki2, Ana. Joaquina Pérez-Bemá1, and Eva Pereiro1

1MISTRAL Beamline, Experiments Division, Alba Light Source

2Beamline B24, Diamond Light Source, Harwell Science and Innovation Campus

☐   All author names are spelled correctly, and the affiliations are correct (city/state/country information not included in video title page).

# Corresponding Author:

# Eva Pereiro

# epereiro@cells.es

# Co-authors:

jgroen@cells.es

asorrentino@cells.es

laballe@cells.es

roliete@cells.es

rvalcarcel@cells.es

chidinma.okolo@diamond.ac.uk

ilias.kounatidis@diamond.ac.uk

maria.harkiolaki@diamond.ac.uk

anperez@cells.es

## Author Questionnaire

1. Microscopy: Does your protocol demonstrate the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or similar?  N

2. Software: Does the part of your protocol being filmed demonstrate software usage?  Y

If Yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps. Please upload all screen captured video files to your project page as soon as possible.

3. Interview statements: Considering the Covid-19-imposed mask-wearing and social distancing recommendations, which interview statement filming option is the most appropriate for your group? Please select one.

☐ 	Interviewees wear masks until the videographer steps away (≥6 ft/2 m) and begins filming. The interviewee then removes the mask for line delivery only. When the shot is acquired, the interviewee puts the mask back on. Statements can be filmed outside if weather permits.

☐ 	Interviewees self-record interview statements outside of the filming date. JoVE can provide support for this option.

☐ 	Interview Statements are read by JoVE’s voiceover talent.

☐ 	Author interview statement opt out (interview statements removed completely from paper).

3. Filming location: Will the filming need to take place in multiple locations (greater than walking distance)?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

- To ensure that your protocol can be filmed in one day, the protocol is restricted to 25 steps and/or 55 shots.

Protocol Length

Number of Shots: 35

# Introduction

- Introductory Interview Statements

Your answers to these questions will become author interview statements, which will be incorporated into the script. Authors will memorize the statements and then deliver them on camera.

- Enter the full name of the author who will deliver the statement on camera.

- Each author should deliver no more than two statements.

- Fill out both required statements. One optional statement may also be selected.

- Please answer in full sentences in a style suitable for being spoken aloud.

- Limit the length of each statement to 30 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: Why is your protocol significant? OR What key questions can this method help answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

REQUIRED: What is the main advantage of this technique?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Do the implications of this technique extend toward the therapy (or diagnosis) of a particular disease, disability, or challenge? How so?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Are there any specific areas of research that this method could provide insight into? OR Can this method be applied to any other systems?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: How would you expect an individual who has never performed this technique to struggle? Do you have any advice to offer to somebody who is trying this technique for the first time?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL: Why is visual demonstration of this method critical?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Has any author been assigned more than two statements? ☐ No

# Introduction of Demonstrator on Camera

Complete this statement ONLY if any of the individuals who will be demonstrating the procedure on camera will not be delivering an Introductory Interview Statement.

  - Enter name of author who will introduce demonstrator: Demonstrating the procedure will be Click here to enter name of demonstrator(s)., a Click here to enter demonstrator job title. from my laboratory. Include additional demonstrators as needed.[1][2].

    - INTERVIEW: Author saying the above

    - The named demonstrator(s) looks up from workbench or desk or microscope and acknowledges the camera

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The one-digit numbers represent sections of the video. The text will appear onscreen.

- The two-digit numbers (e.g. 2.1., 2.2.) represent steps of your protocol. The text will be recorded by a professional voiceover talent.

- The three-digit numbers (e.g. 2.1.1., 2.2.2.) represent the shots that our videographer will capture at your lab.

Please use this draft script to help you prepare for filming day.

- Filming should take no more than 10 minutes per step. If a step will take more than 10 minutes, prepare the product from that step in advance.

- Transmission X-Ray Microscope (TXM) Loading

  - To load samples into the transmission X-ray microscope, cool the transfer chamber with liquid nitrogen until it reaches less than 100 Kelvin [1], fill the workstation [2], and turn on the heater of the workstation rim [3].

    - WIDE: Talent cooling chamber

    - Talent filling workstation

    - Talent turning on heater

  - When the workstation rim stops boiling, place the cryo-box-containing grids into the appropriate locations in the workstation under cryo conditions [1] and load the grids onto the previously cooled sample holders [2].

    - Talent placing grid(s) into workstation

    - Grids being loaded onto sample holders

  - Load the holders onto the shuttle [1] and protect them with the covers [2].

    - Talent loading holder onto shuttle

    - Talent protecting holder with cover

  - Load the shuttle into the transfer chamber at less than 100 Kelvin [1] and pump the chamber down to low vacuum [2].

    - Talent loading shuttle into chamber

    - Talent pumping chamber

  - Attach the transfer chamber to the microscope [1] and follow the vacuum procedure on the screen to load the transfer chamber shuttle into the microscope [2].

    - Talent attaching chamber to microscope

    - Talent following procedure/loading shuttle into microscope

  - Once the shuttle is the microscope with the samples, use the microscope robot arm to transfer one sample holder to the sample stage [1].

    - Arm transferring holder to stage

- TXM Brightfield Imaging

  - For brightfield imaging of the grid with the on-line visible light microscope, select the visible light microscope camera [1] and turn on the visible light microscope-led source for brightfield imaging [2].

    - WIDE: Talent selecting microscope camera, with monitor visible in frame

    - SCREEN: To be provided by Authors: Microscope, Acquisition, Acquisition Settings, Source Settings, and Transmission being selected

  - Click Motion, Control, Sample, and Sample theta to rotate the sample to minus 60 degrees so that it faces the visible light microscope objective and select Motion Control and Sample and change Sample X to Sample Y to move the sample to the expected centered positions [1].

    - SCREEN: To be provided by Authors: Motion, Control, Sample, and Sample theta being clicked, then Motion Control and Sample being clicked, then Sample X being changed to Sample y

  - Select Microscope, Acquisition, Acquisition Settings, Acquisition Modes, Continuous, and Start and click Motion Control and Sample to select Sample Z to refine the focus with smaller steps down to 5 microns until the cells and/or the holes of the carbon support film are in focus [1].

    - SCREEN: To be provided by Authors: Microscope, Acquisition, Acquisition Settings, Acquisition Modes, and Continuous Start being selected, then Motion Control and Sample being clicked and Sample Z being selected/cells or holes coming into focus

  - Then select Microscope, Acquisition, Acquisition Settings, Acquisition Modes, Mosaic, and Start to start the acquisition of a full mosaic map of the grid in brightfield mode using the default values for the mosaic [1].

    - SCREEN: To be provided by Authors: Microscope, Acquisition, Acquisition Settings, Acquisition Modes, Mosaic, and Start being selected/map being acquired

- Fluorescence Mode Mosaic Acquisition

  - For fluorescence mode mosaic imaging, turn off the visible light microscope-led source for brightfield imaging [1] and select the LED light source corresponding to the desired excitation wavelength and the corresponding optical filter manually on the set-up [2].

    - WIDE: Talent turning off light source, with monitor visible in frame

    - SCREEN: To be provided by Authors: LED light source being selected

  - Select Microscope, Acquisition, Acquisition Settings, Acquisition Modes, Continuous, and Start and click Motion Control and Sample to select Sample Z to refine the focus on the fluorescence image [1].

    - SCREEN: To be provided by Authors: Microscope, Acquisition, Acquisition Settings, Acquisition Modes, Continuous, and Start being selected, then Motion Control and Sample being clicked, then Sample Z being selected

  - Then click Microscope, Acquisition, Acquisition Settings, Acquisition Modes, Mosaic, and Start to acquire a mosaic map retaining the positional X and Y parameters from the bright field mosaic [1].

    - SCREEN: To be provided by Authors: Microscope, Acquisition, Acquisition Settings, Acquisition Modes, Mosaic, and Start being clicked/map being acquired

  - Then switch off the LED light-source [1].

    - SCREEN: To be provided by Authors: LED light-source being switched off

- X-Ray Imaging

  - For X-ray mosaic acquisition, select a binning of 2 [1], an exit slit of 5 microns to minimize the irradiation, 1-second exposure at Mistral, a binning of 1 and a B24 (B-twenty-four) of 60 microns [2].

    - WIDE: Talent setting binning, with monitor visible in frame

    - SCREEN: To be provided by Authors: Exit slit, exposure, binning, and B24 being set

  - Select Microscope, Acquisition, Acquisition Settings, and Camera Settings to select binning to adjust the focus using sample Z translation. Then select Motion Control and XS (X-S) to select XS [1].

    - SCREEN: To be provided by Authors: Microscope, Acquisition, Acquisition Settings, Camera Settings, and binning being selected, then Motion Control and XS and XS being selected

  - Select Microscope, Acquisition, Acquisition Settings, Acquisition Modes, Continuous, and Start. Select Motion Control and Sample to select Sample Z. Starting in steps of 5 microns, refine the focus down to steps of 0.5 microns until the cell or the carbon foil holes are well in focus [1].

    - SCREEN: To be provided by Authors: Microscope, Acquisition, Acquisition Settings, Acquisition Modes, Continuous, and Start being selected, then Motion Control, Sample, and Sample Z being selected/focus being refined/cells or holes coming into focus

  - To acquire a mosaic map of the mesh square, click Microscope, Acquisition, Acquisition Settings, Acquisition Mode, Mosaic, and Start [1].

    - SCREEN: To be provided by Authors: Microscope, Acquisition, Acquisition Settings, Acquisition Mode, Mosaic, and Start being clicked

  - When the map has been acquired, click Motion Control and Sample and change Sample X to Sample Y to move the sample to a Flat-Field position [1].

    - SCREEN: To be provided by Authors: Motion Control and Sample being clicked, then Sample X being changed to Sample Y

  - Set the exposure time to 1 second at Mistral and half a second at B24 and acquire a single image [1].

    - SCREEN: To be provided by Authors: Parameters being set and image being acquired

  - Then normalize the acquired mosaic by the flat-field image to obtain the transmission and save the normalized mosaic [1].

    - SCREEN: To be provided by Authors: Mosaic being normalized/transmission being obtained, then mosaic being saved

- Tomography Acquisition

  - To acquire a tomogram, click Motion Control and Sample [1] and select Sample theta to move the negative maximum angle to plus 0.1 [2].

    - WIDE: Talent clicking Motion Control and Sample

    - SCREEN: To be provided by Authors: Sample theta being selected/negative maximum angle being moved to plus 0.1

  - Click Microscope, Acquisition, Acquisition Settings, Acquisition Modes, and Tomography to set the number of images as the total number of angles, taking into account the image at angle zero and the angular range, and set the number of images [1].

    - SCREEN: To be provided by Authors: Microscope, Acquisition, Acquisition Settings, Acquisition Modes, and Tomography being selected, then number of images being set, then number of images being set

  - Select Angle Start and Angle End, click Microscope, Acquisition, Acquisition Settings, Camera Settings, and set the Exposure Time [1].

    - SCREEN: To be provided by Authors: Angle Start and Angle End being selected, then Microscope, Acquisition, Acquisition Settings, Camera Settings, and Exposure time being set

  - Then click Start to start the acquisition [1].

    - SCREEN: To be provided by Authors: Start being clicked, then acquisition being started

## Protocol Script Questions

# Authors: Please use the step numbers from the script above (not step numbers from the manuscript) when answering the questions below. Please do not include steps that will be screen-captured and do not list entire sections.

A. Which steps from the protocol are the most important for viewers to see? Please list 4 to 6 individual steps.

Click here to list 4 to 6 individual steps, using the step numbers from the protocol section of the video script.

B. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1 or 2 individual steps from the script above.

Click here to list 1 or 2 individual steps, using the step numbers from the protocol section of the video script.

# Results

Please review this section to make sure that it accurately reflects your findings.

- Use Track Changes when making edits or revisions.

- If you would like the video to include different results, please revise this section.

- When revising, please keep the length of the voiceover below 200 words. Current word count: 112. (Voiceover is the text that follows the two-digit numbers.)

- Please note that the video cannot include voiceover without an accompanying visual.

- Notes in blue italics are for JoVE’s video editor.

- Results: Representative Cryo Soft X-Ray Tomograms

  - The ideal sample [1] should have single cells at the center of a square mesh embedded in a thin layer of ice and surrounded by well-dispersed gold fiducial markers [2].

    - LAB MEDIA: Figure 5A

    - LAB MEDIA: Figure 5A Video Editor: please emphasize red square/cells in red square

  - Many different organelles [1], such as mitochondria [2], endoplasmic reticula [3], vesicles [4], and the nucleus should be able to be distinguished thanks to the quantitative reconstruction of the linear absorption coefficients [5].

    - LAB MEDIA: Figures 5A and 5B Video Editor: please zoom into red square to show magnification in 5B or similar

    - LAB MEDIA: Figure 5B Video Editor: please emphasize M texts and organelles

    - LAB MEDIA: Figure 5B Video Editor: please emphasize ER text and organelle

    - LAB MEDIA: Figure 5B Video Editor: please emphasize V texts and organelles

    - LAB MEDIA: Figure 5B Video Editor: please emphasize N text

  - In this image, a square with higher cell density can be observed [1], as observed, the blotting was less efficient, leading to a thicker ice layer with cracks [2].

    - LAB MEDIA: Figure 5C

    - LAB MEDIA: Figure 5C Video Editor: please emphasize red arrows

  - Even though some larger structures can be recognized in this preparation [1], fine details were lost within the noise and grainy texture due to the poor vitrification quality of the thick ice [2].

    - LAB MEDIA: Figures 5C and 5D Video Editor: please zoom into red square to show magnification in Figure 5D or similar

    - LAB MEDIA: Figure 5D Video Editor: please red arrow

# Conclusion

- Conclusion Interview Statements

Below are prompts for interview statements that can be used to further emphasize the significance of your protocol.

- Please answer one or two of the statements.

- Each statement is limited to 30 words.

- Answer the questions in full sentences; you will need to memorize and deliver the sentences as spoken interview statements during filming.

- Indicate the full name of the author who will give each statement.

# What is the most important thing to remember when attempting this procedure? Please indicate the steps (e.g., 2.4., 2.5.) in the Protocol section of the script that this advice applies to.

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.[1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera (Enter step numbers referred to.)

# Following this procedure, what other methods can be performed? What questions would these additional methods answer?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

# After its development, did this technique pave the way for researchers to explore new questions within a specific scientific field? If so, how?

  - Enter author name: Click here to answer. Please use language that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words. [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Is each interview statement 30 words or fewer? ☐ Yes

Thank you for addressing our questions. We will incorporate your answers and suggestions and send you the final script before your filming day. You will also receive detailed preparation instructions in the email accompanying the final script.