Editorial comments:
1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues.

The manuscript has been proofread closely again and all identified grammatical and formatting errors corrected, to the best of our knowledge.

2. Please provide an institutional email address for each author.

This has been added

3. JoVE cannot publish manuscripts containing commercial language. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials: e.g., Spotiton, Picosystem, Hydro System and Supplies, etc. We must maintain our scientific integrity and prevent the subsequent video from becoming a commercial advertisement.

Specific information relating to make/model of the purified water system to the supplemental materials document has been moved and we do not believe any other references to specific products remain in the protocol.

4. Lines 76-87: Please consider removing the protocol outline.

This has been removed.

5. Line 129-134/184-188/ 349-362: The Protocol should contain only action items that direct the reader to do something. Please move the discussion about the protocol to the Discussion.

Non-instructional text from these three sections has been modified and/or relocated.

6. Line 204-223: Please add more details to your protocol steps. Please ensure you answer the “how” question, i.e., how is the step performed? Alternatively, add references to published material specifying how to perform the protocol action.

A change to step 2.1.1. was also requested by reviewer #2 and has been modified to a note. We feel that steps 2.1.2 through 2.1.4 deal with sample dilution, filling a standard cryogen bowl with liquid nitrogen, and plasma cleaning a grid, actions that are not specific to this protocol, rather fall within the realm of general cryo-EM laboratory practice.

7. Please highlight up to 3 pages of the Protocol (including headings and spacing) that identifies the essential steps of the protocol for the video, i.e., the steps that should be visualized to tell the most cohesive story of the Protocol. Remember that non-highlighted Protocol steps will remain in the manuscript, and therefore will still be available to the reader.

In the submitted Word file, roughly three pages of the protocol is highlighted in yellow as requested. Did this highlighting not appear in the document?

I will say that this task was difficult as we view all steps as “essential” to this protocol. If the highlighted portion is still too much for the video, we are of course willing to work with the editorial staff to tweak the selection to just those steps that will appear in the video.

8. Please do not use the &-sign or the word “and” when listing authors. Authors should be listed as last name author 1, initials author 1, last name author 2, initials author 2, etc. End the list of authors with a period. Example: Bedford, C. D., Harris, R. N., Howd, R. A., Goff, D. A., Koolpe, G. A. Quaternary salts of 2-[(hydroxyimino)methyl]limidazole. Journal of Medicinal Chemistry. 32 (2), 493-503 (1998).

Done.

9. Please reduce the number of figures presented. Some of the figures may be unnecessary with the video of the protocol. Please consider moving some figure to the supplementary figures section.

Of the 18 figures submitted, I have selected 5 for the main article, 5 as supplementary material, and removed 8. I am not clear on the specific number that is allowed, but hope these adjustments are acceptable.

10. Please ensure all figures showing data includes measurement definitions and scale bars (if applicable).

Scale bars have been added to figures that warrant them.
____________________________________
Reviewers' comments:
Reviewer #1:
1. It is not clear who is affiliated with 3Engineering Arts LLC, Arizona, USA from the author list.

The author list and affiliations have been updated.

2. It would be great if the authors could add in the introduction the method description paper from Frank and White (Time-resolved cryo-electron microscopy using a microfluidic chip S Kaledhonkar, Z Fu, H White, J Frank, Protein Complex Assembly, 59-71) and two relevant applications from Frank group (Late steps in bacterial translation initiation visualized using time-resolved cryo-EM; The structural basis for release-factor activation during translation termination revealed by time-resolved cryogenic electron microscopy).

We thank the reviewer for suggesting the inclusion of these papers and have included reference to them in the introduction.

3. The protocol is straightforward and clear. However, it would be great if the authors could address if there are practical differences comparing Spotiton and Chameleon. It would be more helpful for users of Chameleon to develop their time-resolved cryo-EM experiments.

At this time, the conduct of time-resolved studies in the manner reported here is not possible using chameleon. In the discussion (lines 532-535), we raise the possibility that such experiments could be carried out on chameleon should a second sample dispenser be added.

4. In the discussion, it would be nice to discuss what steps are new to improve the current protocol's reproducibility compared to the previous nature method paper.

Additional text describing specific upgrades to the system made since the previous publication reporting time-resolved Spotiton (Dandey et al, 2020), has been added to the Discussion. 

Reviewer #2:
Manuscript Summary:
The manuscript by Budell et al. provides a detailed protocol on how to prepare cryo-EM grids for time-resolved studies, using Spotiton. Time-resolved cryo-EM is under active development and of high importance due to the additional dimension of time, missing in conventional cryo-EM studies. The technology on its own is of high relevance, and the protocol is clear and well written, however, the applicability of this step-by-step protocol is limited due to the limited access to the Spotiton system. Regardless of that, I would still recommend this protocol for publication as it provides a nice overview of the procedure and the troubleshooting section together with the figures of representative results may be useful for Chameleon users.

Major concerns:
This protocol is supposed to be focused on sample preparation for time-resolved studies, however, no guidelines are provided on how to control the mixing time of the two samples, which is the crucial aspect for such studies.

Section 2.1 has been edited to describe how to modify estimated mixing times. 

Minor concerns:
Line 197: "1.3.11. If both tips fired successfully and as expected, the system is ready to prepare sample grids." But what if not? Simply repeat the procedure? It would be good to add troubleshooting here.

Section 1.3.10 has been edited to address this concern.

"2.1.1. Put aliquots of your samples on ice now for the ~20-30 minutes it will take before the first grid is prepared." Is this the requirement for the successful vitrification with Spotiton? There might be cases where scientist would prefer to keep the sample at room temperature or even at 37 degrees prior to grid preparation in order to activate a certain state, therefore, if possible, I would totally remove this sentence and allow the users to decide themselves how to store their samples prior to grid preparation, unless it compromises the subsequent protocol in some way.

Section 2.2 has been edited to address this concern

„2.1.4. Plasma clean 3-4 nanowire grids. Use 5W, hydrogen and oxygen, 1.5 minutes as a starting point." Not everyone has a hydrogen generator in the lab (ignoring the fact that even fewer have Spotiton/Chameleon), so from practical point of view, it would be useful to comment on the suitability of other mixes of gasses. Would simply glow-discharging work for nanowire grids as well?

It is certainly likely that either plasma cleaning with an alternate mixture of gases or glow discharging would effectively treat the surfaces of nanowire grids for use in Spotiton. We have added a sentence to the note under Section 2.2.3 indicating the potential effectiveness of other surface treatment methods; however, for the current protocol and for all experimentation reported in our previous publication, we used only the treatment protocol described here.  Further investigation into the effect of alternative surface treatment protocols is of interest to us, but we feel is beyond the scope and intent of this report. 

Reviewer #3:
Manuscript Summary:
The manuscript is well written and describes many useful pieces of information for using the Spotiton system for time resolved studies.

Concerns:
From the images presented, the extent of specimen mixing is not clear to me. Is it possible that the first specimen is wicked completely by the grid bars while the second specimen forms the layer of ice (that doesn't contain the first specimen) observed in the figures? Can the authors point to other results that show that the two specimens are thoroughly mixed? Alternatively, it may be possible to do an Unwin & Berriman-type experiment where two easily identifiable particles are used as the different specimens to show that there is thorough mixing on the grid.

We do not feel it is possible that the first sample is wicked completely prior to application of the second sample. Two points support this idea: first, in our prior publication of this method (Dandey et al, 2020), we present EM data for 4 time-resolved experiments for which on-grid mixing is the only explanation for the observed results; second, were the first sample to wick completely in the approximately 7 ms before the second sample is applied, by the time the grid enters the cryogen roughly 100 ms later, the aqueous layer comprised of both samples would have thinned so much that the holes would be emptied of liquid. In light of this, we believe the second sample is hitting a bulk volume of the first already present on the grid and mixing between them occurs as the liquid thins on the grid until the cryogen is reached.

The authors should make clear whether this experiment can be performed with a commercially available Chameleon device. Is the spotiton device used for these experiments modified relative to the publications describing the spotiton device?

In the Discussion (lines 532-535), we indicate that chameleon is not currently capable of two-sample mixing experiments but could be made capable by addition of a second piezo dispenser. Mention of modifications to the system made subsequent to the previous publication reporting the time-resolved Spotiton method have been added to the Discussion (lines 592-596.)

Reviewer #4:
Manuscript Summary:
This paper describes the procedures for using the Spotiton system for making cryoEM grids, including its exciting new capability of performing fast mixing for time-resolved experiments. This is an important new technique and very much warrants description in this way.

Major Concerns:
I'd like to see how the time delay between the two sprays is controlled, as this is a key feature of the protocol. Also, how did you determine the listed 90ms minimum mixing time?

Both dispensers initiate and cease firing simultaneously under all conditions. The reported mixing times are calculated using distances between different machine components and conditions used when a grid is made, such as acceleration, deceleration, and maximum velocity of the grid robot. The adjustment of these conditions has been newly addressed in section 2.1.

Minor Concerns:
Step 2.2. Does higher humidity decrease the wicking speed? Maybe point out.

The third paragraph of the Discussion addresses the effect of higher or lower humidity on the speed of wicking on nanowire grids.

Step 2.4.10. I take it that "hopping" the grid is done by the robot, not the user? So the robot deposits the grids onto the shelf.

Yes. A slight edit was made to this section (now: 2.5.10) to clarify this point.

Fig. 2b(3) and Fig. 6(right) is the alternating pattern of large and small drops normal? Desirable?

This pattern of droplet size is typical and expected of this type of piezo dispenser.

Fig. 9A I couldn't see the black arrow in the copy of the figure I printed from the jpeg I was sent. Likewise Fig. 15, I couldn't see white arrowheads.

Figure 9 has been removed from the manuscript and white arrowheads have been added to the indicated figure.

Reviewer #5:
Manuscript Summary:
This work by Budell et al present a step-by-step protocol for producing time-resolved cryo-EM grids using the Spotiton system, following from their paper "Time-resolved cryo-EM using Spotiton" in Nature Methods last year. Importantly provide in depth technical detail on optimization and quality control while using the system which are of great importance to the community. To validate their protocol they provide several examples of how a high-quality time-resolved grid can be produced using the method.
This is important addition to their valuable body of work and clearly deserves publication and attention.
In my opinion, it will be of further value to the community, if the authors could use this forum to further elaborate on some of the points below, which remain insufficiently discussed in the original paper and also are not further clarified in this work.

Major Concerns:
Unlike other time-resolved preparation methods, the technique used here practically combines the blotting (by means of self wicking), mixing and incubation times into one motion of the machine that cannot be separated easily. This raises several idiosyncratic issues which are worthy of further discussion:

* It is my understanding that the method outlined here does not easily generalize to interrogate a wide range of timepoints without first optimizing every step of the grid making process individually for the respective reaction timepoint, potentially requiring the user to even go back to the initial step of making the nanowire grid to find usable conditions. It would be helpful, if the authors would comment on this potential issue - both in terms of addressing it by experts as themselves but also by independent users of the system.

An additional sentence at the top of the third discussion paragraph connects the ideas of sample mixing times and wicking speed. The subsequent explanation of how adjusting specific factors can impact wicking speed should provide the reader guidance on the steps that can be taken to facilitate sample mixing of a particular duration.

* Even more importantly, a more detailed discussion of the mixing of reactants will facilitate better judgement of the method's strengths and weaknesses. If one was to assume free diffusion in a aqueous layer of less than 100nm thick, it is true that mixing would be nearly instant, however analysis of grids prepared by spotiton using tomography have revealed that up to 94% of proteins tend to be associated with the air-water interface (Muench et al 2020). This association is less of a problem for microfluidic systems where mixing and any potential structural rearrangements have already taken place before the sample lands on the grid (and likely gets stuck at the air-water interface). However here, the interaction between sample 1 and 2 takes place on the grid itself and it remains unclear whether the mobility of these proteins is sufficiently similar to the state in solution: once they have been blotted sufficiently, they are most likely to be stuck at the air-water interface, according to the above cited work, limiting their diffusion substantially.

(The reviewer is also directed to our response to reviewer 3..) Briefly, we agree the on-grid mixing we describe here is distinct from the ‘pre-grid’ two-sample mixing that characterizes microfluidic systems. However, we do not feel the first sample has sufficiently thinned (been ‘blotted’ by the nanowires) for the effects of AWI to come into play in a significant way before the second sample arrives. The second sample lands on, and mixes with, the first sample in a bulk state.

* Following from this, it may be increasingly more challenging to compare data from different incubation times, especially with regards to comparatively very short vs very long time points. Beyond the conceptual concern, it would be equally valuable if the authors could please comment on the potential to "tune" nanowire blotting for a reaction time of several seconds or more.

We agree that such a direct comparison between particularly short or extended sample mixing durations presents unique experimental challenges. However, the intended range of mixing times using Spotiton is between 90-500ms. Experiments with timescales in the second or multi-second range are not possible using Spotiton due to limitations of nanowire grid wicking capacity, so we feel a discussion of such longer duration experiments is not warranted at this time.

Minor Concerns:
Lines 210-212 - The note here says that the sample does not have time to migrate into the holes, requiring more sample to be used. However, there is substantial evidence in the time-resolved EM field that the concentration effect from standard blotting comes from proteins getting stuck at the air-water interface during blotting, e.g. Muench et al 2020, thus giving the appearance of more concentrated sample. In ohter words, is the simpler (and valid) explanation that migration into the holes is not the primary attractor. What we see in time-resolved preparation has lower concentrations because there is not enough time for a denatured layer at the air-water interface, a la Glaeser, to serve as the attractor for sample enrichment. Given the prominence of the authors in the field, it may be a great opportunity to make their nuanced opinion on the matter more clear.

The reviewer raises a valid point and the explanation in the note accompanying section 2.2.1 has been modified.