# Submission ID #:  62271

# Scriptwriter Name: Domnic Colvin

# Supervisor Name: Anastasia Gomez

# Project Page Link: https://www.jove.com/account/file-uploader?src=18992703

# Title:   Cryo-Electron Microscopic Grid Preparation for Time-resolved Studies Using a Novel Robotic System, Spotiton

# Authors and Affiliations:

William C. Budell1, Luis Allegri2, Venkat Dandey1, Clinton S. Potter1, Bridget Carragher1

1The National Resource for Automated Molecular Microscopy, Simons Electron Microscopy Center, New York Structural Biology Center, New York, NY, USA

2piTree Software, Metuchen, NJ, USA

# Corresponding Authors:

Bridget Carragher (bcarr@nysbc.org)

# Email Addresses for All Authors:

# wbudell@nysbc.org

# luisallegri@msn.com

# ven820212@gmail.com

cpotter@nysbc.org

bcarr@nysbc.org

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes, all done

3. Interview statements: Considering the COVID-19-imposed mask-wearing and social distancing recommendations, which interview statement filming option is the most appropriate for your group? Please select one.

☒ 	Interviewees self-record interview statements. JoVE can provide support for this option.

4. Filming location: Will the filming need to take place in multiple locations?   No

Current Protocol Length

Number of Steps:   19

Number of Shots:   43

# Introduction

- Introductory Interview Statements

REQUIRED:

  - Chase Budell: Using the Spotiton robotic system, one can mix and vitrify a protein of interest with an interacting partner on an electron microscope grid in as fast as 90 milliseconds.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. B-roll: Figure 3, 3.7.1. and 3.8.2.

  - Chase Budell: This protocol allows the capture of intermediate protein conformations that are too transient to be captured by standard grid preparation techniques.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. B-roll: Figure 3.11.1.

OPTIONAL:

  - Chase Budell: Time-resolved Spotiton can inform sub second biological or biochemical systems such as membrane channel activation, DNA and RNA synthesis, or early interaction of a drug or antibody with its target protein.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Chase Budell: The user must concurrently manage several components that directly impact the quality of a grid. It is important to understand the system before using and have patience when preparing grids.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. B-roll: 3.2.1., 3.10.2.

# Protocol

- Set up the Spotiton Machine and Software

  - To begin, open the main valve on the nitrogen supply tank and ensure the system reservoir is filled with degassed, ultrapure water [1]. Turn on the computer [2] and the Spotiton system at the multioutlet powerstrip [3].

    - Talent opening the main valve on the nitrogen supply tank.

    - Talent turning on the computer. Videographer: Obtain a few shots of talent clicking the mouse and typing on the keyboard to use as b-roll throughout the video.

    - Talent turning on the Spotiton system.

  - Click on the desktop icon to open the user interface of the Spotiton software. In the Tools menu, select Initialize Stages to initialize and home the 3-axis robots and the rotating dispenser head assembly [1], ensuring the dispenser tips are pointing down prior to initialization and homing [2].

    - SCREEN: 62271_screenshot_1.mp4. 00:00-00:40.

    - Robot initializing (robot action).

  - On the main window, click on Go to SafePosition [1] to send the robots to the SafePosition [2]. On the Aspirate tab, select Prime to flush the dispenser heads multiple times with water from the reservoir. Continue until two uninterrupted streams of water can be seen emerging from the tips [3].

    - SCREEN: 62271_screenshot_2.mp4. 00:01-00:02.

    - Robot moving to safe position (robot action).

    - SCREEN: 62271_screenshot_2.mp4. 00:03-00:10.

  - On the Inspect tab, send Tip 1 to the inspection camera and test-fire water, adjusting amplitude until discrete droplets are produced [1]. Press the Record button in the upper camera monitor to record a video of Tip 1 firing [2].

    - SCREEN: 62271_screenshot_3.mp4. 00:00-00:30.

    - Tip 1 moving to inspection camera (robot action).

  - Send Tip 2 to the inspection camera. Play back the video of Tip 1 firing in the right-side monitor at the same time Tip 2 is fired in the upper camera monitor [1]. Match the pattern of droplet production from the two dispensers [2].

    - SCREEN: 62271_screenshot_3.mp4. 00:30-01:07.

    - Tip 2 moving to the inspection camera (robot action).

  - Remove the tweezers from the mount on the grid robot using the supplied Allen key [1]. On a nearby benchtop, position a test grid, nanowire side up, on the edge of the grid block. Carefully grab the rim of the grid and position it correctly in the tweezers [2], then remount the tweezers [3].

    - Talent removing the tweezers form the mount using Allen key.

    - Talent grabbing and positioning the rim of the grid with tweezers.

    - Talent remounting the tweezers.

  - On the Cryo tab, click on Tip to Camera to move Tip 1 into the field of view of the upper plunge path camera [1], ensuring that Live is selected in the upper camera monitor [2]. Turn on and adjust the upper camera light [2].

    - SCREEN: 62271_screenshot_4.mp4. 00:00-00:11.

    - Tip moving to camera (robot action).

    - Talent adjusting upper camera light or SCREEN: 62271_screenshot_4.mp4. from 00:11-00:14

  - Position Tip 1, visible in the upper camera monitor, by clicking the mouse within the monitor [1]. Click Grid to Camera [2] to position the grid in front of the upper camera [3], then adjust Tip 1 position again if needed [4].

    - SCREEN: 62271_screenshot_4.mp4. 00:14-00:19.

    - SCREEN: 62271_screenshot_5.mp4. 00:00-00:08.

    - Grid moving to camera (robot action).

    - SCREEN: 62271_screenshot_5.mp4. 00:08-00:14.

  - On the Cryo tab, ensure Vitrify Grid is not selected, click on Queue Target [1], then on Plunge [2]. Evaluate the upper and lower images [3].

    - SCREEN: 62271_screenshot_6.mp4. 00:00-00:17.

    - Test plunge (robot action).

    - SCREEN: 62271_screenshot_6.mp4. 00:17-00:20.

- Prepare Two-sample, Mixed Grids

  - Dilute two samples to the desired concentrations with an appropriate buffer, ideally using the same for both [1], and fill the cryogen bowl with liquid nitrogen [2].

    - Talent diluting the two samples with buffer.

    - Talent filling the cryogen bowl with liquid nitrogen.

  - Plasma clean 3 to 4 nanowire grids using 5 Watts, hydrogen and oxygen, and 1.5 minutes as a starting point [1]. Plug in the nebulizer, and observe the vapor exit the central port on the nebulizer cap [2].

    - Talent plasma cleaning the nanowire grids.

    - Talent plugging in the nebulizer; show vapor exiting the nebulizer.

  - Observe the live humidity monitor in the main window or open the ambient humidity tracker under Reports and Ambient. Check humidity levels in the Chamber and Shroud zones [1].

    - SCREEN: 62271_screenshot_7.mp4. 00:13-00:24.

  - Add 5 microliters of each sample into the sample cups [1]. Load the sample cups into the holding tray with the sample for Tip 1 on the left and for Tip 2 on the right [2], then push the tray back into the machine until it seats [3].

    - Talent adding samples into small cups. Videographer: This step is important!

    - Talent loading the samples into the holding tray.

    - Talent pushing the holding tray back into the machine.

  - On the Aspirate tab, select 3 microliters for the volume to be aspirated by each tip. Ensure the sample tray has been seated securely, click Aspirate [1], and observe the pipette stage move the dispenser heads into the sample cups [2].

    - SCREEN: 62271_screenshot_8.mp4. 00:01-00:20.

    - Pipette stage moving dispenser into sample cups (robot action).

  - Verify successful aspiration of both samples by removing the sample cups and observing a drop in the liquid level [1]. On the Inspect tab, send each tip to the inspection camera to confirm unobstructed dispensing. Adjust the amplitude as needed to match droplet formation from each tip [2].

    - Talent removing the sample cups to observe drop in liquid.

    - SCREEN: Sending of each tip to the inspection camera and adjusting amplitude. Videographer: Film the screen for this shot, it’s important!

  - Load a freshly plasma cleaned grid into the tweezers, but do not mount the tweezers yet. Ensure that the humidity level is elevated to approximately 90 to 95% [1]. Fill the ethane cup and perform a final test-fire of both tips in front of the inspection camera, confirming no obstruction [2]. On the Cryo tab, click on Tip to camera [3].

    - Talent loading plasma cleaned grid into tweezers.

    - Talent filling the ethane cup.

    - SCREEN: 62271_screenshot_9.mp4. 00:01-00:10.

  - Check ethane cup. If ethane ice has formed, melt as needed with additional ethane gas [1]. Mount the tweezers with the grid onto the grid stage [2]. On the Cryo tab, click on Grid to camera [3], ensuring Tip 1 is positioned correctly in the upper camera monitor [4].

    - Talent melting the ethane ice forms using ethane gas.

    - Talent mounting the tweezers with the grid.

    - SCREEN: 62271_screenshot_9.mp4. 00:11-00:22.

    - Properly positioned Tip 1 (robot action).

  - Click on Vitrify Grid, Queue Target, then Plunge. Click on OK when prompted to command the grid robot [1] to hop the grid from ethane into liquid nitrogen and release it onto the submerged shelf [2]. Examine images of the grid to decide if it should be kept or discarded [3]. Videographer: This step is important!

    - SCREEN: 62271_screenshot_9.mp4. 00:23-00:46.

    - Grid moving into liquid nitrogen (robot action).

    - Talent examining the grid images.

  - If keeping the grid, pre-cool fine-tipped forceps [1], gently grasp the grid by the edge [2], and place it into a grid box slot, starting with the first slot left of the notch and going clockwise [3].

    - Talent pre-cooling forceps.

    - Talent grasping the grid by edge.

    - Talent placing grid into grid box.

  - Use the experiment viewer to review and compare the grid images from the upper and lower cameras along with the machine settings and humidity measurements at the time of the plunge by selecting Reports and Experiment [1].

    - SCREEN: 62271_screenshot_10.mp4. 00:01-00:52. Video editor: Speed up the video.

# Results

- System Camera Images and the Pattern of Ice Deposition on the Spotiton-prepared Grid.

  - Images of grids prepared during a single time-resolved Spotiton session by mixing RNA (R-N-A) polymerase and a 105 base pair DNA (D-N-A) oligomer carrying a promoter sequence for 150 milliseconds prior to vitrification are shown here [1]. Of the six grids, only one shows suboptimal wicking [2].

    - LAB MEDIA: Figure 4.

    - LAB MEDIA: Figure 4F.

  - The pattern of ice deposition on a vitrified grid [1] matches closely with the pattern of deposited liquid seen in the upper camera image [2].

    - LAB MEDIA: Figure 5C.

    - LAB MEDIA: Figure 5B.

  - Effective wicking of the mixed samples occurs along the nanowire-covered grid bars where the sample rarely overflows into squares adjacent to those in which it landed [1].

    - LAB MEDIA: Figure 5D.

  - In ice-filled squares, the ice is typically thickest within holes at the center of the square and becomes thinner in holes closer to the grid bars [1].

    - LAB MEDIA: Figure 5E.

  - Holes immediately adjacent to the grid bars are often empty due to proximity to the nanowires [1].

    - LAB MEDIA: Figure 5F.

# Conclusion

- Conclusion Interview Statements

  - Chase Budell: Proper preparation and handling of the nanowire grids will ensure good ice thickness on mixed sample grids.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B-roll: 3.2.1 and 3.8.2-3.9.1.

  - Chase Budell: The Spotiton system also allows the user to deposit the two samples separately onto a single grid, enabling the collection of an unmixed control during the same gridmaking session.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. B-roll: Figure 3.4.1.

  - Chase Budell: Spotiton has enabled the capture of the first intermediates in bacterial gene expression in real time. Because they form on a subsecond timescale, their structures have been unknown until now.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.